## convers.+compens.
file name: PSt1_Eq4_Sal |  |  | 2008-03-29
author: Dr. Christian Huber
programm for temperature compensated oxygen calculation of membrane PSt1
Fit of the calibration curve is optimized between 0 and 250 % air-saturation (0 and 388 Torr)
enter your values here
input of F0,T0 | 58.87 | phase angle of cal 0
input of F100,T100 | 30.69 | phase angle of cal 100 (water-vapor saturated air)
input of Fm,Tm | 15 | measured phase angle
input of T0 | 17.7 | temperature of cal 0
input of T100 | 17.8 | temperatur of cal100
input of Tm | 17.6 | temperature at measurement
air pressure patm | 1013 | air pressure
Salinity | 32.2 | Salinity
Results of calculation + compensation
[%] air saturation | 494.3167 | %
[%] O2 | 103.3122 | %
pO2 [hPa] | 1028.1589 | hPa
pO2 [Torr] | 771.1847 | Torr
cO2 [mg/L] | 38.861 | mg/L
cO2 [ppm] | 38.861 | ppm
cO2 [µmol/L] | 1214.406 | µmol/L
Internal calculated parameter (DO NOT CHANGE !!)
f1 | 0.801
DF/K | -0.08
DKSV/K | 0.000383
m | 22.9
tan(F0,T100) | 1.6552
tan(F0,Tm) | 1.6563
tan(F100,T100) | 0.59352
tan(Fm,Tm) | 0.26795
KSV,T100 | 0.033331
KSV,Tm | 0.033254
A | 156.5815
B | 14.0252
C | -0.64143
a | 7.8124e-06
b | -0.0021661
c | -0.83822
Chlorinity | 17.8227

## Recalculate Fibox values
file name: PSt1_Eq4_Sal |  |  |  | 2008-03-29
author: Dr. Christian Huber
programm for temperature compensated oxygen calculation of membrane PSt1
Fit of the calibration curve is optimized between 0 and 250 % air-saturation (0 and 388 Torr)
Insert your calibration values
cal0 | 58.87 | ° | T0 | 17.7 | °C
cal100 | 30.69 | ° | T100 | 17.8 | °C
air pressure | 1013 | mbar
Salinity | 32.2 | °/°°
Internal calculated parameter (DO NOT CHANGE !!)
f1 | 0.801 | tan(F0,T100) | 1.6552 | A | 156.5815 | KSV,T100 | 0.033331 | a | 1.7034e-05
DF/K | -0.08 | tan(F0,Tm) | 1.6568 | B | 14.0252 | KSV,Tm | 0.033216 | b | 0.0044844 |  |  |  |  | INSERT YOUR RESPECTIVE DATA
DKSV/K | 0.000383 | tan(F100,T100) | 0.59352 | C | -0.64143 |  |  | c | -0.64645 |  |  |  |  | chamber volume [L] | FW alga [g] | chl a thallus -1
m | 22.9 | tan(Fm,Tm) | 0.58576 |  |  |  |  | Chlorinity | 17.8227 |  |  |  |  | 0.026 | 0.0707 | 0.02667
Insert the data you get from the Ascii File recorded with the OxyView TX3-V5.31.exe software | [7 empty] | The oxygen contents are calculated automatically from the raw data
date | time/hh:mm:ss | logtime/min | oxygen/% airsatur. | phase/° | amp | temp/°C |  | % air-sat. | % oxygen | pO2 (hPa) | pO2 (Torr) | cO2 [mg/L] | cO2 [µmol/L] | cO2 [µmol/L 10 min] | cO2 [µmol/L h] | cO2 [µmol/ h] | cO2 [µmol/ h gFW] | cO2 [µmol/ h mg chl a]
2010-08-12 | 14:49:02 | 0 | 309.015 | 30.36 | 6245 | 17.5 |  | 103.4804 | 21.6274 | 215.2626 | 161.4607 | 8.1506 | 254.7062 | 11.85 | 71.1 | 1.8486 | 26.1471 | 69.3138
2010-08-12 | 14:49:13 | 0.183 | 309.824 | 30.33 | 6240 | 17.5 |  | 103.7512 | 21.684 | 215.826 | 161.8833 | 8.1719 | 255.3728
2010-08-12 | 14:49:23 | 0.35 | 312.537 | 30.23 | 6242 | 17.5 |  | 104.6599 | 21.8739 | 217.7162 | 163.301 | 8.2435 | 257.6093 | change data input according to column N
2010-08-12 | 14:49:33 | 0.517 | 309.824 | 30.33 | 6248 | 17.5 |  | 103.7512 | 21.684 | 215.826 | 161.8833 | 8.1719 | 255.3728
2010-08-12 | 14:49:43 | 0.684 | 310.236 | 30.27 | 6244 | 17.6 |  | 104.1037 | 21.7577 | 216.5316 | 162.4125 | 8.1842 | 255.7553
2010-08-12 | 14:49:53 | 0.851 | 310.236 | 30.27 | 6237 | 17.6 |  | 104.1037 | 21.7577 | 216.5316 | 162.4125 | 8.1842 | 255.7553
2010-08-12 | 14:50:03 | 1.018 | 307.812 | 30.36 | 6244 | 17.6 |  | 103.2902 | 21.5877 | 214.8396 | 161.1434 | 8.1202 | 253.7568
2010-08-12 | 14:50:13 | 1.185 | 309.425 | 30.3 | 6241 | 17.6 |  | 103.8318 | 21.7008 | 215.9659 | 161.9882 | 8.1628 | 255.0872
2010-08-12 | 14:50:23 | 1.352 | 307.544 | 30.37 | 6248 | 17.6 |  | 103.2003 | 21.5689 | 214.6525 | 161.0031 | 8.1131 | 253.5358
2010-08-12 | 14:50:33 | 1.519 | 308.08 | 30.35 | 6252 | 17.6 |  | 103.3803 | 21.6065 | 215.0268 | 161.2838 | 8.1273 | 253.978
2010-08-12 | 14:50:43 | 1.686 | 308.617 | 30.33 | 6248 | 17.6 |  | 103.5606 | 21.6442 | 215.4019 | 161.5652 | 8.1415 | 254.421
2010-08-12 | 14:50:54 | 1.853 | 305.143 | 30.46 | 6250 | 17.6 |  | 102.3948 | 21.4005 | 212.9771 | 159.7464 | 8.0498 | 251.557
2010-08-12 | 14:51:04 | 2.02 | 310.507 | 30.26 | 6252 | 17.6 |  | 104.1946 | 21.7767 | 216.7205 | 162.5542 | 8.1913 | 255.9785
2010-08-12 | 14:51:14 | 2.186 | 307.812 | 30.36 | 6261 | 17.6 |  | 103.2902 | 21.5877 | 214.8396 | 161.1434 | 8.1202 | 253.7568
2010-08-12 | 14:51:24 | 2.353 | 309.965 | 30.28 | 6253 | 17.6 |  | 104.013 | 21.7387 | 216.3428 | 162.2709 | 8.177 | 255.5324
2010-08-12 | 14:51:34 | 2.52 | 307.276 | 30.38 | 6246 | 17.6 |  | 103.1105 | 21.5501 | 214.4656 | 160.8629 | 8.1061 | 253.3151
2010-08-12 | 14:51:44 | 2.687 | 307.276 | 30.38 | 6252 | 17.6 |  | 103.1105 | 21.5501 | 214.4656 | 160.8629 | 8.1061 | 253.3151
2010-08-12 | 14:51:54 | 2.854 | 308.886 | 30.32 | 6248 | 17.6 |  | 103.6509 | 21.663 | 215.5897 | 161.706 | 8.1486 | 254.6428
2010-08-12 | 14:52:04 | 3.021 | 306.474 | 30.41 | 6259 | 17.6 |  | 102.8414 | 21.4939 | 213.9061 | 160.4432 | 8.0849 | 252.6542
2010-08-12 | 14:52:14 | 3.188 | 309.156 | 30.31 | 6246 | 17.6 |  | 103.7413 | 21.6819 | 215.7777 | 161.847 | 8.1557 | 254.8649
2010-08-12 | 14:52:24 | 3.355 | 305.282 | 30.41 | 6258 | 17.7 |  | 102.6526 | 21.4544 | 213.4858 | 160.1279 | 8.0548 | 251.7132
2010-08-12 | 14:52:34 | 3.522 | 309.299 | 30.26 | 6245 | 17.7 |  | 104.0033 | 21.7367 | 216.2949 | 162.235 | 8.1608 | 255.0254
2010-08-12 | 14:52:44 | 3.689 | 306.348 | 30.37 | 6248 | 17.7 |  | 103.0108 | 21.5293 | 214.2308 | 160.6868 | 8.0829 | 252.5917 |  | regression formula | time
2010-08-12 | 14:52:54 | 3.855 | 307.417 | 30.33 | 6261 | 17.7 |  | 103.3705 | 21.6044 | 214.9788 | 161.2478 | 8.1112 | 253.4736 | value for T=26 min. | 272.74 | T11
2010-08-12 | 14:53:04 | 4.022 | 311.195 | 30.19 | 6257 | 17.7 |  | 104.6405 | 21.8699 | 217.6202 | 163.229 | 8.2108 | 256.588 | value for T=1 min. | 260.89 | T1
2010-08-12 | 14:53:14 | 4.189 | 308.491 | 30.29 | 6256 | 17.7 |  | 103.7315 | 21.6799 | 215.7297 | 161.811 | 8.1395 | 254.359 | difference between T25 and T1 | 11.85 | 10minutes
2010-08-12 | 14:53:24 | 4.356 | 309.569 | 30.25 | 6239 | 17.7 |  | 104.0941 | 21.7557 | 216.4837 | 162.3765 | 8.1679 | 255.2479 | calculate from regression curve values for 10 minutes of photosynthesis or respiration
2010-08-12 | 14:53:34 | 4.523 | 309.03 | 30.27 | 6264 | 17.7 |  | 103.9126 | 21.7177 | 216.1063 | 162.0935 | 8.1537 | 254.803
2010-08-12 | 14:53:44 | 4.69 | 307.954 | 30.31 | 6262 | 17.7 |  | 103.5508 | 21.6421 | 215.3539 | 161.5291 | 8.1253 | 253.9159
2010-08-12 | 14:53:54 | 4.857 | 307.685 | 30.32 | 6255 | 17.7 |  | 103.4606 | 21.6233 | 215.1662 | 161.3884 | 8.1182 | 253.6946
2010-08-12 | 14:54:04 | 5.024 | 308.222 | 30.3 | 6245 | 17.7 |  | 103.6411 | 21.661 | 215.5417 | 161.67 | 8.1324 | 254.1373
2010-08-12 | 14:54:14 | 5.19 | 309.299 | 30.26 | 6246 | 17.7 |  | 104.0033 | 21.7367 | 216.2949 | 162.235 | 8.1608 | 255.0254
2010-08-12 | 14:54:24 | 5.357 | 306.615 | 30.36 | 6251 | 17.7 |  | 103.1006 | 21.548 | 214.4175 | 160.8268 | 8.09 | 252.8118
2010-08-12 | 14:54:34 | 5.524 | 308.886 | 30.32 | 6250 | 17.6 |  | 103.6509 | 21.663 | 215.5897 | 161.706 | 8.1486 | 254.6428
2010-08-12 | 14:54:44 | 5.691 | 308.349 | 30.34 | 6255 | 17.6 |  | 103.4704 | 21.6253 | 215.2143 | 161.4244 | 8.1344 | 254.1994
2010-08-12 | 14:54:54 | 5.858 | 307.812 | 30.36 | 6251 | 17.6 |  | 103.2902 | 21.5877 | 214.8396 | 161.1434 | 8.1202 | 253.7568
2010-08-12 | 14:55:04 | 6.025 | 307.008 | 30.39 | 6254 | 17.6 |  | 103.0207 | 21.5313 | 214.2789 | 160.7228 | 8.099 | 253.0946
2010-08-12 | 14:55:14 | 6.192 | 307.008 | 30.39 | 6254 | 17.6 |  | 103.0207 | 21.5313 | 214.2789 | 160.7228 | 8.099 | 253.0946
2010-08-12 | 14:55:24 | 6.359 | 308.886 | 30.32 | 6246 | 17.6 |  | 103.6509 | 21.663 | 215.5897 | 161.706 | 8.1486 | 254.6428
2010-08-12 | 14:55:34 | 6.526 | 307.404 | 30.42 | 6257 | 17.5 |  | 102.9411 | 21.5147 | 214.1408 | 160.6193 | 8.1081 | 253.3788
2010-08-12 | 14:55:44 | 6.693 | 306.603 | 30.45 | 6259 | 17.5 |  | 102.6727 | 21.4586 | 213.5824 | 160.2004 | 8.087 | 252.7181
2010-08-12 | 14:55:54 | 6.86 | 309.015 | 30.36 | 6248 | 17.5 |  | 103.4804 | 21.6274 | 215.2626 | 161.4607 | 8.1506 | 254.7062
2010-08-12 | 14:56:04 | 7.027 | 309.554 | 30.34 | 6252 | 17.5 |  | 103.6609 | 21.6651 | 215.638 | 161.7423 | 8.1648 | 255.1504
2010-08-12 | 14:56:14 | 7.194 | 311.448 | 30.27 | 6261 | 17.5 |  | 104.2953 | 21.7977 | 216.9579 | 162.7322 | 8.2148 | 256.712
2010-08-12 | 14:56:24 | 7.36 | 308.746 | 30.37 | 6257 | 17.5 |  | 103.3903 | 21.6086 | 215.0752 | 161.3201 | 8.1435 | 254.4844
2010-08-12 | 14:56:34 | 7.527 | 307.94 | 30.4 | 6258 | 17.5 |  | 103.1205 | 21.5522 | 214.514 | 160.8992 | 8.1223 | 253.8204
2010-08-12 | 14:56:44 | 7.694 | 308.477 | 30.38 | 6254 | 17.5 |  | 103.3003 | 21.5898 | 214.8879 | 161.1797 | 8.1364 | 254.2628
2010-08-12 | 14:56:54 | 7.861 | 310.906 | 30.29 | 6252 | 17.5 |  | 104.1136 | 21.7597 | 216.5798 | 162.4487 | 8.2005 | 256.2647
2010-08-12 | 14:57:04 | 8.028 | 309.284 | 30.35 | 6257 | 17.5 |  | 103.5706 | 21.6463 | 215.4502 | 161.6014 | 8.1577 | 254.9282
2010-08-12 | 14:57:14 | 8.195 | 308.208 | 30.39 | 6262 | 17.5 |  | 103.2104 | 21.571 | 214.7009 | 161.0393 | 8.1293 | 254.0415
2010-08-12 | 14:57:24 | 8.362 | 310.364 | 30.31 | 6262 | 17.5 |  | 103.9323 | 21.7218 | 216.2026 | 162.1657 | 8.1862 | 255.8183
2010-08-12 | 14:57:34 | 8.529 | 307.672 | 30.41 | 6266 | 17.5 |  | 103.0308 | 21.5334 | 214.3273 | 160.7591 | 8.1152 | 253.5995
2010-08-12 | 14:57:44 | 8.696 | 308.349 | 30.34 | 6246 | 17.6 |  | 103.4704 | 21.6253 | 215.2143 | 161.4244 | 8.1344 | 254.1994
2010-08-12 | 14:57:53 | 8.846 | 306.741 | 30.4 | 6253 | 17.6 |  | 102.931 | 21.5126 | 214.0924 | 160.5829 | 8.092 | 252.8742
2010-08-12 | 14:58:03 | 9.013 | 307.276 | 30.38 | 6264 | 17.6 |  | 103.1105 | 21.5501 | 214.4656 | 160.8629 | 8.1061 | 253.3151
2010-08-12 | 14:58:13 | 9.18 | 305.941 | 30.43 | 6260 | 17.6 |  | 102.6625 | 21.4565 | 213.5339 | 160.1641 | 8.0709 | 252.2146
2010-08-12 | 14:58:23 | 9.347 | 309.695 | 30.29 | 6261 | 17.6 |  | 103.9223 | 21.7198 | 216.1543 | 162.1295 | 8.1699 | 255.3096
2010-08-12 | 14:58:33 | 9.514 | 309.965 | 30.28 | 6260 | 17.6 |  | 104.013 | 21.7387 | 216.3428 | 162.2709 | 8.177 | 255.5324
2010-08-12 | 14:58:43 | 9.68 | 309.695 | 30.29 | 6252 | 17.6 |  | 103.9223 | 21.7198 | 216.1543 | 162.1295 | 8.1699 | 255.3096
2010-08-12 | 14:58:53 | 9.848 | 306.474 | 30.41 | 6255 | 17.6 |  | 102.8414 | 21.4939 | 213.9061 | 160.4432 | 8.0849 | 252.6542
2010-08-12 | 14:59:03 | 10.014 | 309.695 | 30.29 | 6245 | 17.6 |  | 103.9223 | 21.7198 | 216.1543 | 162.1295 | 8.1699 | 255.3096
2010-08-12 | 14:59:13 | 10.181 | 308.617 | 30.33 | 6249 | 17.6 |  | 103.5606 | 21.6442 | 215.4019 | 161.5652 | 8.1415 | 254.421
2010-08-12 | 14:59:23 | 10.348 | 306.207 | 30.42 | 6254 | 17.6 |  | 102.7519 | 21.4752 | 213.7199 | 160.3036 | 8.0779 | 252.4343
2010-08-12 | 14:59:33 | 10.515 | 308.617 | 30.33 | 6270 | 17.6 |  | 103.5606 | 21.6442 | 215.4019 | 161.5652 | 8.1415 | 254.421
2010-08-12 | 14:59:43 | 10.682 | 307.544 | 30.37 | 6253 | 17.6 |  | 103.2003 | 21.5689 | 214.6525 | 161.0031 | 8.1131 | 253.5358
2010-08-12 | 14:59:53 | 10.849 | 308.349 | 30.34 | 6264 | 17.6 |  | 103.4704 | 21.6253 | 215.2143 | 161.4244 | 8.1344 | 254.1994
2010-08-12 | 15:00:03 | 11.016 | 307.276 | 30.38 | 6260 | 17.6 |  | 103.1105 | 21.5501 | 214.4656 | 160.8629 | 8.1061 | 253.3151
2010-08-12 | 15:00:13 | 11.183 | 308.349 | 30.34 | 6254 | 17.6 |  | 103.4704 | 21.6253 | 215.2143 | 161.4244 | 8.1344 | 254.1994
2010-08-12 | 15:00:23 | 11.35 | 305.675 | 30.44 | 6253 | 17.6 |  | 102.5732 | 21.4378 | 213.3481 | 160.0247 | 8.0638 | 251.9952
2010-08-12 | 15:00:33 | 11.517 | 308.349 | 30.34 | 6263 | 17.6 |  | 103.4704 | 21.6253 | 215.2143 | 161.4244 | 8.1344 | 254.1994
2010-08-12 | 15:00:43 | 11.683 | 306.207 | 30.42 | 6258 | 17.6 |  | 102.7519 | 21.4752 | 213.7199 | 160.3036 | 8.0779 | 252.4343
2010-08-12 | 15:00:53 | 11.85 | 307.812 | 30.36 | 6239 | 17.6 |  | 103.2902 | 21.5877 | 214.8396 | 161.1434 | 8.1202 | 253.7568
2010-08-12 | 15:01:03 | 12.017 | 307.544 | 30.37 | 6262 | 17.6 |  | 103.2003 | 21.5689 | 214.6525 | 161.0031 | 8.1131 | 253.5358
2010-08-12 | 15:01:13 | 12.184 | 308.617 | 30.33 | 6265 | 17.6 |  | 103.5606 | 21.6442 | 215.4019 | 161.5652 | 8.1415 | 254.421
2010-08-12 | 15:01:23 | 12.351 | 307.008 | 30.39 | 6256 | 17.6 |  | 103.0207 | 21.5313 | 214.2789 | 160.7228 | 8.099 | 253.0946
2010-08-12 | 15:01:33 | 12.518 | 306.741 | 30.4 | 6258 | 17.6 |  | 102.931 | 21.5126 | 214.0924 | 160.5829 | 8.092 | 252.8742
2010-08-12 | 15:01:43 | 12.685 | 308.617 | 30.33 | 6251 | 17.6 |  | 103.5606 | 21.6442 | 215.4019 | 161.5652 | 8.1415 | 254.421
2010-08-12 | 15:01:53 | 12.852 | 307.404 | 30.42 | 6255 | 17.5 |  | 102.9411 | 21.5147 | 214.1408 | 160.6193 | 8.1081 | 253.3788
2010-08-12 | 15:02:03 | 13.018 | 307.94 | 30.4 | 6260 | 17.5 |  | 103.1205 | 21.5522 | 214.514 | 160.8992 | 8.1223 | 253.8204
2010-08-12 | 15:02:13 | 13.185 | 308.208 | 30.39 | 6245 | 17.5 |  | 103.2104 | 21.571 | 214.7009 | 161.0393 | 8.1293 | 254.0415
2010-08-12 | 15:02:23 | 13.352 | 307.672 | 30.41 | 6261 | 17.5 |  | 103.0308 | 21.5334 | 214.3273 | 160.7591 | 8.1152 | 253.5995
2010-08-12 | 15:02:34 | 13.519 | 306.87 | 30.44 | 6265 | 17.5 |  | 102.7621 | 21.4773 | 213.7683 | 160.3399 | 8.094 | 252.9381
2010-08-12 | 15:02:44 | 13.686 | 306.87 | 30.44 | 6269 | 17.5 |  | 102.7621 | 21.4773 | 213.7683 | 160.3399 | 8.094 | 252.9381
2010-08-12 | 15:02:54 | 13.853 | 310.094 | 30.32 | 6247 | 17.5 |  | 103.8417 | 21.7029 | 216.0142 | 162.0244 | 8.1791 | 255.5955
2010-08-12 | 15:03:04 | 14.02 | 307.672 | 30.41 | 6246 | 17.5 |  | 103.0308 | 21.5334 | 214.3273 | 160.7591 | 8.1152 | 253.5995
2010-08-12 | 15:03:14 | 14.187 | 310.094 | 30.32 | 6253 | 17.5 |  | 103.8417 | 21.7029 | 216.0142 | 162.0244 | 8.1791 | 255.5955
2010-08-12 | 15:03:24 | 14.354 | 308.746 | 30.37 | 6254 | 17.5 |  | 103.3903 | 21.6086 | 215.0752 | 161.3201 | 8.1435 | 254.4844
2010-08-12 | 15:03:34 | 14.521 | 307.672 | 30.41 | 6270 | 17.5 |  | 103.0308 | 21.5334 | 214.3273 | 160.7591 | 8.1152 | 253.5995
2010-08-12 | 15:03:44 | 14.688 | 307.94 | 30.4 | 6267 | 17.5 |  | 103.1205 | 21.5522 | 214.514 | 160.8992 | 8.1223 | 253.8204
2010-08-12 | 15:03:54 | 14.855 | 311.448 | 30.27 | 6261 | 17.5 |  | 104.2953 | 21.7977 | 216.9579 | 162.7322 | 8.2148 | 256.712
2010-08-12 | 15:04:04 | 15.021 | 309.824 | 30.33 | 6254 | 17.5 |  | 103.7512 | 21.684 | 215.826 | 161.8833 | 8.1719 | 255.3728
2010-08-12 | 15:04:14 | 15.188 | 307.404 | 30.42 | 6266 | 17.5 |  | 102.9411 | 21.5147 | 214.1408 | 160.6193 | 8.1081 | 253.3788
2010-08-12 | 15:04:24 | 15.355 | 308.746 | 30.37 | 6258 | 17.5 |  | 103.3903 | 21.6086 | 215.0752 | 161.3201 | 8.1435 | 254.4844
2010-08-12 | 15:04:34 | 15.522 | 308.477 | 30.38 | 6262 | 17.5 |  | 103.3003 | 21.5898 | 214.8879 | 161.1797 | 8.1364 | 254.2628
2010-08-12 | 15:04:44 | 15.689 | 305.272 | 30.5 | 6255 | 17.5 |  | 102.2271 | 21.3655 | 212.6554 | 159.5051 | 8.0519 | 251.6212
2010-08-12 | 15:04:54 | 15.856 | 307.404 | 30.42 | 6259 | 17.5 |  | 102.9411 | 21.5147 | 214.1408 | 160.6193 | 8.1081 | 253.3788
2010-08-12 | 15:05:04 | 16.023 | 308.08 | 30.35 | 6261 | 17.6 |  | 103.3803 | 21.6065 | 215.0268 | 161.2838 | 8.1273 | 253.978
2010-08-12 | 15:05:14 | 16.19 | 308.08 | 30.35 | 6252 | 17.6 |  | 103.3803 | 21.6065 | 215.0268 | 161.2838 | 8.1273 | 253.978
2010-08-12 | 15:05:24 | 16.357 | 308.349 | 30.34 | 6255 | 17.6 |  | 103.4704 | 21.6253 | 215.2143 | 161.4244 | 8.1344 | 254.1994
2010-08-12 | 15:05:34 | 16.524 | 308.08 | 30.35 | 6272 | 17.6 |  | 103.3803 | 21.6065 | 215.0268 | 161.2838 | 8.1273 | 253.978
2010-08-12 | 15:05:44 | 16.691 | 308.349 | 30.34 | 6264 | 17.6 |  | 103.4704 | 21.6253 | 215.2143 | 161.4244 | 8.1344 | 254.1994
2010-08-12 | 15:05:54 | 16.857 | 305.675 | 30.44 | 6263 | 17.6 |  | 102.5732 | 21.4378 | 213.3481 | 160.0247 | 8.0638 | 251.9952
2010-08-12 | 15:06:04 | 17.024 | 305.941 | 30.43 | 6254 | 17.6 |  | 102.6625 | 21.4565 | 213.5339 | 160.1641 | 8.0709 | 252.2146
2010-08-12 | 15:06:14 | 17.191 | 308.349 | 30.34 | 6253 | 17.6 |  | 103.4704 | 21.6253 | 215.2143 | 161.4244 | 8.1344 | 254.1994
2010-08-12 | 15:06:24 | 17.358 | 309.156 | 30.31 | 6259 | 17.6 |  | 103.7413 | 21.6819 | 215.7777 | 161.847 | 8.1557 | 254.8649
2010-08-12 | 15:06:34 | 17.525 | 310.236 | 30.27 | 6245 | 17.6 |  | 104.1037 | 21.7577 | 216.5316 | 162.4125 | 8.1842 | 255.7553
2010-08-12 | 15:06:44 | 17.692 | 305.409 | 30.45 | 6259 | 17.6 |  | 102.484 | 21.4191 | 213.1625 | 159.8855 | 8.0568 | 251.776
2010-08-12 | 15:06:54 | 17.859 | 307.008 | 30.39 | 6250 | 17.6 |  | 103.0207 | 21.5313 | 214.2789 | 160.7228 | 8.099 | 253.0946
2010-08-12 | 15:07:04 | 18.026 | 308.08 | 30.35 | 6246 | 17.6 |  | 103.3803 | 21.6065 | 215.0268 | 161.2838 | 8.1273 | 253.978
2010-08-12 | 15:07:14 | 18.193 | 308.08 | 30.35 | 6247 | 17.6 |  | 103.3803 | 21.6065 | 215.0268 | 161.2838 | 8.1273 | 253.978
2010-08-12 | 15:07:24 | 18.36 | 308.349 | 30.34 | 6251 | 17.6 |  | 103.4704 | 21.6253 | 215.2143 | 161.4244 | 8.1344 | 254.1994
2010-08-12 | 15:07:34 | 18.527 | 305.675 | 30.44 | 6256 | 17.6 |  | 102.5732 | 21.4378 | 213.3481 | 160.0247 | 8.0638 | 251.9952
2010-08-12 | 15:07:44 | 18.694 | 307.15 | 30.34 | 6251 | 17.7 |  | 103.2804 | 21.5856 | 214.7915 | 161.1073 | 8.1041 | 253.2528
2010-08-12 | 15:07:54 | 18.86 | 306.348 | 30.37 | 6238 | 17.7 |  | 103.0108 | 21.5293 | 214.2308 | 160.6868 | 8.0829 | 252.5917
2010-08-12 | 15:08:04 | 19.027 | 303.428 | 30.48 | 6252 | 17.7 |  | 102.029 | 21.3241 | 212.189 | 159.1553 | 8.0059 | 250.1843
2010-08-12 | 15:08:14 | 19.194 | 307.685 | 30.32 | 6247 | 17.7 |  | 103.4606 | 21.6233 | 215.1662 | 161.3884 | 8.1182 | 253.6946
2010-08-12 | 15:08:24 | 19.361 | 306.882 | 30.35 | 6254 | 17.7 |  | 103.1904 | 21.5668 | 214.6044 | 160.967 | 8.097 | 253.0322
2010-08-12 | 15:08:34 | 19.528 | 307.954 | 30.31 | 6245 | 17.7 |  | 103.5508 | 21.6421 | 215.3539 | 161.5291 | 8.1253 | 253.9159
2010-08-12 | 15:08:44 | 19.695 | 308.222 | 30.3 | 6245 | 17.7 |  | 103.6411 | 21.661 | 215.5417 | 161.67 | 8.1324 | 254.1373
2010-08-12 | 15:08:54 | 19.862 | 308.76 | 30.28 | 6239 | 17.7 |  | 103.822 | 21.6988 | 215.9179 | 161.9522 | 8.1466 | 254.5809
2010-08-12 | 15:09:04 | 20.029 | 310.381 | 30.22 | 6227 | 17.7 |  | 104.3669 | 21.8127 | 217.0511 | 162.8021 | 8.1893 | 255.917
2010-08-12 | 15:09:14 | 20.196 | 306.615 | 30.36 | 6232 | 17.7 |  | 103.1006 | 21.548 | 214.4175 | 160.8268 | 8.09 | 252.8118
2010-08-12 | 15:09:24 | 20.362 | 311.195 | 30.19 | 6231 | 17.7 |  | 104.6405 | 21.8699 | 217.6202 | 163.229 | 8.2108 | 256.588
2010-08-12 | 15:09:33 | 20.513 | 310.652 | 30.21 | 6240 | 17.7 |  | 104.458 | 21.8317 | 217.2406 | 162.9443 | 8.1965 | 256.1404
2010-08-12 | 15:09:43 | 20.68 | 309.425 | 30.3 | 6220 | 17.6 |  | 103.8318 | 21.7008 | 215.9659 | 161.9882 | 8.1628 | 255.0872
2010-08-12 | 15:09:53 | 20.846 | 307.812 | 30.36 | 6222 | 17.6 |  | 103.2902 | 21.5877 | 214.8396 | 161.1434 | 8.1202 | 253.7568
2010-08-12 | 15:10:03 | 21.013 | 309.156 | 30.31 | 6229 | 17.6 |  | 103.7413 | 21.6819 | 215.7777 | 161.847 | 8.1557 | 254.8649
2010-08-12 | 15:10:13 | 21.18 | 309.965 | 30.28 | 6225 | 17.6 |  | 104.013 | 21.7387 | 216.3428 | 162.2709 | 8.177 | 255.5324
2010-08-12 | 15:10:23 | 21.347 | 309.965 | 30.28 | 6218 | 17.6 |  | 104.013 | 21.7387 | 216.3428 | 162.2709 | 8.177 | 255.5324
2010-08-12 | 15:10:33 | 21.514 | 311.049 | 30.24 | 6220 | 17.6 |  | 104.3765 | 21.8147 | 217.099 | 162.8381 | 8.2056 | 256.4255
2010-08-12 | 15:10:43 | 21.681 | 311.864 | 30.21 | 6228 | 17.6 |  | 104.6501 | 21.8719 | 217.668 | 163.2649 | 8.2271 | 257.0976
2010-08-12 | 15:10:53 | 21.848 | 310.236 | 30.27 | 6237 | 17.6 |  | 104.1037 | 21.7577 | 216.5316 | 162.4125 | 8.1842 | 255.7553
2010-08-12 | 15:11:03 | 22.015 | 309.695 | 30.29 | 6217 | 17.6 |  | 103.9223 | 21.7198 | 216.1543 | 162.1295 | 8.1699 | 255.3096
2010-08-12 | 15:11:13 | 22.182 | 309.695 | 30.29 | 6210 | 17.6 |  | 103.9223 | 21.7198 | 216.1543 | 162.1295 | 8.1699 | 255.3096
2010-08-12 | 15:11:23 | 22.348 | 310.094 | 30.32 | 6227 | 17.5 |  | 103.8417 | 21.7029 | 216.0142 | 162.0244 | 8.1791 | 255.5955
2010-08-12 | 15:11:33 | 22.516 | 309.156 | 30.31 | 6227 | 17.6 |  | 103.7413 | 21.6819 | 215.7777 | 161.847 | 8.1557 | 254.8649
2010-08-12 | 15:11:43 | 22.682 | 311.864 | 30.21 | 6214 | 17.6 |  | 104.6501 | 21.8719 | 217.668 | 163.2649 | 8.2271 | 257.0976
2010-08-12 | 15:11:53 | 22.849 | 312.955 | 30.17 | 6229 | 17.6 |  | 105.0162 | 21.9484 | 218.4294 | 163.836 | 8.2559 | 257.9969
2010-08-12 | 15:12:03 | 23.016 | 312.136 | 30.2 | 6211 | 17.6 |  | 104.7415 | 21.891 | 217.8581 | 163.4075 | 8.2343 | 257.3221
2010-08-12 | 15:12:13 | 23.183 | 309.965 | 30.28 | 6226 | 17.6 |  | 104.013 | 21.7387 | 216.3428 | 162.2709 | 8.177 | 255.5324
2010-08-12 | 15:12:23 | 23.35 | 310.507 | 30.26 | 6212 | 17.6 |  | 104.1946 | 21.7767 | 216.7205 | 162.5542 | 8.1913 | 255.9785
2010-08-12 | 15:12:33 | 23.517 | 312.682 | 30.18 | 6202 | 17.6 |  | 104.9245 | 21.9292 | 218.2388 | 163.693 | 8.2487 | 257.7718
2010-08-12 | 15:12:43 | 23.684 | 312.136 | 30.2 | 6201 | 17.6 |  | 104.7415 | 21.891 | 217.8581 | 163.4075 | 8.2343 | 257.3221
2010-08-12 | 15:12:53 | 23.851 | 310.507 | 30.26 | 6213 | 17.6 |  | 104.1946 | 21.7767 | 216.7205 | 162.5542 | 8.1913 | 255.9785
2010-08-12 | 15:13:03 | 24.018 | 313.502 | 30.15 | 6199 | 17.6 |  | 105.1997 | 21.9867 | 218.8113 | 164.1224 | 8.2703 | 258.4479
2010-08-12 | 15:13:13 | 24.185 | 314.05 | 30.13 | 6206 | 17.6 |  | 105.3837 | 22.0252 | 219.1939 | 164.4094 | 8.2848 | 258.8998
2010-08-12 | 15:13:23 | 24.352 | 312.955 | 30.17 | 6196 | 17.6 |  | 105.0162 | 21.9484 | 218.4294 | 163.836 | 8.2559 | 257.9969
2010-08-12 | 15:13:33 | 24.518 | 311.049 | 30.24 | 6204 | 17.6 |  | 104.3765 | 21.8147 | 217.099 | 162.8381 | 8.2056 | 256.4255
2010-08-12 | 15:13:43 | 24.685 | 313.228 | 30.16 | 6199 | 17.6 |  | 105.1079 | 21.9676 | 218.6203 | 163.9791 | 8.2631 | 258.2223
2010-08-12 | 15:13:54 | 24.852 | 313.502 | 30.15 | 6202 | 17.6 |  | 105.1997 | 21.9867 | 218.8113 | 164.1224 | 8.2703 | 258.4479
2010-08-12 | 15:14:03 | 25.019 | 313.776 | 30.14 | 6193 | 17.6 |  | 105.2917 | 22.006 | 219.0025 | 164.2658 | 8.2776 | 258.6738
2010-08-12 | 15:14:14 | 25.186 | 312.136 | 30.2 | 6182 | 17.6 |  | 104.7415 | 21.891 | 217.8581 | 163.4075 | 8.2343 | 257.3221
2010-08-12 | 15:14:24 | 25.353 | 313.228 | 30.16 | 6190 | 17.6 |  | 105.1079 | 21.9676 | 218.6203 | 163.9791 | 8.2631 | 258.2223
2010-08-12 | 15:14:34 | 25.52 | 315.977 | 30.06 | 6175 | 17.6 |  | 106.0304 | 22.1604 | 220.539 | 165.4183 | 8.3356 | 260.4886
2010-08-12 | 15:14:44 | 25.687 | 312.682 | 30.18 | 6192 | 17.6 |  | 104.9245 | 21.9292 | 218.2388 | 163.693 | 8.2487 | 257.7718
2010-08-12 | 15:14:54 | 25.854 | 312.829 | 30.13 | 6195 | 17.7 |  | 105.1903 | 21.9848 | 218.7635 | 164.0866 | 8.254 | 257.936
2010-08-12 | 15:15:04 | 26.021 | 313.651 | 30.1 | 6190 | 17.7 |  | 105.4664 | 22.0425 | 219.3377 | 164.5173 | 8.2756 | 258.6131
2010-08-12 | 15:15:14 | 26.188 | 315.3 | 30.04 | 6176 | 17.7 |  | 106.0211 | 22.1584 | 220.4914 | 165.3826 | 8.3191 | 259.9733
2010-08-12 | 15:15:24 | 26.355 | 316.406 | 30 | 6168 | 17.7 |  | 106.3928 | 22.2361 | 221.2643 | 165.9623 | 8.3483 | 260.8846
2010-08-12 | 15:15:34 | 26.521 | 315.3 | 30.04 | 6168 | 17.7 |  | 106.0211 | 22.1584 | 220.4914 | 165.3826 | 8.3191 | 259.9733
2010-08-12 | 15:15:44 | 26.688 | 317.515 | 29.96 | 6167 | 17.7 |  | 106.7659 | 22.3141 | 222.0404 | 166.5444 | 8.3776 | 261.7996
2010-08-12 | 15:15:54 | 26.855 | 316.129 | 30.01 | 6163 | 17.7 |  | 106.2997 | 22.2166 | 221.0708 | 165.8172 | 8.341 | 260.6565
2010-08-12 | 15:16:04 | 27.022 | 315.852 | 30.02 | 6156 | 17.7 |  | 106.2068 | 22.1972 | 220.8775 | 165.6722 | 8.3337 | 260.4285
2010-08-12 | 15:16:14 | 27.189 | 313.925 | 30.09 | 6167 | 17.7 |  | 105.5586 | 22.0618 | 219.5295 | 164.6611 | 8.2829 | 258.8392
2010-08-12 | 15:16:24 | 27.356 | 314.474 | 30.07 | 6157 | 17.7 |  | 105.7434 | 22.1004 | 219.9137 | 164.9493 | 8.2973 | 259.2922
2010-08-12 | 15:16:34 | 27.523 | 318.63 | 29.92 | 6147 | 17.7 |  | 107.1406 | 22.3924 | 222.8195 | 167.1288 | 8.407 | 262.7183
2010-08-12 | 15:16:44 | 27.69 | 316.96 | 29.98 | 6143 | 17.7 |  | 106.5792 | 22.275 | 221.652 | 166.2531 | 8.3629 | 261.3417
2010-08-12 | 15:16:54 | 27.856 | 317.238 | 29.97 | 6137 | 17.7 |  | 106.6725 | 22.2946 | 221.8461 | 166.3987 | 8.3703 | 261.5706
2010-08-12 | 15:17:04 | 28.023 | 317.238 | 29.97 | 6143 | 17.7 |  | 106.6725 | 22.2946 | 221.8461 | 166.3987 | 8.3703 | 261.5706
2010-08-12 | 15:17:14 | 28.19 | 318.909 | 29.91 | 6149 | 17.7 |  | 107.2345 | 22.412 | 223.0148 | 167.2753 | 8.4144 | 262.9486
2010-08-12 | 15:17:24 | 28.357 | 315.3 | 30.04 | 6131 | 17.7 |  | 106.0211 | 22.1584 | 220.4914 | 165.3826 | 8.3191 | 259.9733
2010-08-12 | 15:17:34 | 28.524 | 318.197 | 29.98 | 6137 | 17.6 |  | 106.775 | 22.316 | 222.0878 | 166.58 | 8.3942 | 262.318
2010-08-12 | 15:17:44 | 28.691 | 319.873 | 29.92 | 6135 | 17.6 |  | 107.3375 | 22.4335 | 223.2576 | 167.4574 | 8.4384 | 263.6997
2010-08-12 | 15:17:54 | 28.858 | 319.873 | 29.92 | 6137 | 17.6 |  | 107.3375 | 22.4335 | 223.2576 | 167.4574 | 8.4384 | 263.6997
2010-08-12 | 15:18:04 | 29.025 | 321.277 | 29.87 | 6135 | 17.6 |  | 107.8087 | 22.532 | 224.2378 | 168.1927 | 8.4754 | 264.8575
2010-08-12 | 15:18:14 | 29.192 | 319.313 | 29.94 | 6129 | 17.6 |  | 107.1496 | 22.3943 | 222.8669 | 167.1644 | 8.4236 | 263.2382
2010-08-12 | 15:18:24 | 29.359 | 318.475 | 29.97 | 6134 | 17.6 |  | 106.8685 | 22.3355 | 222.2823 | 166.7259 | 8.4015 | 262.5477
2010-08-12 | 15:18:34 | 29.526 | 321.559 | 29.86 | 6118 | 17.6 |  | 107.9033 | 22.5518 | 224.4345 | 168.3402 | 8.4829 | 265.0898
2010-08-12 | 15:18:44 | 29.693 | 319.593 | 29.93 | 6115 | 17.6 |  | 107.2435 | 22.4139 | 223.0622 | 167.3108 | 8.431 | 263.4688
2010-08-12 | 15:18:54 | 29.859 | 319.873 | 29.92 | 6121 | 17.6 |  | 107.3375 | 22.4335 | 223.2576 | 167.4574 | 8.4384 | 263.6997
2010-08-12 | 15:19:04 | 30.026 | 319.033 | 29.95 | 6118 | 17.6 |  | 107.0558 | 22.3747 | 222.6718 | 167.0181 | 8.4163 | 263.0078
2010-08-12 | 15:19:14 | 30.193 | 320.434 | 29.9 | 6119 | 17.6 |  | 107.5257 | 22.4729 | 223.6491 | 167.7511 | 8.4532 | 264.1621
2010-08-12 | 15:19:24 | 30.36 | 320.434 | 29.9 | 6101 | 17.6 |  | 107.5257 | 22.4729 | 223.6491 | 167.7511 | 8.4532 | 264.1621
2010-08-12 | 15:19:34 | 30.527 | 320.996 | 29.88 | 6119 | 17.6 |  | 107.7143 | 22.5123 | 224.0414 | 168.0453 | 8.468 | 264.6255
2010-08-12 | 15:19:44 | 30.694 | 318.754 | 29.96 | 6093 | 17.6 |  | 106.9621 | 22.3551 | 222.477 | 166.8719 | 8.4089 | 262.7777
2010-08-12 | 15:19:54 | 30.861 | 321.559 | 29.86 | 6105 | 17.6 |  | 107.9033 | 22.5518 | 224.4345 | 168.3402 | 8.4829 | 265.0898
2010-08-12 | 15:20:04 | 31.028 | 320.996 | 29.88 | 6097 | 17.6 |  | 107.7143 | 22.5123 | 224.0414 | 168.0453 | 8.468 | 264.6255
2010-08-12 | 15:20:14 | 31.195 | 322.972 | 29.81 | 6091 | 17.6 |  | 108.3774 | 22.6509 | 225.4206 | 169.0798 | 8.5201 | 266.2545
2010-08-12 | 15:20:24 | 31.362 | 323.539 | 29.79 | 6106 | 17.6 |  | 108.5677 | 22.6907 | 225.8165 | 169.3768 | 8.5351 | 266.7221
2010-08-12 | 15:20:34 | 31.529 | 325.533 | 29.72 | 6099 | 17.6 |  | 109.2368 | 22.8305 | 227.2083 | 170.4207 | 8.5877 | 268.366
2010-08-12 | 15:20:44 | 31.695 | 322.123 | 29.84 | 6092 | 17.6 |  | 108.0926 | 22.5914 | 224.8283 | 168.6356 | 8.4978 | 265.555
2010-08-12 | 15:20:54 | 31.862 | 325.247 | 29.73 | 6086 | 17.6 |  | 109.141 | 22.8105 | 227.0088 | 170.2711 | 8.5802 | 268.1305
2010-08-12 | 15:21:03 | 32.013 | 322.972 | 29.81 | 6094 | 17.6 |  | 108.3774 | 22.6509 | 225.4206 | 169.0798 | 8.5201 | 266.2545
2010-08-12 | 15:21:13 | 32.179 | 324.392 | 29.76 | 6084 | 17.6
2010-08-12 | 15:21:23 | 32.346 | 324.392 | 29.76 | 6079 | 17.6
2010-08-12 | 15:21:33 | 32.513 | 323.539 | 29.79 | 6083 | 17.6
2010-08-12 | 15:21:43 | 32.68 | 322.972 | 29.81 | 6071 | 17.6
2010-08-12 | 15:21:53 | 32.847 | 323.699 | 29.74 | 6077 | 17.7
2010-08-12 | 15:22:03 | 33.014 | 322.282 | 29.79 | 6071 | 17.7
2010-08-12 | 15:22:13 | 33.181 | 321.999 | 29.8 | 6063 | 17.7
2010-08-12 | 15:22:23 | 33.348 | 323.699 | 29.74 | 6062 | 17.7
2010-08-12 | 15:22:33 | 33.515 | 325.696 | 29.67 | 6061 | 17.7
2010-08-12 | 15:22:43 | 33.682 | 326.269 | 29.65 | 6054 | 17.7
2010-08-12 | 15:22:53 | 33.849 | 326.269 | 29.65 | 6069 | 17.7
2010-08-12 | 15:23:03 | 34.016 | 323.415 | 29.75 | 6050 | 17.7
2010-08-12 | 15:23:13 | 34.183 | 325.696 | 29.67 | 6043 | 17.7
2010-08-12 | 15:23:23 | 34.349 | 325.982 | 29.66 | 6038 | 17.7
2010-08-12 | 15:23:33 | 34.516 | 325.696 | 29.67 | 6049 | 17.7
2010-08-12 | 15:23:43 | 34.683 | 326.843 | 29.63 | 6036 | 17.7
2010-08-12 | 15:23:53 | 34.85 | 328.862 | 29.56 | 6049 | 17.7
2010-08-12 | 15:24:03 | 35.017 | 331.601 | 29.51 | 6017 | 17.6
2010-08-12 | 15:24:13 | 35.184 | 330.435 | 29.55 | 6023 | 17.6
2010-08-12 | 15:24:23 | 35.351 | 328.695 | 29.61 | 6025 | 17.6
2010-08-12 | 15:24:33 | 35.518 | 330.435 | 29.55 | 6018 | 17.6
2010-08-12 | 15:24:43 | 35.685 | 330.435 | 29.55 | 6018 | 17.6
2010-08-12 | 15:24:53 | 35.851 | 330.144 | 29.56 | 6018 | 17.6
2010-08-12 | 15:25:03 | 36.018 | 334.241 | 29.42 | 6015 | 17.6
2010-08-12 | 15:25:13 | 36.185 | 331.017 | 29.53 | 6024 | 17.6
2010-08-12 | 15:25:23 | 36.352 | 331.725 | 29.55 | 6015 | 17.5
2010-08-12 | 15:25:33 | 36.519 | 336.138 | 29.4 | 6014 | 17.5
2010-08-12 | 15:25:44 | 36.686 | 334.955 | 29.44 | 6002 | 17.5
2010-08-12 | 15:25:54 | 36.853 | 334.07 | 29.47 | 6013 | 17.5
2010-08-12 | 15:26:04 | 37.02 | 334.955 | 29.44 | 6000 | 17.5
2010-08-12 | 15:26:14 | 37.187 | 334.07 | 29.47 | 6010 | 17.5
2010-08-12 | 15:26:24 | 37.353 | 332.602 | 29.52 | 6005 | 17.5
2010-08-12 | 15:26:34 | 37.52 | 331.893 | 29.5 | 5978 | 17.6
2010-08-12 | 15:26:44 | 37.687 | 333.358 | 29.45 | 5986 | 17.6
2010-08-12 | 15:26:54 | 37.854 | 335.718 | 29.37 | 5998 | 17.6