## convers.+compens.
file name: PSt1_Eq4_Sal |  |  | 2008-03-29
author: Dr. Christian Huber
programm for temperature compensated oxygen calculation of membrane PSt1
Fit of the calibration curve is optimized between 0 and 250 % air-saturation (0 and 388 Torr)
enter your values here
input of F0,T0 | 58.2 | phase angle of cal 0
input of F100,T100 | 28.64 | phase angle of cal 100 (water-vapor saturated air)
input of Fm,Tm | 15 | measured phase angle
input of T0 | 17.6 | temperature of cal 0
input of T100 | 17.7 | temperatur of cal100
input of Tm | 17.4 | temperature at measurement
air pressure patm | 1013 | air pressure
Salinity | 32.2 | Salinity
Results of calculation + compensation
[%] air saturation | 417.4737 | %
[%] O2 | 87.252 | %
pO2 [hPa] | 868.5499 | hPa
pO2 [Torr] | 651.4678 | Torr
cO2 [mg/L] | 32.9446 | mg/L
cO2 [ppm] | 32.9446 | ppm
cO2 [µmol/L] | 1029.5196 | µmol/L
Internal calculated parameter (DO NOT CHANGE !!)
f1 | 0.801
DF/K | -0.08
DKSV/K | 0.000383
m | 22.9
tan(F0,T100) | 1.6123
tan(F0,Tm) | 1.6138
tan(F100,T100) | 0.54612
tan(Fm,Tm) | 0.26795
KSV,T100 | 0.037719
KSV,Tm | 0.037604
A | 147.9112
B | 11.953
C | -0.66128
a | 1.0252e-05
b | -0.0022824
c | -0.83397
Chlorinity | 17.8227

## Recalculate Fibox values
file name: PSt1_Eq4_Sal |  |  |  | 2008-03-29
author: Dr. Christian Huber
programm for temperature compensated oxygen calculation of membrane PSt1
Fit of the calibration curve is optimized between 0 and 250 % air-saturation (0 and 388 Torr)
Insert your calibration values
cal0 | 58.2 | ° | T0 | 17.6 | °C
cal100 | 28.64 | ° | T100 | 17.7 | °C
air pressure | 1013 | mbar
Salinity | 32.2 | °/°°
Internal calculated parameter (DO NOT CHANGE !!)
f1 | 0.801 | tan(F0,T100) | 1.6123 | A | 147.9112 | KSV,T100 | 0.037719 | a | 2.0948e-05
DF/K | -0.08 | tan(F0,Tm) | 1.6138 | B | 11.953 | KSV,Tm | 0.037604 | b | 0.0045154 |  |  |  |  | INSERT YOUR RESPECTIVE DATA
DKSV/K | 0.000383 | tan(F100,T100) | 0.54612 | C | -0.66128 |  |  | c | -0.66076 |  |  |  |  | chamber volume [L] | FW alga [g] | mg Chl a Thallus -1
m | 22.9 | tan(Fm,Tm) | 0.54748 |  |  |  |  | Chlorinity | 17.8227 |  |  |  |  | 0.026 | 0.058 | 0.023876
Insert the data you get from the Ascii File recorded with the OxyView TX3-V5.31.exe software | [7 empty] | The oxygen contents are calculated automatically from the raw data
date | time/hh:mm:ss | logtime/min | oxygen/% airsatur. | phase/° | amp | temp/°C |  | % air-sat. | % oxygen | pO2 (hPa) | pO2 (Torr) | cO2 [mg/L] | cO2 [µmol/L] | cO2 [µmol/L 10 min] | cO2 [µmol/L h] | cO2 [µmol/ h] | cO2 [µmol/ h gFW] | cO2 [µmol/ h mg Chl a]
2010-08-12 | 18:02:21 | 0 | 300.354 | 28.7 | 2927 | 17.4 |  | 99.9698 | 20.8937 | 207.9862 | 156.0029 | 7.889 | 246.5324 | 11.97 | 71.82 | 1.8673 | 32.1952 | 78.2091
2010-08-12 | 18:02:33 | 0.201 | 301.701 | 28.65 | 2923 | 17.4 |  | 100.4184 | 20.9874 | 208.9194 | 156.7029 | 7.9244 | 247.6386 | change data input according to column N
2010-08-12 | 18:02:43 | 0.368 | 301.161 | 28.67 | 2911 | 17.4 |  | 100.2387 | 20.9499 | 208.5455 | 156.4224 | 7.9103 | 247.1954
2010-08-12 | 18:02:53 | 0.535 | 304.965 | 28.53 | 2915 | 17.4 |  | 101.5045 | 21.2145 | 211.1792 | 158.3979 | 8.0102 | 250.3172
2010-08-12 | 18:03:03 | 0.702 | 299.281 | 28.74 | 2913 | 17.4 |  | 99.6126 | 20.819 | 207.2431 | 155.4455 | 7.8609 | 245.6516
2010-08-12 | 18:03:13 | 0.869 | 304.691 | 28.54 | 2913 | 17.4 |  | 101.4135 | 21.1954 | 210.9898 | 158.2558 | 8.003 | 250.0927
2010-08-12 | 18:03:23 | 1.036 | 301.701 | 28.65 | 2915 | 17.4 |  | 100.4184 | 20.9874 | 208.9194 | 156.7029 | 7.9244 | 247.6386
2010-08-12 | 18:03:33 | 1.203 | 303.328 | 28.59 | 2907 | 17.4 |  | 100.9597 | 21.1006 | 210.0458 | 157.5477 | 7.9672 | 248.9737
2010-08-12 | 18:03:43 | 1.37 | 302.513 | 28.62 | 2913 | 17.4 |  | 100.6886 | 21.0439 | 209.4817 | 157.1246 | 7.9458 | 248.3051
2010-08-12 | 18:03:53 | 1.537 | 301.431 | 28.66 | 2912 | 17.4 |  | 100.3285 | 20.9686 | 208.7324 | 156.5626 | 7.9173 | 247.4169
2010-08-12 | 18:04:03 | 1.704 | 302.242 | 28.63 | 2913 | 17.4 |  | 100.5984 | 21.0251 | 209.2941 | 156.9839 | 7.9386 | 248.0827
2010-08-12 | 18:04:13 | 1.871 | 303.6 | 28.58 | 2909 | 17.4 |  | 101.0503 | 21.1195 | 210.2342 | 157.689 | 7.9743 | 249.197
2010-08-12 | 18:04:23 | 2.037 | 303.328 | 28.59 | 2916 | 17.4 |  | 100.9597 | 21.1006 | 210.0458 | 157.5477 | 7.9672 | 248.9737
2010-08-12 | 18:04:33 | 2.204 | 301.161 | 28.67 | 2922 | 17.4 |  | 100.2387 | 20.9499 | 208.5455 | 156.4224 | 7.9103 | 247.1954
2010-08-12 | 18:04:43 | 2.371 | 305.239 | 28.52 | 2920 | 17.4 |  | 101.5957 | 21.2335 | 211.3688 | 158.5401 | 8.0173 | 250.542
2010-08-12 | 18:04:53 | 2.538 | 303.6 | 28.58 | 2907 | 17.4 |  | 101.0503 | 21.1195 | 210.2342 | 157.689 | 7.9743 | 249.197
2010-08-12 | 18:05:03 | 2.705 | 302.513 | 28.62 | 2908 | 17.4 |  | 100.6886 | 21.0439 | 209.4817 | 157.1246 | 7.9458 | 248.3051
2010-08-12 | 18:05:13 | 2.872 | 303.6 | 28.58 | 2910 | 17.4 |  | 101.0503 | 21.1195 | 210.2342 | 157.689 | 7.9743 | 249.197
2010-08-12 | 18:05:23 | 3.039 | 303.872 | 28.57 | 2910 | 17.4 |  | 101.141 | 21.1385 | 210.4228 | 157.8305 | 7.9815 | 249.4206
2010-08-12 | 18:05:33 | 3.206 | 302.242 | 28.63 | 2911 | 17.4 |  | 100.5984 | 21.0251 | 209.2941 | 156.9839 | 7.9386 | 248.0827
2010-08-12 | 18:05:43 | 3.373 | 302.785 | 28.61 | 2916 | 17.4 |  | 100.7789 | 21.0628 | 209.6695 | 157.2655 | 7.9529 | 248.5277
2010-08-12 | 18:05:53 | 3.54 | 300.085 | 28.71 | 2911 | 17.4 |  | 99.8804 | 20.875 | 207.8001 | 155.8633 | 7.882 | 246.3119 |  | regression formula | time
2010-08-12 | 18:06:03 | 3.706 | 300.623 | 28.69 | 2912 | 17.4 |  | 100.0593 | 20.9124 | 208.1724 | 156.1426 | 7.8961 | 246.7532 | value for T=26 min. | 263.11 | T11
2010-08-12 | 18:06:13 | 3.873 | 306.062 | 28.49 | 2912 | 17.4 |  | 101.8697 | 21.2908 | 211.9388 | 158.9676 | 8.039 | 251.2176 | value for T=1 min. | 251.14 | T1
2010-08-12 | 18:06:23 | 4.04 | 301.972 | 28.64 | 2906 | 17.4 |  | 100.5084 | 21.0062 | 209.1066 | 156.8433 | 7.9315 | 247.8605 | difference between T25 and T1 | 11.97 | 10minutes
2010-08-12 | 18:06:33 | 4.207 | 301.972 | 28.64 | 2911 | 17.4 |  | 100.5084 | 21.0062 | 209.1066 | 156.8433 | 7.9315 | 247.8605 | calculate from regression curve values for 10 minutes of photosynthesis or respiration
2010-08-12 | 18:06:43 | 4.374 | 302.513 | 28.62 | 2913 | 17.4 |  | 100.6886 | 21.0439 | 209.4817 | 157.1246 | 7.9458 | 248.3051
2010-08-12 | 18:06:53 | 4.541 | 302.785 | 28.61 | 2902 | 17.4 |  | 100.7789 | 21.0628 | 209.6695 | 157.2655 | 7.9529 | 248.5277
2010-08-12 | 18:07:03 | 4.708 | 299.548 | 28.73 | 2907 | 17.4 |  | 99.7018 | 20.8377 | 207.4286 | 155.5847 | 7.8679 | 245.8715
2010-08-12 | 18:07:13 | 4.875 | 301.161 | 28.67 | 2901 | 17.4 |  | 100.2387 | 20.9499 | 208.5455 | 156.4224 | 7.9103 | 247.1954
2010-08-12 | 18:07:23 | 5.042 | 302.242 | 28.63 | 2902 | 17.4 |  | 100.5984 | 21.0251 | 209.2941 | 156.9839 | 7.9386 | 248.0827
2010-08-12 | 18:07:33 | 5.209 | 298.479 | 28.77 | 2903 | 17.4 |  | 99.3457 | 20.7633 | 206.6878 | 155.029 | 7.8398 | 244.9934
2010-08-12 | 18:07:43 | 5.376 | 303.056 | 28.6 | 2903 | 17.4 |  | 100.8693 | 21.0817 | 209.8575 | 157.4065 | 7.96 | 248.7506
2010-08-12 | 18:07:53 | 5.542 | 302.785 | 28.61 | 2900 | 17.4 |  | 100.7789 | 21.0628 | 209.6695 | 157.2655 | 7.9529 | 248.5277
2010-08-12 | 18:08:03 | 5.709 | 303.6 | 28.58 | 2903 | 17.4 |  | 101.0503 | 21.1195 | 210.2342 | 157.689 | 7.9743 | 249.197
2010-08-12 | 18:08:13 | 5.876 | 298.212 | 28.78 | 2909 | 17.4 |  | 99.257 | 20.7447 | 206.5031 | 154.8905 | 7.8328 | 244.7745
2010-08-12 | 18:08:23 | 6.043 | 299.013 | 28.75 | 2902 | 17.4 |  | 99.5236 | 20.8004 | 207.0578 | 155.3066 | 7.8538 | 245.432
2010-08-12 | 18:08:33 | 6.21 | 304.145 | 28.56 | 2899 | 17.4 |  | 101.2317 | 21.1574 | 210.6116 | 157.9721 | 7.9886 | 249.6444
2010-08-12 | 18:08:43 | 6.377 | 301.701 | 28.65 | 2904 | 17.4 |  | 100.4184 | 20.9874 | 208.9194 | 156.7029 | 7.9244 | 247.6386
2010-08-12 | 18:08:52 | 6.527 | 304.418 | 28.55 | 2900 | 17.4 |  | 101.3226 | 21.1764 | 210.8006 | 158.1139 | 7.9958 | 249.8684
2010-08-12 | 18:09:02 | 6.694 | 300.623 | 28.69 | 2903 | 17.4 |  | 100.0593 | 20.9124 | 208.1724 | 156.1426 | 7.8961 | 246.7532
2010-08-12 | 18:09:12 | 6.861 | 301.972 | 28.64 | 2899 | 17.4 |  | 100.5084 | 21.0062 | 209.1066 | 156.8433 | 7.9315 | 247.8605
2010-08-12 | 18:09:22 | 7.028 | 303.6 | 28.58 | 2900 | 17.4 |  | 101.0503 | 21.1195 | 210.2342 | 157.689 | 7.9743 | 249.197
2010-08-12 | 18:09:32 | 7.195 | 301.431 | 28.66 | 2901 | 17.4 |  | 100.3285 | 20.9686 | 208.7324 | 156.5626 | 7.9173 | 247.4169
2010-08-12 | 18:09:43 | 7.362 | 301.972 | 28.64 | 2904 | 17.4 |  | 100.5084 | 21.0062 | 209.1066 | 156.8433 | 7.9315 | 247.8605
2010-08-12 | 18:09:53 | 7.529 | 300.354 | 28.7 | 2898 | 17.4 |  | 99.9698 | 20.8937 | 207.9862 | 156.0029 | 7.889 | 246.5324
2010-08-12 | 18:10:03 | 7.696 | 304.145 | 28.56 | 2906 | 17.4 |  | 101.2317 | 21.1574 | 210.6116 | 157.9721 | 7.9886 | 249.6444
2010-08-12 | 18:10:13 | 7.862 | 302.785 | 28.61 | 2902 | 17.4 |  | 100.7789 | 21.0628 | 209.6695 | 157.2655 | 7.9529 | 248.5277
2010-08-12 | 18:10:23 | 8.029 | 301.431 | 28.66 | 2902 | 17.4 |  | 100.3285 | 20.9686 | 208.7324 | 156.5626 | 7.9173 | 247.4169
2010-08-12 | 18:10:33 | 8.196 | 303.328 | 28.59 | 2897 | 17.4 |  | 100.9597 | 21.1006 | 210.0458 | 157.5477 | 7.9672 | 248.9737
2010-08-12 | 18:10:43 | 8.363 | 301.161 | 28.67 | 2898 | 17.4 |  | 100.2387 | 20.9499 | 208.5455 | 156.4224 | 7.9103 | 247.1954
2010-08-12 | 18:10:53 | 8.53 | 300.085 | 28.71 | 2897 | 17.4 |  | 99.8804 | 20.875 | 207.8001 | 155.8633 | 7.882 | 246.3119
2010-08-12 | 18:11:03 | 8.697 | 300.892 | 28.68 | 2897 | 17.4 |  | 100.1489 | 20.9311 | 208.3589 | 156.2824 | 7.9032 | 246.9742
2010-08-12 | 18:11:13 | 8.864 | 296.882 | 28.83 | 2894 | 17.4 |  | 98.8144 | 20.6522 | 205.5824 | 154.1999 | 7.7979 | 243.6832
2010-08-12 | 18:11:23 | 9.031 | 301.972 | 28.64 | 2901 | 17.4 |  | 100.5084 | 21.0062 | 209.1066 | 156.8433 | 7.9315 | 247.8605
2010-08-12 | 18:11:33 | 9.198 | 300.354 | 28.7 | 2901 | 17.4 |  | 99.9698 | 20.8937 | 207.9862 | 156.0029 | 7.889 | 246.5324
2010-08-12 | 18:11:43 | 9.365 | 299.548 | 28.73 | 2899 | 17.4 |  | 99.7018 | 20.8377 | 207.4286 | 155.5847 | 7.8679 | 245.8715
2010-08-12 | 18:11:53 | 9.532 | 297.946 | 28.79 | 2899 | 17.4 |  | 99.1683 | 20.7262 | 206.3186 | 154.7521 | 7.8258 | 244.5558
2010-08-12 | 18:12:03 | 9.698 | 299.817 | 28.72 | 2904 | 17.4 |  | 99.791 | 20.8563 | 207.6143 | 155.7239 | 7.8749 | 246.0916
2010-08-12 | 18:12:13 | 9.865 | 302.785 | 28.61 | 2899 | 17.4 |  | 100.7789 | 21.0628 | 209.6695 | 157.2655 | 7.9529 | 248.5277
2010-08-12 | 18:12:23 | 10.032 | 301.161 | 28.67 | 2901 | 17.4 |  | 100.2387 | 20.9499 | 208.5455 | 156.4224 | 7.9103 | 247.1954
2010-08-12 | 18:12:33 | 10.199 | 301.972 | 28.64 | 2909 | 17.4 |  | 100.5084 | 21.0062 | 209.1066 | 156.8433 | 7.9315 | 247.8605
2010-08-12 | 18:12:43 | 10.366 | 300.623 | 28.69 | 2908 | 17.4 |  | 100.0593 | 20.9124 | 208.1724 | 156.1426 | 7.8961 | 246.7532
2010-08-12 | 18:12:53 | 10.533 | 297.679 | 28.8 | 2911 | 17.4 |  | 99.0797 | 20.7077 | 206.1343 | 154.6138 | 7.8188 | 244.3373
2010-08-12 | 18:13:03 | 10.7 | 302.785 | 28.61 | 2916 | 17.4 |  | 100.7789 | 21.0628 | 209.6695 | 157.2655 | 7.9529 | 248.5277
2010-08-12 | 18:13:13 | 10.867 | 301.701 | 28.65 | 2906 | 17.4 |  | 100.4184 | 20.9874 | 208.9194 | 156.7029 | 7.9244 | 247.6386
2010-08-12 | 18:13:23 | 11.034 | 301.431 | 28.66 | 2911 | 17.4 |  | 100.3285 | 20.9686 | 208.7324 | 156.5626 | 7.9173 | 247.4169
2010-08-12 | 18:13:33 | 11.2 | 298.746 | 28.76 | 2915 | 17.4 |  | 99.4346 | 20.7818 | 206.8727 | 155.1677 | 7.8468 | 245.2126
2010-08-12 | 18:13:43 | 11.368 | 304.145 | 28.56 | 2916 | 17.4 |  | 101.2317 | 21.1574 | 210.6116 | 157.9721 | 7.9886 | 249.6444
2010-08-12 | 18:13:53 | 11.534 | 299.281 | 28.74 | 2913 | 17.4 |  | 99.6126 | 20.819 | 207.2431 | 155.4455 | 7.8609 | 245.6516
2010-08-12 | 18:14:03 | 11.701 | 297.946 | 28.79 | 2915 | 17.4 |  | 99.1683 | 20.7262 | 206.3186 | 154.7521 | 7.8258 | 244.5558
2010-08-12 | 18:14:13 | 11.868 | 301.972 | 28.64 | 2914 | 17.4 |  | 100.5084 | 21.0062 | 209.1066 | 156.8433 | 7.9315 | 247.8605
2010-08-12 | 18:14:23 | 12.035 | 298.746 | 28.76 | 2915 | 17.4 |  | 99.4346 | 20.7818 | 206.8727 | 155.1677 | 7.8468 | 245.2126
2010-08-12 | 18:14:33 | 12.202 | 304.418 | 28.55 | 2917 | 17.4 |  | 101.3226 | 21.1764 | 210.8006 | 158.1139 | 7.9958 | 249.8684
2010-08-12 | 18:14:43 | 12.369 | 298.746 | 28.76 | 2919 | 17.4 |  | 99.4346 | 20.7818 | 206.8727 | 155.1677 | 7.8468 | 245.2126
2010-08-12 | 18:14:53 | 12.536 | 301.701 | 28.65 | 2914 | 17.4 |  | 100.4184 | 20.9874 | 208.9194 | 156.7029 | 7.9244 | 247.6386
2010-08-12 | 18:15:03 | 12.703 | 298.746 | 28.76 | 2919 | 17.4 |  | 99.4346 | 20.7818 | 206.8727 | 155.1677 | 7.8468 | 245.2126
2010-08-12 | 18:15:13 | 12.87 | 302.785 | 28.61 | 2925 | 17.4 |  | 100.7789 | 21.0628 | 209.6695 | 157.2655 | 7.9529 | 248.5277
2010-08-12 | 18:15:23 | 13.037 | 305.239 | 28.52 | 2920 | 17.4 |  | 101.5957 | 21.2335 | 211.3688 | 158.5401 | 8.0173 | 250.542
2010-08-12 | 18:15:33 | 13.203 | 301.701 | 28.65 | 2923 | 17.4 |  | 100.4184 | 20.9874 | 208.9194 | 156.7029 | 7.9244 | 247.6386
2010-08-12 | 18:15:43 | 13.37 | 300.623 | 28.69 | 2917 | 17.4 |  | 100.0593 | 20.9124 | 208.1724 | 156.1426 | 7.8961 | 246.7532
2010-08-12 | 18:15:53 | 13.537 | 300.354 | 28.7 | 2923 | 17.4 |  | 99.9698 | 20.8937 | 207.9862 | 156.0029 | 7.889 | 246.5324
2010-08-12 | 18:16:03 | 13.704 | 302.242 | 28.63 | 2919 | 17.4 |  | 100.5984 | 21.0251 | 209.2941 | 156.9839 | 7.9386 | 248.0827
2010-08-12 | 18:16:13 | 13.871 | 302.513 | 28.62 | 2920 | 17.4 |  | 100.6886 | 21.0439 | 209.4817 | 157.1246 | 7.9458 | 248.3051
2010-08-12 | 18:16:23 | 14.038 | 302.513 | 28.62 | 2920 | 17.4 |  | 100.6886 | 21.0439 | 209.4817 | 157.1246 | 7.9458 | 248.3051
2010-08-12 | 18:16:33 | 14.205 | 301.972 | 28.64 | 2918 | 17.4 |  | 100.5084 | 21.0062 | 209.1066 | 156.8433 | 7.9315 | 247.8605
2010-08-12 | 18:16:43 | 14.372 | 301.431 | 28.66 | 2926 | 17.4 |  | 100.3285 | 20.9686 | 208.7324 | 156.5626 | 7.9173 | 247.4169
2010-08-12 | 18:16:53 | 14.539 | 299.281 | 28.74 | 2917 | 17.4 |  | 99.6126 | 20.819 | 207.2431 | 155.4455 | 7.8609 | 245.6516
2010-08-12 | 18:17:03 | 14.706 | 300.892 | 28.68 | 2924 | 17.4 |  | 100.1489 | 20.9311 | 208.3589 | 156.2824 | 7.9032 | 246.9742
2010-08-12 | 18:17:13 | 14.873 | 302.513 | 28.62 | 2925 | 17.4 |  | 100.6886 | 21.0439 | 209.4817 | 157.1246 | 7.9458 | 248.3051
2010-08-12 | 18:17:23 | 15.039 | 300.623 | 28.69 | 2923 | 17.4 |  | 100.0593 | 20.9124 | 208.1724 | 156.1426 | 7.8961 | 246.7532
2010-08-12 | 18:17:33 | 15.206 | 301.431 | 28.66 | 2917 | 17.4 |  | 100.3285 | 20.9686 | 208.7324 | 156.5626 | 7.9173 | 247.4169
2010-08-12 | 18:17:43 | 15.373 | 299.013 | 28.75 | 2913 | 17.4 |  | 99.5236 | 20.8004 | 207.0578 | 155.3066 | 7.8538 | 245.432
2010-08-12 | 18:17:53 | 15.54 | 301.161 | 28.67 | 2914 | 17.4 |  | 100.2387 | 20.9499 | 208.5455 | 156.4224 | 7.9103 | 247.1954
2010-08-12 | 18:18:03 | 15.707 | 301.972 | 28.64 | 2914 | 17.4 |  | 100.5084 | 21.0062 | 209.1066 | 156.8433 | 7.9315 | 247.8605
2010-08-12 | 18:18:13 | 15.874 | 299.281 | 28.74 | 2913 | 17.4 |  | 99.6126 | 20.819 | 207.2431 | 155.4455 | 7.8609 | 245.6516
2010-08-12 | 18:18:23 | 16.041 | 302.242 | 28.63 | 2919 | 17.4 |  | 100.5984 | 21.0251 | 209.2941 | 156.9839 | 7.9386 | 248.0827
2010-08-12 | 18:18:33 | 16.208 | 302.785 | 28.61 | 2915 | 17.4 |  | 100.7789 | 21.0628 | 209.6695 | 157.2655 | 7.9529 | 248.5277
2010-08-12 | 18:18:43 | 16.375 | 300.354 | 28.7 | 2913 | 17.4 |  | 99.9698 | 20.8937 | 207.9862 | 156.0029 | 7.889 | 246.5324
2010-08-12 | 18:18:53 | 16.542 | 300.354 | 28.7 | 2913 | 17.4 |  | 99.9698 | 20.8937 | 207.9862 | 156.0029 | 7.889 | 246.5324
2010-08-12 | 18:19:03 | 16.709 | 301.431 | 28.66 | 2921 | 17.4 |  | 100.3285 | 20.9686 | 208.7324 | 156.5626 | 7.9173 | 247.4169
2010-08-12 | 18:19:13 | 16.876 | 300.354 | 28.7 | 2913 | 17.4 |  | 99.9698 | 20.8937 | 207.9862 | 156.0029 | 7.889 | 246.5324
2010-08-12 | 18:19:23 | 17.043 | 303.056 | 28.6 | 2908 | 17.4 |  | 100.8693 | 21.0817 | 209.8575 | 157.4065 | 7.96 | 248.7506
2010-08-12 | 18:19:33 | 17.209 | 300.623 | 28.69 | 2911 | 17.4 |  | 100.0593 | 20.9124 | 208.1724 | 156.1426 | 7.8961 | 246.7532
2010-08-12 | 18:19:43 | 17.376 | 301.161 | 28.67 | 2901 | 17.4 |  | 100.2387 | 20.9499 | 208.5455 | 156.4224 | 7.9103 | 247.1954
2010-08-12 | 18:19:53 | 17.543 | 301.972 | 28.64 | 2908 | 17.4 |  | 100.5084 | 21.0062 | 209.1066 | 156.8433 | 7.9315 | 247.8605
2010-08-12 | 18:20:03 | 17.71 | 303.6 | 28.58 | 2910 | 17.4 |  | 101.0503 | 21.1195 | 210.2342 | 157.689 | 7.9743 | 249.197
2010-08-12 | 18:20:13 | 17.877 | 300.892 | 28.68 | 2903 | 17.4 |  | 100.1489 | 20.9311 | 208.3589 | 156.2824 | 7.9032 | 246.9742
2010-08-12 | 18:20:22 | 18.027 | 300.354 | 28.7 | 2908 | 17.4 |  | 99.9698 | 20.8937 | 207.9862 | 156.0029 | 7.889 | 246.5324
2010-08-12 | 18:20:32 | 18.194 | 301.701 | 28.65 | 2906 | 17.4 |  | 100.4184 | 20.9874 | 208.9194 | 156.7029 | 7.9244 | 247.6386
2010-08-12 | 18:20:42 | 18.361 | 301.431 | 28.66 | 2904 | 17.4 |  | 100.3285 | 20.9686 | 208.7324 | 156.5626 | 7.9173 | 247.4169
2010-08-12 | 18:20:52 | 18.528 | 303.056 | 28.6 | 2903 | 17.4 |  | 100.8693 | 21.0817 | 209.8575 | 157.4065 | 7.96 | 248.7506
2010-08-12 | 18:21:02 | 18.695 | 305.787 | 28.5 | 2902 | 17.4 |  | 101.7782 | 21.2717 | 211.7486 | 158.825 | 8.0317 | 250.9922
2010-08-12 | 18:21:13 | 18.862 | 301.701 | 28.65 | 2904 | 17.4 |  | 100.4184 | 20.9874 | 208.9194 | 156.7029 | 7.9244 | 247.6386
2010-08-12 | 18:21:23 | 19.029 | 300.354 | 28.7 | 2906 | 17.4 |  | 99.9698 | 20.8937 | 207.9862 | 156.0029 | 7.889 | 246.5324
2010-08-12 | 18:21:33 | 19.196 | 301.431 | 28.66 | 2909 | 17.4 |  | 100.3285 | 20.9686 | 208.7324 | 156.5626 | 7.9173 | 247.4169
2010-08-12 | 18:21:43 | 19.363 | 299.817 | 28.72 | 2903 | 17.4 |  | 99.791 | 20.8563 | 207.6143 | 155.7239 | 7.8749 | 246.0916
2010-08-12 | 18:21:53 | 19.529 | 300.892 | 28.68 | 2900 | 17.4 |  | 100.1489 | 20.9311 | 208.3589 | 156.2824 | 7.9032 | 246.9742
2010-08-12 | 18:22:03 | 19.696 | 304.691 | 28.54 | 2894 | 17.4 |  | 101.4135 | 21.1954 | 210.9898 | 158.2558 | 8.003 | 250.0927
2010-08-12 | 18:22:13 | 19.863 | 300.892 | 28.68 | 2898 | 17.4 |  | 100.1489 | 20.9311 | 208.3589 | 156.2824 | 7.9032 | 246.9742
2010-08-12 | 18:22:23 | 20.03 | 298.746 | 28.76 | 2898 | 17.4 |  | 99.4346 | 20.7818 | 206.8727 | 155.1677 | 7.8468 | 245.2126
2010-08-12 | 18:22:33 | 20.197 | 303.328 | 28.59 | 2897 | 17.4 |  | 100.9597 | 21.1006 | 210.0458 | 157.5477 | 7.9672 | 248.9737
2010-08-12 | 18:22:43 | 20.364 | 300.623 | 28.69 | 2894 | 17.4 |  | 100.0593 | 20.9124 | 208.1724 | 156.1426 | 7.8961 | 246.7532
2010-08-12 | 18:22:53 | 20.531 | 304.418 | 28.55 | 2897 | 17.4 |  | 101.3226 | 21.1764 | 210.8006 | 158.1139 | 7.9958 | 249.8684
2010-08-12 | 18:23:03 | 20.698 | 303.872 | 28.57 | 2893 | 17.4 |  | 101.141 | 21.1385 | 210.4228 | 157.8305 | 7.9815 | 249.4206
2010-08-12 | 18:23:13 | 20.865 | 303.872 | 28.57 | 2903 | 17.4 |  | 101.141 | 21.1385 | 210.4228 | 157.8305 | 7.9815 | 249.4206
2010-08-12 | 18:23:23 | 21.032 | 305.787 | 28.5 | 2897 | 17.4 |  | 101.7782 | 21.2717 | 211.7486 | 158.825 | 8.0317 | 250.9922
2010-08-12 | 18:23:33 | 21.198 | 303.328 | 28.59 | 2896 | 17.4 |  | 100.9597 | 21.1006 | 210.0458 | 157.5477 | 7.9672 | 248.9737
2010-08-12 | 18:23:43 | 21.365 | 304.965 | 28.53 | 2898 | 17.4 |  | 101.5045 | 21.2145 | 211.1792 | 158.3979 | 8.0102 | 250.3172
2010-08-12 | 18:23:53 | 21.532 | 304.691 | 28.54 | 2897 | 17.4 |  | 101.4135 | 21.1954 | 210.9898 | 158.2558 | 8.003 | 250.0927
2010-08-12 | 18:24:03 | 21.699 | 304.418 | 28.55 | 2900 | 17.4 |  | 101.3226 | 21.1764 | 210.8006 | 158.1139 | 7.9958 | 249.8684
2010-08-12 | 18:24:13 | 21.866 | 306.062 | 28.49 | 2902 | 17.4 |  | 101.8697 | 21.2908 | 211.9388 | 158.9676 | 8.039 | 251.2176
2010-08-12 | 18:24:23 | 22.033 | 305.787 | 28.5 | 2890 | 17.4 |  | 101.7782 | 21.2717 | 211.7486 | 158.825 | 8.0317 | 250.9922
2010-08-12 | 18:24:33 | 22.2 | 304.418 | 28.55 | 2894 | 17.4 |  | 101.3226 | 21.1764 | 210.8006 | 158.1139 | 7.9958 | 249.8684
2010-08-12 | 18:24:43 | 22.367 | 305.513 | 28.51 | 2894 | 17.4 |  | 101.6869 | 21.2526 | 211.5586 | 158.6824 | 8.0245 | 250.7669
2010-08-12 | 18:24:53 | 22.534 | 305.291 | 28.56 | 2891 | 17.3 |  | 101.4031 | 21.1933 | 210.9949 | 158.2596 | 8.0174 | 250.5426
2010-08-12 | 18:25:03 | 22.701 | 307.215 | 28.49 | 2894 | 17.3 |  | 102.0421 | 21.3268 | 212.3245 | 159.2569 | 8.0679 | 252.1213
2010-08-12 | 18:25:13 | 22.868 | 306.114 | 28.53 | 2897 | 17.3 |  | 101.6764 | 21.2504 | 211.5635 | 158.6861 | 8.039 | 251.2178
2010-08-12 | 18:25:23 | 23.034 | 306.94 | 28.5 | 2894 | 17.3 |  | 101.9506 | 21.3077 | 212.1339 | 159.114 | 8.0606 | 251.8951
2010-08-12 | 18:25:33 | 23.201 | 305.018 | 28.57 | 2889 | 17.3 |  | 101.3122 | 21.1743 | 210.8058 | 158.1178 | 8.0102 | 250.318
2010-08-12 | 18:25:43 | 23.368 | 302.568 | 28.66 | 2892 | 17.3 |  | 100.4984 | 21.0042 | 209.1124 | 156.8476 | 7.9458 | 248.3072
2010-08-12 | 18:25:53 | 23.535 | 307.768 | 28.47 | 2894 | 17.3 |  | 102.2256 | 21.3651 | 212.7062 | 159.5432 | 8.0824 | 252.5746
2010-08-12 | 18:26:03 | 23.702 | 305.291 | 28.56 | 2883 | 17.3 |  | 101.4031 | 21.1933 | 210.9949 | 158.2596 | 8.0174 | 250.5426
2010-08-12 | 18:26:13 | 23.869 | 308.044 | 28.46 | 2898 | 17.3 |  | 102.3174 | 21.3843 | 212.8973 | 159.6865 | 8.0896 | 252.8015
2010-08-12 | 18:26:23 | 24.036 | 308.598 | 28.44 | 2892 | 17.3 |  | 102.5014 | 21.4228 | 213.2802 | 159.9737 | 8.1042 | 253.2562
2010-08-12 | 18:26:33 | 24.203 | 303.926 | 28.61 | 2893 | 17.3 |  | 100.9496 | 21.0985 | 210.0512 | 157.5518 | 7.9815 | 249.4219
2010-08-12 | 18:26:43 | 24.37 | 305.84 | 28.54 | 2891 | 17.3 |  | 101.5852 | 21.2313 | 211.3738 | 158.5438 | 8.0318 | 250.9925
2010-08-12 | 18:26:53 | 24.537 | 310.546 | 28.37 | 2886 | 17.3 |  | 103.1485 | 21.558 | 214.6266 | 160.9837 | 8.1554 | 254.855
2010-08-12 | 18:27:03 | 24.704 | 305.84 | 28.54 | 2887 | 17.3 |  | 101.5852 | 21.2313 | 211.3738 | 158.5438 | 8.0318 | 250.9925
2010-08-12 | 18:27:13 | 24.87 | 308.598 | 28.44 | 2885 | 17.3 |  | 102.5014 | 21.4228 | 213.2802 | 159.9737 | 8.1042 | 253.2562
2010-08-12 | 18:27:23 | 25.037 | 309.153 | 28.42 | 2883 | 17.3 |  | 102.6858 | 21.4613 | 213.6639 | 160.2615 | 8.1188 | 253.7118
2010-08-12 | 18:27:33 | 25.204 | 307.768 | 28.47 | 2884 | 17.3 |  | 102.2256 | 21.3651 | 212.7062 | 159.5432 | 8.0824 | 252.5746
2010-08-12 | 18:27:43 | 25.371 | 309.102 | 28.38 | 2876 | 17.4 |  | 102.8817 | 21.5023 | 214.0443 | 160.5469 | 8.1188 | 253.7133
2010-08-12 | 18:27:53 | 25.538 | 306.887 | 28.46 | 2882 | 17.4 |  | 102.1445 | 21.3482 | 212.5106 | 159.3965 | 8.0607 | 251.8954
2010-08-12 | 18:28:03 | 25.705 | 307.716 | 28.43 | 2880 | 17.4 |  | 102.4202 | 21.4058 | 213.0842 | 159.8268 | 8.0824 | 252.5753
2010-08-12 | 18:28:13 | 25.872 | 310.496 | 28.33 | 2884 | 17.4 |  | 103.3456 | 21.5992 | 215.0095 | 161.2708 | 8.1554 | 254.8573
2010-08-12 | 18:28:23 | 26.039 | 310.217 | 28.34 | 2871 | 17.4 |  | 103.2526 | 21.5798 | 214.816 | 161.1257 | 8.1481 | 254.628
2010-08-12 | 18:28:33 | 26.206 | 309.938 | 28.35 | 2877 | 17.4 |  | 103.1597 | 21.5604 | 214.6228 | 160.9808 | 8.1408 | 254.399
2010-08-12 | 18:28:43 | 26.373 | 307.163 | 28.45 | 2879 | 17.4 |  | 102.2363 | 21.3674 | 212.7016 | 159.5398 | 8.0679 | 252.1218
2010-08-12 | 18:28:53 | 26.539 | 307.716 | 28.43 | 2875 | 17.4 |  | 102.4202 | 21.4058 | 213.0842 | 159.8268 | 8.0824 | 252.5753
2010-08-12 | 18:29:03 | 26.706 | 311.616 | 28.29 | 2871 | 17.4 |  | 103.7185 | 21.6772 | 215.7853 | 161.8527 | 8.1849 | 255.7769
2010-08-12 | 18:29:13 | 26.873 | 307.716 | 28.43 | 2872 | 17.4 |  | 102.4202 | 21.4058 | 213.0842 | 159.8268 | 8.0824 | 252.5753
2010-08-12 | 18:29:23 | 27.04 | 306.887 | 28.46 | 2868 | 17.4 |  | 102.1445 | 21.3482 | 212.5106 | 159.3965 | 8.0607 | 251.8954
2010-08-12 | 18:29:33 | 27.207 | 310.217 | 28.34 | 2870 | 17.4 |  | 103.2526 | 21.5798 | 214.816 | 161.1257 | 8.1481 | 254.628
2010-08-12 | 18:29:43 | 27.374 | 310.496 | 28.33 | 2874 | 17.4 |  | 103.3456 | 21.5992 | 215.0095 | 161.2708 | 8.1554 | 254.8573
2010-08-12 | 18:29:53 | 27.541 | 309.938 | 28.35 | 2868 | 17.4 |  | 103.1597 | 21.5604 | 214.6228 | 160.9808 | 8.1408 | 254.399
2010-08-12 | 18:30:03 | 27.708 | 316.717 | 28.11 | 2872 | 17.4 |  | 105.4162 | 22.032 | 219.3174 | 164.502 | 8.3188 | 259.9636
2010-08-12 | 18:30:13 | 27.875 | 314.155 | 28.2 | 2871 | 17.4 |  | 104.5633 | 21.8537 | 217.5429 | 163.171 | 8.2515 | 257.8603
2010-08-12 | 18:30:23 | 28.042 | 315.29 | 28.16 | 2867 | 17.4 |  | 104.9414 | 21.9327 | 218.3294 | 163.761 | 8.2814 | 258.7926
2010-08-12 | 18:30:33 | 28.209 | 310.496 | 28.33 | 2864 | 17.4 |  | 103.3456 | 21.5992 | 215.0095 | 161.2708 | 8.1554 | 254.8573
2010-08-12 | 18:30:43 | 28.375 | 308.824 | 28.39 | 2867 | 17.4 |  | 102.7892 | 21.4829 | 213.8519 | 160.4026 | 8.1115 | 253.4852
2010-08-12 | 18:30:53 | 28.542 | 313.306 | 28.23 | 2865 | 17.4 |  | 104.2808 | 21.7947 | 216.9551 | 162.7302 | 8.2292 | 257.1636
2010-08-12 | 18:31:03 | 28.709 | 313.588 | 28.22 | 2866 | 17.4 |  | 104.3749 | 21.8143 | 217.1508 | 162.877 | 8.2367 | 257.3956
2010-08-12 | 18:31:13 | 28.876 | 313.871 | 28.21 | 2860 | 17.4 |  | 104.469 | 21.834 | 217.3468 | 163.0239 | 8.2441 | 257.6278
2010-08-12 | 18:31:23 | 29.043 | 318.198 | 28.1 | 2866 | 17.3 |  | 105.6899 | 22.0892 | 219.9146 | 164.95 | 8.3563 | 261.1341
2010-08-12 | 18:31:33 | 29.21 | 316.478 | 28.16 | 2860 | 17.3 |  | 105.1189 | 21.9698 | 218.7264 | 164.0587 | 8.3111 | 259.7232
2010-08-12 | 18:31:43 | 29.377 | 315.054 | 28.21 | 2853 | 17.3 |  | 104.6458 | 21.871 | 217.742 | 163.3204 | 8.2737 | 258.5543
2010-08-12 | 18:31:52 | 29.527 | 314.77 | 28.22 | 2857 | 17.3 |  | 104.5514 | 21.8512 | 217.5457 | 163.1731 | 8.2663 | 258.3212
2010-08-12 | 18:32:02 | 29.694 | 317.623 | 28.12 | 2855 | 17.3 |  | 105.4992 | 22.0493 | 219.5177 | 164.6523 | 8.3412 | 260.6628
2010-08-12 | 18:32:12 | 29.861 | 312.509 | 28.3 | 2855 | 17.3 |  | 103.8005 | 21.6943 | 215.9831 | 162.0011 | 8.2069 | 256.4657
2010-08-12 | 18:32:22 | 30.028 | 315.623 | 28.19 | 2847 | 17.3 |  | 104.8347 | 21.9105 | 218.1351 | 163.6152 | 8.2887 | 259.0211
2010-08-12 | 18:32:32 | 30.195 | 318.773 | 28.08 | 2855 | 17.3 |  | 105.8811 | 22.1291 | 220.3124 | 165.2483 | 8.3714 | 261.6065
2010-08-12 | 18:32:43 | 30.362 | 315.908 | 28.18 | 2853 | 17.3 |  | 104.9293 | 21.9302 | 218.332 | 163.7629 | 8.2962 | 259.2549
2010-08-12 | 18:32:53 | 30.529 | 316.764 | 28.15 | 2847 | 17.3 |  | 105.2138 | 21.9897 | 218.9239 | 164.2069 | 8.3186 | 259.9577
2010-08-12 | 18:33:03 | 30.695 | 316.193 | 28.17 | 2848 | 17.3 |  | 105.024 | 21.95 | 218.5291 | 163.9107 | 8.3036 | 259.4889
2010-08-12 | 18:33:13 | 30.862 | 319.639 | 28.05 | 2846 | 17.3 |  | 106.1686 | 22.1892 | 220.9106 | 165.6971 | 8.3941 | 262.3168
2010-08-12 | 18:33:23 | 31.029 | 313.92 | 28.25 | 2841 | 17.3 |  | 104.2691 | 21.7922 | 216.9582 | 162.7325 | 8.244 | 257.6236
2010-08-12 | 18:33:33 | 31.196 | 315.338 | 28.2 | 2847 | 17.3 |  | 104.7402 | 21.8907 | 217.9384 | 163.4677 | 8.2812 | 258.7875
2010-08-12 | 18:33:43 | 31.363 | 317.623 | 28.12 | 2852 | 17.3 |  | 105.4992 | 22.0493 | 219.5177 | 164.6523 | 8.3412 | 260.6628
2010-08-12 | 18:33:53 | 31.53 | 320.797 | 28.01 | 2842 | 17.3 |  | 106.5534 | 22.2697 | 221.7113 | 166.2976 | 8.4246 | 263.2676
2010-08-12 | 18:34:03 | 31.697 | 315.908 | 28.18 | 2844 | 17.3 |  | 104.9293 | 21.9302 | 218.332 | 163.7629 | 8.2962 | 259.2549
2010-08-12 | 18:34:13 | 31.864 | 321.378 | 27.99 | 2836 | 17.3 |  | 106.7464 | 22.31 | 222.1129 | 166.5988 | 8.4398 | 263.7444
2010-08-12 | 18:34:23 | 32.031 | 319.639 | 28.05 | 2840 | 17.3 |  | 106.1686 | 22.1892 | 220.9106 | 165.6971 | 8.3941 | 262.3168
2010-08-12 | 18:34:33 | 32.198 | 318.485 | 28.09 | 2836 | 17.3 |  | 105.7855 | 22.1092 | 220.1134 | 165.0991 | 8.3638 | 261.3702
2010-08-12 | 18:34:43 | 32.365 | 321.961 | 27.97 | 2838 | 17.3 |  | 106.9398 | 22.3504 | 222.5154 | 166.9007 | 8.4551 | 264.2223
2010-08-12 | 18:34:53 | 32.531 | 319.061 | 28.07 | 2835 | 17.3 |  | 105.9768 | 22.1492 | 220.5116 | 165.3978 | 8.379 | 261.843
2010-08-12 | 18:35:03 | 32.698 | 321.961 | 27.97 | 2830 | 17.3 |  | 106.9398 | 22.3504 | 222.5154 | 166.9007 | 8.4551 | 264.2223
2010-08-12 | 18:35:13 | 32.865 | 319.061 | 28.07 | 2828 | 17.3 |  | 105.9768 | 22.1492 | 220.5116 | 165.3978 | 8.379 | 261.843
2010-08-12 | 18:35:23 | 33.032 | 321.378 | 27.99 | 2830 | 17.3 |  | 106.7464 | 22.31 | 222.1129 | 166.5988 | 8.4398 | 263.7444
2010-08-12 | 18:35:33 | 33.199 | 321.961 | 27.97 | 2830 | 17.3 |  | 106.9398 | 22.3504 | 222.5154 | 166.9007 | 8.4551 | 264.2223
2010-08-12 | 18:35:43 | 33.366 | 323.129 | 27.93 | 2834 | 17.3 |  | 107.3279 | 22.4315 | 223.3229 | 167.5064 | 8.4858 | 265.1812
2010-08-12 | 18:35:53 | 33.533 | 319.35 | 28.06 | 2828 | 17.3 |  | 106.0727 | 22.1692 | 220.711 | 165.5473 | 8.3866 | 262.0798
2010-08-12 | 18:36:03 | 33.7 | 324.302 | 27.89 | 2822 | 17.3 |  | 107.7177 | 22.513 | 224.1339 | 168.1147 | 8.5166 | 266.1442
2010-08-12 | 18:36:13 | 33.867 | 327.555 | 27.78 | 2821 | 17.3 |  | 108.7982 | 22.7388 | 226.3821 | 169.801 | 8.602 | 268.8139
2010-08-12 | 18:36:23 | 34.034 | 322.836 | 27.94 | 2821 | 17.3 |  | 107.2307 | 22.4112 | 223.1207 | 167.3547 | 8.4781 | 264.9411
2010-08-12 | 18:36:33 | 34.201 | 322.836 | 27.94 | 2820 | 17.3 |  | 107.2307 | 22.4112 | 223.1207 | 167.3547 | 8.4781 | 264.9411
2010-08-12 | 18:36:43 | 34.367 | 322.836 | 27.94 | 2818 | 17.3 |  | 107.2307 | 22.4112 | 223.1207 | 167.3547 | 8.4781 | 264.9411
2010-08-12 | 18:36:53 | 34.534 | 320.797 | 28.01 | 2825 | 17.3 |  | 106.5534 | 22.2697 | 221.7113 | 166.2976 | 8.4246 | 263.2676
2010-08-12 | 18:37:03 | 34.701 | 321.917 | 27.93 | 2813 | 17.4 |  | 107.1468 | 22.3937 | 222.9178 | 167.2026 | 8.4554 | 264.2314
2010-08-12 | 18:37:13 | 34.868 | 319.016 | 28.03 | 2819 | 17.4 |  | 106.1813 | 22.1919 | 220.909 | 165.6959 | 8.3792 | 261.8503