## convers.+compens.
file name: PSt1_Eq4_Sal |  |  | 2008-03-29
author: Dr. Christian Huber
programm for temperature compensated oxygen calculation of membrane PSt1
Fit of the calibration curve is optimized between 0 and 250 % air-saturation (0 and 388 Torr)
enter your values here
input of F0,T0 | 58.87 | phase angle of cal 0
input of F100,T100 | 30.69 | phase angle of cal 100 (water-vapor saturated air)
input of Fm,Tm | 15 | measured phase angle
input of T0 | 17.7 | temperature of cal 0
input of T100 | 17.8 | temperatur of cal100
input of Tm | 17.7 | temperature at measurement
air pressure patm | 1013 | air pressure
Salinity | 32.2 | Salinity
Results of calculation + compensation
[%] air saturation | 493.4732 | %
[%] O2 | 103.1359 | %
pO2 [hPa] | 1026.2727 | hPa
pO2 [Torr] | 769.7699 | Torr
cO2 [mg/L] | 38.7213 | mg/L
cO2 [ppm] | 38.7213 | ppm
cO2 [µmol/L] | 1210.0413 | µmol/L
Internal calculated parameter (DO NOT CHANGE !!)
f1 | 0.801
DF/K | -0.08
DKSV/K | 0.000383
m | 22.9
tan(F0,T100) | 1.6552
tan(F0,Tm) | 1.6558
tan(F100,T100) | 0.59352
tan(Fm,Tm) | 0.26795
KSV,T100 | 0.033331
KSV,Tm | 0.033293
A | 156.5815
B | 14.0252
C | -0.64143
a | 7.8328e-06
b | -0.0021668
c | -0.83817
Chlorinity | 17.8227

## Recalculate Fibox values
file name: PSt1_Eq4_Sal |  |  |  | 2008-03-29
author: Dr. Christian Huber
programm for temperature compensated oxygen calculation of membrane PSt1
Fit of the calibration curve is optimized between 0 and 250 % air-saturation (0 and 388 Torr)
Insert your calibration values
cal0 | 58.87 | ° | T0 | 17.7 | °C
cal100 | 30.69 | ° | T100 | 17.8 | °C
air pressure | 1013 | mbar
Salinity | 32.2 | °/°°
Internal calculated parameter (DO NOT CHANGE !!)
f1 | 0.801 | tan(F0,T100) | 1.6552 | A | 156.5815 | KSV,T100 | 0.033331 | a | 1.6587e-05
DF/K | -0.08 | tan(F0,Tm) | 1.6578 | B | 14.0252 | KSV,Tm | 0.03314 | b | 0.0042086 |  |  |  |  | INSERT YOUR RESPECTIVE DATA
DKSV/K | 0.000383 | tan(F100,T100) | 0.59352 | C | -0.64143 |  |  | c | -0.65413 |  |  |  |  | chamber volume [L] | FW alga [g] | mg Chl a Thallus-1
m | 22.9 | tan(Fm,Tm) | 0.5734 |  |  |  |  | Chlorinity | 17.8227 |  |  |  |  | 0.026 | 0.0827 | 0.026479
Insert the data you get from the Ascii File recorded with the OxyView TX3-V5.31.exe software | [7 empty] | The oxygen contents are calculated automatically from the raw data
date | time/hh:mm:ss | logtime/min | oxygen/% airsatur. | phase/° | amp | temp/°C |  | % air-sat. | % oxygen | pO2 (hPa) | pO2 (Torr) | cO2 [mg/L] | cO2 [µmol/L] | cO2 [µmol/L 10 min] | cO2 [µmol/L h] | cO2 [µmol/ h] | cO2 [µmol/ h gFW] | cO2 [µmol/ h g Chl a]
2010-08-12 | 13:26:34 | 0 | 108.781 | 29.83 | 6080 | 17.3 |  | 108.7856 | 22.7362 | 226.3561 | 169.7815 | 8.6011 | 268.7829 | 24 | 144 | 3.744 | 45.2721 | 141.3924
2010-08-12 | 13:26:45 | 0.184 | 109.643 | 29.74 | 6068 | 17.3 |  | 109.648 | 22.9164 | 228.1503 | 171.1273 | 8.6692 | 270.9135
2010-08-12 | 13:26:55 | 0.351 | 110.805 | 29.62 | 6088 | 17.3 |  | 110.81 | 23.1593 | 230.5682 | 172.9409 | 8.7611 | 273.7846 | change data input according to column N
2010-08-12 | 13:27:05 | 0.517 | 110.416 | 29.66 | 6079 | 17.3 |  | 110.4211 | 23.078 | 229.759 | 172.3339 | 8.7304 | 272.8237
2010-08-12 | 13:27:15 | 0.684 | 109.451 | 29.76 | 6096 | 17.3 |  | 109.4557 | 22.8762 | 227.7502 | 170.8272 | 8.654 | 270.4384
2010-08-12 | 13:27:25 | 0.851 | 109.932 | 29.71 | 6086 | 17.3 |  | 109.9371 | 22.9769 | 228.7521 | 171.5786 | 8.6921 | 271.628
2010-08-12 | 13:27:35 | 1.018 | 110.805 | 29.62 | 6084 | 17.3 |  | 110.81 | 23.1593 | 230.5682 | 172.9409 | 8.7611 | 273.7846
2010-08-12 | 13:27:45 | 1.185 | 110.61 | 29.64 | 6092 | 17.3 |  | 110.6153 | 23.1186 | 230.1632 | 172.6371 | 8.7457 | 273.3036
2010-08-12 | 13:27:56 | 1.352 | 110.805 | 29.62 | 6079 | 17.3 |  | 110.81 | 23.1593 | 230.5682 | 172.9409 | 8.7611 | 273.7846
2010-08-12 | 13:28:06 | 1.519 | 110.61 | 29.64 | 6090 | 17.3 |  | 110.6153 | 23.1186 | 230.1632 | 172.6371 | 8.7457 | 273.3036
2010-08-12 | 13:28:16 | 1.686 | 111.286 | 29.55 | 6074 | 17.4 |  | 111.2896 | 23.2595 | 231.537 | 173.6675 | 8.7823 | 274.4479
2010-08-12 | 13:28:26 | 1.852 | 109.155 | 29.77 | 6068 | 17.4 |  | 109.1587 | 22.8142 | 227.1036 | 170.3422 | 8.6142 | 269.1929
2010-08-12 | 13:28:36 | 2.019 | 109.539 | 29.73 | 6091 | 17.4 |  | 109.5426 | 22.8944 | 227.9024 | 170.9413 | 8.6445 | 270.1397
2010-08-12 | 13:28:46 | 2.186 | 111.09 | 29.57 | 6089 | 17.4 |  | 111.094 | 23.2186 | 231.1299 | 173.3621 | 8.7669 | 273.9654
2010-08-12 | 13:28:56 | 2.353 | 109.828 | 29.7 | 6081 | 17.4 |  | 109.8316 | 22.9548 | 228.5036 | 171.3922 | 8.6673 | 270.8523
2010-08-12 | 13:29:06 | 2.52 | 109.443 | 29.74 | 6083 | 17.4 |  | 109.4465 | 22.8743 | 227.7024 | 170.7913 | 8.6369 | 269.9027
2010-08-12 | 13:29:16 | 2.687 | 110.214 | 29.66 | 6082 | 17.4 |  | 110.2183 | 23.0356 | 229.308 | 171.9956 | 8.6978 | 271.8058
2010-08-12 | 13:29:26 | 2.854 | 109.539 | 29.73 | 6084 | 17.4 |  | 109.5426 | 22.8944 | 227.9024 | 170.9413 | 8.6445 | 270.1397
2010-08-12 | 13:29:36 | 3.021 | 110.408 | 29.64 | 6076 | 17.4 |  | 110.4122 | 23.0761 | 229.7114 | 172.2982 | 8.7131 | 272.284
2010-08-12 | 13:29:46 | 3.188 | 109.731 | 29.71 | 6088 | 17.4 |  | 109.7352 | 22.9347 | 228.303 | 171.2418 | 8.6597 | 270.6145
2010-08-12 | 13:29:56 | 3.355 | 110.311 | 29.65 | 6078 | 17.4 |  | 110.3152 | 23.0559 | 229.5096 | 172.1468 | 8.7054 | 272.0448
2010-08-12 | 13:30:06 | 3.522 | 110.11 | 29.65 | 6073 | 17.5 |  | 110.1127 | 23.0136 | 229.0592 | 171.809 | 8.673 | 271.0308
2010-08-12 | 13:30:16 | 3.689 | 109.723 | 29.69 | 6068 | 17.5 |  | 109.7263 | 22.9328 | 228.2555 | 171.2062 | 8.6426 | 270.0799 |  | regression formula | time
2010-08-12 | 13:30:26 | 3.856 | 110.117 | 29.67 | 6074 | 17.4 |  | 110.1214 | 23.0154 | 229.1066 | 171.8445 | 8.6901 | 271.5671 | value for T=26 min. | 316.86 | T11
2010-08-12 | 13:30:36 | 4.022 | 110.311 | 29.65 | 6070 | 17.4 |  | 110.3152 | 23.0559 | 229.5096 | 172.1468 | 8.7054 | 272.0448 | value for T=1 min. | 292.86 | T1
2010-08-12 | 13:30:46 | 4.189 | 110.408 | 29.64 | 6073 | 17.4 |  | 110.4122 | 23.0761 | 229.7114 | 172.2982 | 8.7131 | 272.284 | difference between T25 and T1 | 24 | 10minutes
2010-08-12 | 13:30:56 | 4.356 | 110.311 | 29.65 | 6071 | 17.4 |  | 110.3152 | 23.0559 | 229.5096 | 172.1468 | 8.7054 | 272.0448 | calculate from regression curve values for 10 minutes of photosynthesis or respiration
2010-08-12 | 13:31:06 | 4.523 | 109.347 | 29.75 | 6077 | 17.4 |  | 109.3505 | 22.8543 | 227.5026 | 170.6415 | 8.6293 | 269.6658
2010-08-12 | 13:31:16 | 4.69 | 109.635 | 29.72 | 6070 | 17.4 |  | 109.6389 | 22.9145 | 228.1026 | 171.0915 | 8.6521 | 270.377
2010-08-12 | 13:31:26 | 4.857 | 109.828 | 29.7 | 6080 | 17.4 |  | 109.8316 | 22.9548 | 228.5036 | 171.3922 | 8.6673 | 270.8523
2010-08-12 | 13:31:36 | 5.024 | 111.58 | 29.52 | 6080 | 17.4 |  | 111.5839 | 23.321 | 232.1492 | 174.1267 | 8.8056 | 275.1737
2010-08-12 | 13:31:46 | 5.191 | 110.222 | 29.68 | 6056 | 17.3 |  | 110.2272 | 23.0375 | 229.3556 | 172.0313 | 8.715 | 272.3447
2010-08-12 | 13:31:56 | 5.358 | 111.587 | 29.54 | 6061 | 17.3 |  | 111.5925 | 23.3228 | 232.1965 | 174.1622 | 8.823 | 275.7181
2010-08-12 | 13:32:06 | 5.525 | 111.195 | 29.58 | 6076 | 17.3 |  | 111.2005 | 23.2409 | 231.3807 | 173.5503 | 8.792 | 274.7494
2010-08-12 | 13:32:16 | 5.692 | 111.195 | 29.58 | 6073 | 17.3 |  | 111.2005 | 23.2409 | 231.3807 | 173.5503 | 8.792 | 274.7494
2010-08-12 | 13:32:26 | 5.859 | 110.708 | 29.63 | 6064 | 17.3 |  | 110.7126 | 23.1389 | 230.3656 | 172.7889 | 8.7534 | 273.544
2010-08-12 | 13:32:36 | 6.025 | 111.883 | 29.51 | 6069 | 17.3 |  | 111.8876 | 23.3845 | 232.8106 | 174.6228 | 8.8463 | 276.4472
2010-08-12 | 13:32:46 | 6.192 | 110.513 | 29.65 | 6063 | 17.3 |  | 110.5182 | 23.0983 | 229.961 | 172.4854 | 8.738 | 273.0635
2010-08-12 | 13:32:56 | 6.359 | 111.195 | 29.58 | 6076 | 17.3 |  | 111.2005 | 23.2409 | 231.3807 | 173.5503 | 8.792 | 274.7494
2010-08-12 | 13:33:06 | 6.526 | 110.805 | 29.62 | 6061 | 17.3 |  | 110.81 | 23.1593 | 230.5682 | 172.9409 | 8.7611 | 273.7846
2010-08-12 | 13:33:16 | 6.693 | 111.195 | 29.58 | 6066 | 17.3 |  | 111.2005 | 23.2409 | 231.3807 | 173.5503 | 8.792 | 274.7494
2010-08-12 | 13:33:26 | 6.86 | 110.513 | 29.65 | 6067 | 17.3 |  | 110.5182 | 23.0983 | 229.961 | 172.4854 | 8.738 | 273.0635
2010-08-12 | 13:33:36 | 7.027 | 110.416 | 29.66 | 6081 | 17.3 |  | 110.4211 | 23.078 | 229.759 | 172.3339 | 8.7304 | 272.8237
2010-08-12 | 13:33:46 | 7.194 | 111.293 | 29.57 | 6074 | 17.3 |  | 111.2983 | 23.2614 | 231.5844 | 173.703 | 8.7997 | 274.9912
2010-08-12 | 13:33:56 | 7.361 | 111.489 | 29.55 | 6056 | 17.3 |  | 111.4944 | 23.3023 | 231.9923 | 174.009 | 8.8152 | 275.4755
2010-08-12 | 13:34:06 | 7.527 | 111.489 | 29.55 | 6071 | 17.3 |  | 111.4944 | 23.3023 | 231.9923 | 174.009 | 8.8152 | 275.4755
2010-08-12 | 13:34:16 | 7.694 | 111 | 29.6 | 6056 | 17.3 |  | 111.005 | 23.2001 | 230.974 | 173.2453 | 8.7765 | 274.2665
2010-08-12 | 13:34:26 | 7.861 | 110.319 | 29.67 | 6066 | 17.3 |  | 110.3241 | 23.0577 | 229.5572 | 172.1825 | 8.7227 | 272.5841
2010-08-12 | 13:34:36 | 8.028 | 111.195 | 29.58 | 6067 | 17.3 |  | 111.2005 | 23.2409 | 231.3807 | 173.5503 | 8.792 | 274.7494
2010-08-12 | 13:34:46 | 8.195 | 110.902 | 29.61 | 6068 | 17.3 |  | 110.9075 | 23.1797 | 230.771 | 173.093 | 8.7688 | 274.0254
2010-08-12 | 13:34:56 | 8.362 | 110.708 | 29.63 | 6072 | 17.3 |  | 110.7126 | 23.1389 | 230.3656 | 172.7889 | 8.7534 | 273.544
2010-08-12 | 13:35:06 | 8.529 | 111.195 | 29.58 | 6067 | 17.3 |  | 111.2005 | 23.2409 | 231.3807 | 173.5503 | 8.792 | 274.7494
2010-08-12 | 13:35:16 | 8.696 | 111.188 | 29.56 | 6062 | 17.4 |  | 111.1917 | 23.2391 | 231.3333 | 173.5147 | 8.7746 | 274.2065
2010-08-12 | 13:35:25 | 8.846 | 109.828 | 29.7 | 6063 | 17.4 |  | 109.8316 | 22.9548 | 228.5036 | 171.3922 | 8.6673 | 270.8523
2010-08-12 | 13:35:35 | 9.013 | 111.777 | 29.5 | 6067 | 17.4 |  | 111.7806 | 23.3621 | 232.5584 | 174.4337 | 8.8211 | 275.6587
2010-08-12 | 13:35:45 | 9.18 | 110.992 | 29.58 | 6063 | 17.4 |  | 110.9963 | 23.1982 | 230.9266 | 173.2097 | 8.7592 | 273.7244
2010-08-12 | 13:35:55 | 9.347 | 110.117 | 29.67 | 6057 | 17.4 |  | 110.1214 | 23.0154 | 229.1066 | 171.8445 | 8.6901 | 271.5671
2010-08-12 | 13:36:05 | 9.514 | 110.992 | 29.58 | 6066 | 17.4 |  | 110.9963 | 23.1982 | 230.9266 | 173.2097 | 8.7592 | 273.7244
2010-08-12 | 13:36:15 | 9.68 | 109.539 | 29.73 | 6072 | 17.4 |  | 109.5426 | 22.8944 | 227.9024 | 170.9413 | 8.6445 | 270.1397
2010-08-12 | 13:36:25 | 9.847 | 111.384 | 29.54 | 6051 | 17.4 |  | 111.3876 | 23.28 | 231.7409 | 173.8204 | 8.7901 | 274.6896
2010-08-12 | 13:36:35 | 10.014 | 110.895 | 29.59 | 6049 | 17.4 |  | 110.8987 | 23.1778 | 230.7236 | 173.0574 | 8.7515 | 273.4838
2010-08-12 | 13:36:45 | 10.181 | 111.58 | 29.52 | 6048 | 17.4 |  | 111.5839 | 23.321 | 232.1492 | 174.1267 | 8.8056 | 275.1737
2010-08-12 | 13:36:55 | 10.348 | 111.188 | 29.56 | 6061 | 17.4 |  | 111.1917 | 23.2391 | 231.3333 | 173.5147 | 8.7746 | 274.2065
2010-08-12 | 13:37:05 | 10.515 | 110.992 | 29.58 | 6056 | 17.4 |  | 110.9963 | 23.1982 | 230.9266 | 173.2097 | 8.7592 | 273.7244
2010-08-12 | 13:37:15 | 10.682 | 110.207 | 29.64 | 6046 | 17.5 |  | 110.2095 | 23.0338 | 229.2607 | 171.9601 | 8.6806 | 271.2692
2010-08-12 | 13:37:25 | 10.849 | 111.376 | 29.52 | 6057 | 17.5 |  | 111.3792 | 23.2783 | 231.6939 | 173.7852 | 8.7727 | 274.1482
2010-08-12 | 13:37:35 | 11.016 | 110.887 | 29.57 | 6050 | 17.5 |  | 110.8901 | 23.176 | 230.6764 | 173.022 | 8.7342 | 272.9443
2010-08-12 | 13:37:45 | 11.183 | 110.595 | 29.6 | 6059 | 17.5 |  | 110.5978 | 23.1149 | 230.0685 | 172.566 | 8.7112 | 272.225
2010-08-12 | 13:37:55 | 11.35 | 110.498 | 29.61 | 6057 | 17.5 |  | 110.5006 | 23.0946 | 229.8662 | 172.4143 | 8.7035 | 271.9856
2010-08-12 | 13:38:05 | 11.517 | 111.278 | 29.53 | 6049 | 17.5 |  | 111.2812 | 23.2578 | 231.49 | 173.6322 | 8.765 | 273.9069
2010-08-12 | 13:38:15 | 11.683 | 111.18 | 29.54 | 6053 | 17.5 |  | 111.1833 | 23.2373 | 231.2863 | 173.4795 | 8.7573 | 273.6659
2010-08-12 | 13:38:25 | 11.85 | 111.573 | 29.5 | 6051 | 17.5 |  | 111.5755 | 23.3193 | 232.1023 | 174.0915 | 8.7882 | 274.6315
2010-08-12 | 13:38:35 | 12.017 | 111.769 | 29.48 | 6049 | 17.5 |  | 111.7723 | 23.3604 | 232.5115 | 174.3985 | 8.8037 | 275.1157
2010-08-12 | 13:38:45 | 12.184 | 111.474 | 29.51 | 6054 | 17.5 |  | 111.4773 | 23.2988 | 231.898 | 173.9383 | 8.7805 | 274.3897
2010-08-12 | 13:38:55 | 12.351 | 110.887 | 29.57 | 6060 | 17.5 |  | 110.8901 | 23.176 | 230.6764 | 173.022 | 8.7342 | 272.9443
2010-08-12 | 13:39:05 | 12.518 | 111.777 | 29.5 | 6030 | 17.4 |  | 111.7806 | 23.3621 | 232.5584 | 174.4337 | 8.8211 | 275.6587
2010-08-12 | 13:39:15 | 12.685 | 112.072 | 29.47 | 6031 | 17.4 |  | 112.0764 | 23.424 | 233.1738 | 174.8952 | 8.8444 | 276.3881
2010-08-12 | 13:39:25 | 12.852 | 112.27 | 29.45 | 6035 | 17.4 |  | 112.2741 | 23.4653 | 233.5851 | 175.2037 | 8.86 | 276.8756
2010-08-12 | 13:39:36 | 13.019 | 111.188 | 29.56 | 6043 | 17.4 |  | 111.1917 | 23.2391 | 231.3333 | 173.5147 | 8.7746 | 274.2065
2010-08-12 | 13:39:46 | 13.186 | 111.875 | 29.49 | 6042 | 17.4 |  | 111.8791 | 23.3827 | 232.7634 | 174.5874 | 8.8289 | 275.9016
2010-08-12 | 13:39:56 | 13.353 | 112.667 | 29.41 | 6042 | 17.4 |  | 112.6707 | 23.5482 | 234.4102 | 175.8226 | 8.8913 | 277.8536
2010-08-12 | 13:40:06 | 13.519 | 111.974 | 29.48 | 6035 | 17.4 |  | 111.9777 | 23.4033 | 232.9685 | 174.7412 | 8.8366 | 276.1447
2010-08-12 | 13:40:16 | 13.687 | 111.883 | 29.51 | 6036 | 17.3 |  | 111.8876 | 23.3845 | 232.8106 | 174.6228 | 8.8463 | 276.4472
2010-08-12 | 13:40:26 | 13.853 | 111.098 | 29.59 | 6046 | 17.3 |  | 111.1027 | 23.2205 | 231.1773 | 173.3977 | 8.7842 | 274.5078
2010-08-12 | 13:40:36 | 14.02 | 111.293 | 29.57 | 6036 | 17.3 |  | 111.2983 | 23.2614 | 231.5844 | 173.703 | 8.7997 | 274.9912
2010-08-12 | 13:40:46 | 14.187 | 111.981 | 29.5 | 6037 | 17.3 |  | 111.9862 | 23.4051 | 233.0157 | 174.7766 | 8.8541 | 276.6908
2010-08-12 | 13:40:56 | 14.354 | 111.587 | 29.54 | 6035 | 17.3 |  | 111.5925 | 23.3228 | 232.1965 | 174.1622 | 8.823 | 275.7181
2010-08-12 | 13:41:06 | 14.521 | 110.902 | 29.61 | 6050 | 17.3 |  | 110.9075 | 23.1797 | 230.771 | 173.093 | 8.7688 | 274.0254
2010-08-12 | 13:41:16 | 14.688 | 111.482 | 29.53 | 6042 | 17.4 |  | 111.4857 | 23.3005 | 231.945 | 173.9735 | 8.7978 | 274.9315
2010-08-12 | 13:41:26 | 14.855 | 111.384 | 29.54 | 6030 | 17.4 |  | 111.3876 | 23.28 | 231.7409 | 173.8204 | 8.7901 | 274.6896
2010-08-12 | 13:41:36 | 15.022 | 111.678 | 29.51 | 6029 | 17.4 |  | 111.6822 | 23.3416 | 232.3537 | 174.2801 | 8.8133 | 275.416
2010-08-12 | 13:41:46 | 15.189 | 111.09 | 29.57 | 6036 | 17.4 |  | 111.094 | 23.2186 | 231.1299 | 173.3621 | 8.7669 | 273.9654
2010-08-12 | 13:41:56 | 15.356 | 111.974 | 29.48 | 6033 | 17.4 |  | 111.9777 | 23.4033 | 232.9685 | 174.7412 | 8.8366 | 276.1447
2010-08-12 | 13:42:06 | 15.522 | 111.777 | 29.5 | 6030 | 17.4 |  | 111.7806 | 23.3621 | 232.5584 | 174.4337 | 8.8211 | 275.6587
2010-08-12 | 13:42:16 | 15.689 | 111.482 | 29.53 | 6031 | 17.4 |  | 111.4857 | 23.3005 | 231.945 | 173.9735 | 8.7978 | 274.9315
2010-08-12 | 13:42:26 | 15.856 | 111.482 | 29.53 | 6041 | 17.4 |  | 111.4857 | 23.3005 | 231.945 | 173.9735 | 8.7978 | 274.9315
2010-08-12 | 13:42:36 | 16.023 | 112.072 | 29.47 | 6030 | 17.4 |  | 112.0764 | 23.424 | 233.1738 | 174.8952 | 8.8444 | 276.3881
2010-08-12 | 13:42:46 | 16.19 | 111.58 | 29.52 | 6019 | 17.4 |  | 111.5839 | 23.321 | 232.1492 | 174.1267 | 8.8056 | 275.1737
2010-08-12 | 13:42:56 | 16.357 | 111.384 | 29.54 | 6031 | 17.4 |  | 111.3876 | 23.28 | 231.7409 | 173.8204 | 8.7901 | 274.6896
2010-08-12 | 13:43:06 | 16.524 | 112.072 | 29.47 | 6030 | 17.4 |  | 112.0764 | 23.424 | 233.1738 | 174.8952 | 8.8444 | 276.3881
2010-08-12 | 13:43:16 | 16.691 | 110.992 | 29.58 | 6032 | 17.4 |  | 110.9963 | 23.1982 | 230.9266 | 173.2097 | 8.7592 | 273.7244
2010-08-12 | 13:43:26 | 16.858 | 111.678 | 29.51 | 6039 | 17.4 |  | 111.6822 | 23.3416 | 232.3537 | 174.2801 | 8.8133 | 275.416
2010-08-12 | 13:43:36 | 17.024 | 111.868 | 29.47 | 6022 | 17.5 |  | 111.8708 | 23.381 | 232.7165 | 174.5522 | 8.8115 | 275.3582
2010-08-12 | 13:43:46 | 17.191 | 111.573 | 29.5 | 6015 | 17.5 |  | 111.5755 | 23.3193 | 232.1023 | 174.0915 | 8.7882 | 274.6315
2010-08-12 | 13:43:56 | 17.358 | 111.671 | 29.49 | 6026 | 17.5 |  | 111.6738 | 23.3398 | 232.3068 | 174.2449 | 8.796 | 274.8735
2010-08-12 | 13:44:06 | 17.525 | 111.671 | 29.49 | 6018 | 17.5 |  | 111.6738 | 23.3398 | 232.3068 | 174.2449 | 8.796 | 274.8735
2010-08-12 | 13:44:16 | 17.692 | 111.376 | 29.52 | 6023 | 17.5 |  | 111.3792 | 23.2783 | 231.6939 | 173.7852 | 8.7727 | 274.1482
2010-08-12 | 13:44:26 | 17.859 | 111.474 | 29.51 | 6024 | 17.5 |  | 111.4773 | 23.2988 | 231.898 | 173.9383 | 8.7805 | 274.3897
2010-08-12 | 13:44:36 | 18.026 | 112.065 | 29.45 | 6027 | 17.5 |  | 112.0681 | 23.4222 | 233.127 | 174.8601 | 8.827 | 275.8439
2010-08-12 | 13:44:46 | 18.193 | 111.966 | 29.46 | 6019 | 17.5 |  | 111.9694 | 23.4016 | 232.9216 | 174.7061 | 8.8192 | 275.6009
2010-08-12 | 13:44:56 | 18.36 | 111.474 | 29.51 | 6024 | 17.5 |  | 111.4773 | 23.2988 | 231.898 | 173.9383 | 8.7805 | 274.3897
2010-08-12 | 13:45:06 | 18.527 | 112.164 | 29.44 | 6024 | 17.5 |  | 112.1669 | 23.4429 | 233.3326 | 175.0143 | 8.8348 | 276.0871
2010-08-12 | 13:45:16 | 18.694 | 112.27 | 29.45 | 6027 | 17.4 |  | 112.2741 | 23.4653 | 233.5851 | 175.2037 | 8.86 | 276.8756
2010-08-12 | 13:45:26 | 18.86 | 111.58 | 29.52 | 6027 | 17.4 |  | 111.5839 | 23.321 | 232.1492 | 174.1267 | 8.8056 | 275.1737
2010-08-12 | 13:45:36 | 19.027 | 111.482 | 29.53 | 6025 | 17.4 |  | 111.4857 | 23.3005 | 231.945 | 173.9735 | 8.7978 | 274.9315
2010-08-12 | 13:45:46 | 19.194 | 111.678 | 29.51 | 6023 | 17.4 |  | 111.6822 | 23.3416 | 232.3537 | 174.2801 | 8.8133 | 275.416
2010-08-12 | 13:45:56 | 19.361 | 112.171 | 29.46 | 6017 | 17.4 |  | 112.1752 | 23.4446 | 233.3793 | 175.0494 | 8.8522 | 276.6317
2010-08-12 | 13:46:06 | 19.528 | 112.072 | 29.47 | 6008 | 17.4 |  | 112.0764 | 23.424 | 233.1738 | 174.8952 | 8.8444 | 276.3881
2010-08-12 | 13:46:16 | 19.695 | 112.468 | 29.43 | 6013 | 17.4 |  | 112.4722 | 23.5067 | 233.9972 | 175.5128 | 8.8757 | 277.3641
2010-08-12 | 13:46:26 | 19.862 | 112.773 | 29.42 | 6018 | 17.3 |  | 112.7784 | 23.5707 | 234.664 | 176.0129 | 8.9167 | 278.648
2010-08-12 | 13:46:36 | 20.029 | 112.575 | 29.44 | 5999 | 17.3 |  | 112.5797 | 23.5292 | 234.2506 | 175.7029 | 8.901 | 278.1572
2010-08-12 | 13:46:46 | 20.196 | 112.08 | 29.49 | 6017 | 17.3 |  | 112.0849 | 23.4257 | 233.221 | 174.9306 | 8.8619 | 276.9345
2010-08-12 | 13:46:55 | 20.346 | 113.672 | 29.33 | 6024 | 17.3 |  | 113.6773 | 23.7586 | 236.5344 | 177.4159 | 8.9878 | 280.8691
2010-08-12 | 13:47:05 | 20.513 | 113.072 | 29.39 | 6019 | 17.3 |  | 113.0771 | 23.6331 | 235.2855 | 176.4792 | 8.9404 | 279.3861
2010-08-12 | 13:47:15 | 20.68 | 113.072 | 29.39 | 6004 | 17.3 |  | 113.0771 | 23.6331 | 235.2855 | 176.4792 | 8.9404 | 279.3861
2010-08-12 | 13:47:25 | 20.847 | 113.572 | 29.34 | 6015 | 17.3 |  | 113.577 | 23.7376 | 236.3258 | 177.2594 | 8.9799 | 280.6213
2010-08-12 | 13:47:35 | 21.014 | 112.873 | 29.41 | 6030 | 17.3 |  | 112.8779 | 23.5915 | 234.8709 | 176.1682 | 8.9246 | 278.8938
2010-08-12 | 13:47:45 | 21.18 | 113.172 | 29.38 | 6023 | 17.3 |  | 113.1769 | 23.654 | 235.4932 | 176.6349 | 8.9482 | 279.6326
2010-08-12 | 13:47:55 | 21.347 | 112.674 | 29.43 | 6000 | 17.3 |  | 112.679 | 23.5499 | 234.4572 | 175.8578 | 8.9089 | 278.4025
2010-08-12 | 13:48:05 | 21.514 | 111.875 | 29.49 | 6022 | 17.4 |  | 111.8791 | 23.3827 | 232.7634 | 174.5874 | 8.8289 | 275.9016
2010-08-12 | 13:48:15 | 21.681 | 112.965 | 29.38 | 6026 | 17.4 |  | 112.9692 | 23.6106 | 235.0312 | 176.2884 | 8.9149 | 278.5898
2010-08-12 | 13:48:25 | 21.848 | 112.567 | 29.42 | 6005 | 17.4 |  | 112.5714 | 23.5274 | 234.2036 | 175.6676 | 8.8835 | 277.6087
2010-08-12 | 13:48:35 | 22.015 | 113.364 | 29.34 | 6010 | 17.4 |  | 113.3686 | 23.694 | 235.8622 | 176.9117 | 8.9464 | 279.5748
2010-08-12 | 13:48:45 | 22.182 | 113.665 | 29.31 | 6017 | 17.4 |  | 113.6692 | 23.7569 | 236.4877 | 177.3809 | 8.9701 | 280.3162
2010-08-12 | 13:48:55 | 22.349 | 112.965 | 29.38 | 6006 | 17.4 |  | 112.9692 | 23.6106 | 235.0312 | 176.2884 | 8.9149 | 278.5898
2010-08-12 | 13:49:05 | 22.516 | 112.965 | 29.38 | 6011 | 17.4 |  | 112.9692 | 23.6106 | 235.0312 | 176.2884 | 8.9149 | 278.5898
2010-08-12 | 13:49:15 | 22.683 | 113.065 | 29.37 | 6009 | 17.4 |  | 113.0689 | 23.6314 | 235.2387 | 176.444 | 8.9227 | 278.8356
2010-08-12 | 13:49:25 | 22.85 | 112.766 | 29.4 | 6010 | 17.4 |  | 112.7701 | 23.5689 | 234.617 | 175.9777 | 8.8992 | 278.0988
2010-08-12 | 13:49:35 | 23.017 | 112.766 | 29.4 | 6010 | 17.4 |  | 112.7701 | 23.5689 | 234.617 | 175.9777 | 8.8992 | 278.0988
2010-08-12 | 13:49:45 | 23.184 | 112.072 | 29.47 | 6005 | 17.4 |  | 112.0764 | 23.424 | 233.1738 | 174.8952 | 8.8444 | 276.3881
2010-08-12 | 13:49:55 | 23.35 | 113.165 | 29.36 | 6014 | 17.4 |  | 113.1687 | 23.6523 | 235.4463 | 176.5997 | 8.9306 | 279.0818
2010-08-12 | 13:50:05 | 23.517 | 112.958 | 29.36 | 5997 | 17.5 |  | 112.9611 | 23.6089 | 234.9847 | 176.2535 | 8.8973 | 278.042
2010-08-12 | 13:50:15 | 23.684 | 112.065 | 29.45 | 6006 | 17.5 |  | 112.0681 | 23.4222 | 233.127 | 174.8601 | 8.827 | 275.8439
2010-08-12 | 13:50:25 | 23.851 | 112.858 | 29.37 | 5996 | 17.5 |  | 112.8615 | 23.5881 | 234.7774 | 176.098 | 8.8895 | 277.7967
2010-08-12 | 13:50:35 | 24.018 | 112.858 | 29.37 | 5994 | 17.5 |  | 112.8615 | 23.5881 | 234.7774 | 176.098 | 8.8895 | 277.7967
2010-08-12 | 13:50:45 | 24.185 | 112.56 | 29.4 | 5994 | 17.5 |  | 112.5632 | 23.5257 | 234.1569 | 175.6326 | 8.866 | 277.0626
2010-08-12 | 13:50:56 | 24.352 | 111.376 | 29.52 | 5998 | 17.5 |  | 111.3792 | 23.2783 | 231.6939 | 173.7852 | 8.7727 | 274.1482
2010-08-12 | 13:51:05 | 24.518 | 113.859 | 29.27 | 6003 | 17.5 |  | 113.8624 | 23.7972 | 236.8595 | 177.6597 | 8.9683 | 280.2603
2010-08-12 | 13:51:16 | 24.685 | 112.659 | 29.39 | 5989 | 17.5 |  | 112.6625 | 23.5465 | 234.3635 | 175.7876 | 8.8738 | 277.307
2010-08-12 | 13:51:26 | 24.852 | 112.263 | 29.43 | 5987 | 17.5 |  | 112.2659 | 23.4636 | 233.5383 | 175.1686 | 8.8426 | 276.3306
2010-08-12 | 13:51:36 | 25.019 | 113.558 | 29.3 | 5987 | 17.5 |  | 113.5611 | 23.7343 | 236.2326 | 177.1895 | 8.9446 | 279.5186
2010-08-12 | 13:51:46 | 25.186 | 112.759 | 29.38 | 5990 | 17.5 |  | 112.762 | 23.5673 | 234.5704 | 175.9427 | 8.8817 | 277.5518
2010-08-12 | 13:51:56 | 25.353 | 112.355 | 29.4 | 5996 | 17.6 |  | 112.3569 | 23.4826 | 233.6979 | 175.2883 | 8.833 | 276.0312
2010-08-12 | 13:52:06 | 25.52 | 112.951 | 29.34 | 5988 | 17.6 |  | 112.9533 | 23.6072 | 234.9384 | 176.2188 | 8.8799 | 277.4964
2010-08-12 | 13:52:16 | 25.687 | 112.454 | 29.39 | 5980 | 17.6 |  | 112.4561 | 23.5033 | 233.9041 | 175.443 | 8.8408 | 276.2748
2010-08-12 | 13:52:26 | 25.854 | 114.054 | 29.23 | 5980 | 17.6 |  | 114.0563 | 23.8378 | 237.2325 | 177.9395 | 8.9666 | 280.2061
2010-08-12 | 13:52:36 | 26.021 | 112.852 | 29.35 | 5978 | 17.6 |  | 112.8537 | 23.5864 | 234.7311 | 176.0633 | 8.8721 | 277.2516
2010-08-12 | 13:52:46 | 26.188 | 113.752 | 29.26 | 5958 | 17.6 |  | 113.7543 | 23.7746 | 236.6043 | 177.4683 | 8.9429 | 279.4641
2010-08-12 | 13:52:56 | 26.355 | 113.151 | 29.32 | 5959 | 17.6 |  | 113.153 | 23.649 | 235.3536 | 176.5302 | 8.8956 | 277.9868
2010-08-12 | 13:53:06 | 26.521 | 113.953 | 29.24 | 5954 | 17.6 |  | 113.9555 | 23.8167 | 237.0229 | 177.7823 | 8.9587 | 279.9585
2010-08-12 | 13:53:16 | 26.688 | 113.953 | 29.24 | 5957 | 17.6 |  | 113.9555 | 23.8167 | 237.0229 | 177.7823 | 8.9587 | 279.9585
2010-08-12 | 13:53:26 | 26.855 | 114.256 | 29.21 | 5947 | 17.6 |  | 114.2582 | 23.88 | 237.6524 | 178.2545 | 8.9825 | 280.7021
2010-08-12 | 13:53:36 | 27.022 | 114.566 | 29.2 | 5954 | 17.5 |  | 114.5691 | 23.945 | 238.3297 | 178.7625 | 9.024 | 281.9999
2010-08-12 | 13:53:46 | 27.189 | 115.995 | 29.06 | 5954 | 17.5 |  | 115.998 | 24.2436 | 241.302 | 180.9919 | 9.1365 | 285.5169
2010-08-12 | 13:53:56 | 27.356 | 114.667 | 29.19 | 5939 | 17.5 |  | 114.6705 | 23.9661 | 238.5406 | 178.9206 | 9.032 | 282.2495
2010-08-12 | 13:54:06 | 27.523 | 115.278 | 29.13 | 5946 | 17.5 |  | 115.281 | 24.0937 | 239.8105 | 179.8732 | 9.0801 | 283.7521
2010-08-12 | 13:54:16 | 27.69 | 114.263 | 29.23 | 5945 | 17.5 |  | 114.2656 | 23.8815 | 237.6983 | 178.2889 | 9.0001 | 281.2529
2010-08-12 | 13:54:26 | 27.857 | 114.775 | 29.2 | 5935 | 17.4 |  | 114.7796 | 23.9889 | 238.7977 | 179.1135 | 9.0577 | 283.0543
2010-08-12 | 13:54:36 | 28.024 | 115.386 | 29.14 | 5936 | 17.4 |  | 115.3905 | 24.1166 | 240.0688 | 180.0669 | 9.1059 | 284.5609
2010-08-12 | 13:54:46 | 28.191 | 116.413 | 29.04 | 5933 | 17.4 |  | 116.4171 | 24.3312 | 242.2047 | 181.6689 | 9.187 | 287.0926
2010-08-12 | 13:54:56 | 28.357 | 115.898 | 29.09 | 5925 | 17.4 |  | 115.9025 | 24.2236 | 241.134 | 180.8658 | 9.1464 | 285.8235
2010-08-12 | 13:55:06 | 28.524 | 116.31 | 29.05 | 5913 | 17.4 |  | 116.314 | 24.3096 | 241.9901 | 181.508 | 9.1788 | 286.8383
2010-08-12 | 13:55:16 | 28.691 | 116.104 | 29.07 | 5930 | 17.4 |  | 116.108 | 24.2666 | 241.5616 | 181.1866 | 9.1626 | 286.3304
2010-08-12 | 13:55:26 | 28.858 | 115.693 | 29.11 | 5916 | 17.4 |  | 115.6974 | 24.1807 | 240.7072 | 180.5458 | 9.1302 | 285.3177
2010-08-12 | 13:55:36 | 29.025 | 116.104 | 29.07 | 5921 | 17.4 |  | 116.108 | 24.2666 | 241.5616 | 181.1866 | 9.1626 | 286.3304
2010-08-12 | 13:55:46 | 29.192 | 116.516 | 29.03 | 5907 | 17.4 |  | 116.5204 | 24.3528 | 242.4195 | 181.8301 | 9.1951 | 287.3473
2010-08-12 | 13:55:56 | 29.359 | 115.488 | 29.13 | 5906 | 17.4 |  | 115.4927 | 24.138 | 240.2814 | 180.2263 | 9.114 | 284.8129
2010-08-12 | 13:56:06 | 29.526 | 116.413 | 29.04 | 5892 | 17.4 |  | 116.4171 | 24.3312 | 242.2047 | 181.6689 | 9.187 | 287.0926
2010-08-12 | 13:56:16 | 29.693 | 116.104 | 29.07 | 5914 | 17.4 |  | 116.108 | 24.2666 | 241.5616 | 181.1866 | 9.1626 | 286.3304
2010-08-12 | 13:56:26 | 29.86 | 117.034 | 28.98 | 5897 | 17.4 |  | 117.0382 | 24.461 | 243.4968 | 182.6381 | 9.236 | 288.6243
2010-08-12 | 13:56:36 | 30.027 | 116.723 | 29.01 | 5891 | 17.4 |  | 116.7272 | 24.396 | 242.8497 | 182.1528 | 9.2114 | 287.8573
2010-08-12 | 13:56:46 | 30.194 | 116.516 | 29.03 | 5891 | 17.4 |  | 116.5204 | 24.3528 | 242.4195 | 181.8301 | 9.1951 | 287.3473
2010-08-12 | 13:56:56 | 30.36 | 117.236 | 28.94 | 5884 | 17.5 |  | 117.2392 | 24.503 | 243.8841 | 182.9286 | 9.2343 | 288.5721
2010-08-12 | 13:57:06 | 30.527 | 117.444 | 28.92 | 5881 | 17.5 |  | 117.4476 | 24.5466 | 244.3176 | 183.2537 | 9.2507 | 289.085
2010-08-12 | 13:57:16 | 30.694 | 117.132 | 28.95 | 5867 | 17.5 |  | 117.1352 | 24.4813 | 243.6677 | 182.7663 | 9.2261 | 288.316
2010-08-12 | 13:57:26 | 30.861 | 117.028 | 28.96 | 5859 | 17.5 |  | 117.0313 | 24.4595 | 243.4515 | 182.6042 | 9.2179 | 288.0602
2010-08-12 | 13:57:36 | 31.028 | 118.072 | 28.86 | 5859 | 17.5 |  | 118.0753 | 24.6777 | 245.6234 | 184.2332 | 9.3002 | 290.6301
2010-08-12 | 13:57:46 | 31.195 | 118.387 | 28.83 | 5846 | 17.5 |  | 118.3907 | 24.7437 | 246.2793 | 184.7252 | 9.325 | 291.4062
2010-08-12 | 13:57:56 | 31.362 | 118.915 | 28.78 | 5854 | 17.5 |  | 118.9184 | 24.854 | 247.3772 | 185.5487 | 9.3666 | 292.7052
2010-08-12 | 13:58:06 | 31.529 | 116.407 | 29.02 | 5854 | 17.5 |  | 116.41 | 24.3297 | 242.1591 | 181.6348 | 9.169 | 286.5311
2010-08-12 | 13:58:16 | 31.696 | 118.282 | 28.84 | 5852 | 17.5 |  | 118.2855 | 24.7217 | 246.0605 | 184.561 | 9.3167 | 291.1472
2010-08-12 | 13:58:26 | 31.862 | 119.233 | 28.75 | 5846 | 17.5 |  | 119.2364 | 24.9204 | 248.0386 | 186.0448 | 9.3916 | 293.4879
2010-08-12 | 13:58:35 | 32.013 | 118.809 | 28.79 | 5835 | 17.5 |  | 118.8127 | 24.8318 | 247.1572 | 185.3836 | 9.3582 | 292.4449
2010-08-12 | 13:58:45 | 32.18 | 119.872 | 28.69 | 5838 | 17.5 |  | 119.8753 | 25.0539 | 249.3677 | 187.0417 | 9.4419 | 295.0605
2010-08-12 | 13:58:55 | 32.347 | 118.598 | 28.81 | 5833 | 17.5 |  | 118.6015 | 24.7877 | 246.7178 | 185.0541 | 9.3416 | 291.925
2010-08-12 | 13:59:05 | 32.513 | 119.228 | 28.73 | 5828 | 17.6 |  | 119.2304 | 24.9192 | 247.9945 | 186.0117 | 9.3734 | 292.9175
2010-08-12 | 13:59:15 | 32.68 | 119.547 | 28.7 | 5817 | 17.6 |  | 119.5495 | 24.9858 | 248.6581 | 186.5094 | 9.3984 | 293.7014
2010-08-12 | 13:59:25 | 32.847 | 119.867 | 28.67 | 5815 | 17.6 |  | 119.8695 | 25.0527 | 249.3238 | 187.0088 | 9.4236 | 294.4877
2010-08-12 | 13:59:35 | 33.014 | 120.081 | 28.65 | 5796 | 17.6 |  | 120.0835 | 25.0974 | 249.7688 | 187.3425 | 9.4404 | 295.0133
2010-08-12 | 13:59:45 | 33.181 | 119.654 | 28.69 | 5787 | 17.6 |  | 119.656 | 25.0081 | 248.8798 | 186.6757 | 9.4068 | 293.9632
2010-08-12 | 13:59:55 | 33.348 | 120.51 | 28.61 | 5783 | 17.6 |  | 120.5127 | 25.1871 | 250.6615 | 188.0121 | 9.4742 | 296.0677
2010-08-12 | 14:00:05 | 33.515 | 119.547 | 28.7 | 5798 | 17.6 |  | 119.5495 | 24.9858 | 248.6581 | 186.5094 | 9.3984 | 293.7014
2010-08-12 | 14:00:15 | 33.682 | 120.296 | 28.63 | 5783 | 17.6 |  | 120.2978 | 25.1422 | 250.2147 | 187.677 | 9.4573 | 295.5399
2010-08-12 | 14:00:25 | 33.848 | 121.05 | 28.56 | 5771 | 17.6 |  | 121.0517 | 25.2998 | 251.7827 | 188.8531 | 9.5165 | 297.392
2010-08-12 | 14:00:35 | 34.016 | 120.51 | 28.61 | 5760 | 17.6 |  | 120.5127 | 25.1871 | 250.6615 | 188.0121 | 9.4742 | 296.0677
2010-08-12 | 14:00:45 | 34.182 | 120.834 | 28.58 | 5770 | 17.6 |  | 120.8358 | 25.2547 | 251.3335 | 188.5162 | 9.4996 | 296.8615
2010-08-12 | 14:00:55 | 34.349 | 121.809 | 28.49 | 5746 | 17.6 |  | 121.8111 | 25.4585 | 253.3623 | 190.0379 | 9.5762 | 299.2577
2010-08-12 | 14:01:05 | 34.516 | 122.14 | 28.48 | 5743 | 17.5 |  | 122.1434 | 25.528 | 254.0858 | 190.5806 | 9.6206 | 300.6432
2010-08-12 | 14:01:15 | 34.683 | 122.359 | 28.46 | 5747 | 17.5 |  | 122.362 | 25.5737 | 254.5406 | 190.9217 | 9.6378 | 301.1813
2010-08-12 | 14:01:25 | 34.85 | 121.704 | 28.52 | 5733 | 17.5
2010-08-12 | 14:01:35 | 35.017 | 122.797 | 28.42 | 5733 | 17.5
2010-08-12 | 14:01:45 | 35.184 | 124.236 | 28.29 | 5734 | 17.5
2010-08-12 | 14:01:55 | 35.351 | 124.459 | 28.27 | 5731 | 17.5
2010-08-12 | 14:02:05 | 35.518 | 123.791 | 28.33 | 5714 | 17.5
2010-08-12 | 14:02:16 | 35.686 | 123.902 | 28.32 | 5721 | 17.5
2010-08-12 | 14:02:25 | 35.852 | 122.578 | 28.44 | 5705 | 17.5
2010-08-12 | 14:02:36 | 36.019 | 123.569 | 28.35 | 5703 | 17.5
2010-08-12 | 14:02:46 | 36.186 | 123.906 | 28.34 | 5694 | 17.4
2010-08-12 | 14:02:56 | 36.353 | 125.131 | 28.21 | 5701 | 17.5
2010-08-12 | 14:03:06 | 36.52 | 124.57 | 28.26 | 5685 | 17.5
2010-08-12 | 14:03:16 | 36.687 | 125.018 | 28.22 | 5676 | 17.5
2010-08-12 | 14:03:26 | 36.854 | 124.57 | 28.26 | 5676 | 17.5
2010-08-12 | 14:03:36 | 37.021 | 125.018 | 28.22 | 5665 | 17.5
2010-08-12 | 14:03:46 | 37.188 | 125.243 | 28.2 | 5664 | 17.5
2010-08-12 | 14:03:56 | 37.355 | 125.243 | 28.2 | 5651 | 17.5
2010-08-12 | 14:04:06 | 37.521 | 125.694 | 28.16 | 5657 | 17.5
2010-08-12 | 14:04:16 | 37.688 | 126.374 | 28.1 | 5652 | 17.5
2010-08-12 | 14:04:26 | 37.855 | 125.694 | 28.16 | 5648 | 17.5
2010-08-12 | 14:04:36 | 38.022 | 125.356 | 28.19 | 5638 | 17.5
2010-08-12 | 14:04:46 | 38.189 | 126.033 | 28.13 | 5639 | 17.5
2010-08-12 | 14:04:56 | 38.356 | 127.172 | 28.03 | 5619 | 17.5
2010-08-12 | 14:05:06 | 38.523 | 127.631 | 27.99 | 5619 | 17.5
2010-08-12 | 14:05:16 | 38.69 | 127.4 | 27.99 | 5610 | 17.6
2010-08-12 | 14:05:26 | 38.857 | 127.285 | 28 | 5613 | 17.6
2010-08-12 | 14:05:36 | 39.024 | 128.786 | 27.87 | 5609 | 17.6
2010-08-12 | 14:05:46 | 39.191 | 127.745 | 27.96 | 5588 | 17.6
2010-08-12 | 14:05:56 | 39.357 | 127.629 | 27.97 | 5602 | 17.6
2010-08-12 | 14:06:06 | 39.525 | 129.603 | 27.8 | 5582 | 17.6
2010-08-12 | 14:06:16 | 39.691 | 128.786 | 27.87 | 5575 | 17.6
2010-08-12 | 14:06:26 | 39.858 | 128.091 | 27.93 | 5569 | 17.6
2010-08-12 | 14:06:36 | 40.025 | 129.486 | 27.81 | 5569 | 17.6
2010-08-12 | 14:06:46 | 40.192 | 128.67 | 27.88 | 5556 | 17.6
2010-08-12 | 14:06:56 | 40.359 | 130.191 | 27.75 | 5556 | 17.6
2010-08-12 | 14:07:06 | 40.526 | 129.955 | 27.77 | 5557 | 17.6
2010-08-12 | 14:07:16 | 40.693 | 129.838 | 27.78 | 5549 | 17.6
2010-08-12 | 14:07:26 | 40.86 | 131.018 | 27.68 | 5529 | 17.6
2010-08-12 | 14:07:36 | 41.027 | 130.426 | 27.73 | 5532 | 17.6
2010-08-12 | 14:07:46 | 41.194 | 130.308 | 27.74 | 5519 | 17.6
2010-08-12 | 14:07:56 | 41.36 | 130.9 | 27.69 | 5509 | 17.6
2010-08-12 | 14:08:06 | 41.527 | 130.426 | 27.73 | 5512 | 17.6
2010-08-12 | 14:08:16 | 41.694 | 131.137 | 27.67 | 5506 | 17.6
2010-08-12 | 14:08:26 | 41.861 | 130.663 | 27.71 | 5499 | 17.6
2010-08-12 | 14:08:36 | 42.028 | 131.375 | 27.65 | 5488 | 17.6
2010-08-12 | 14:08:46 | 42.195 | 132.452 | 27.56 | 5494 | 17.6
2010-08-12 | 14:08:56 | 42.362 | 131.256 | 27.66 | 5488 | 17.6
2010-08-12 | 14:09:06 | 42.529 | 133.296 | 27.49 | 5481 | 17.6
2010-08-12 | 14:09:16 | 42.696 | 131.256 | 27.66 | 5474 | 17.6
2010-08-12 | 14:09:26 | 42.863 | 132.934 | 27.54 | 5468 | 17.5
2010-08-12 | 14:09:36 | 43.029 | 133.539 | 27.49 | 5467 | 17.5
2010-08-12 | 14:09:46 | 43.196 | 132.813 | 27.55 | 5461 | 17.5
2010-08-12 | 14:09:56 | 43.363 | 133.782 | 27.47 | 5457 | 17.5
2010-08-12 | 14:10:05 | 43.514 | 133.417 | 27.5 | 5444 | 17.5
2010-08-12 | 14:10:15 | 43.68 | 133.054 | 27.53 | 5443 | 17.5
2010-08-12 | 14:10:25 | 43.848 | 134.392 | 27.42 | 5449 | 17.5
2010-08-12 | 14:10:35 | 44.014 | 133.903 | 27.46 | 5430 | 17.5
2010-08-12 | 14:10:45 | 44.181 | 135.128 | 27.36 | 5435 | 17.5
2010-08-12 | 14:10:55 | 44.348 | 135.128 | 27.36 | 5419 | 17.5
2010-08-12 | 14:11:05 | 44.515 | 134.882 | 27.38 | 5424 | 17.5
2010-08-12 | 14:11:15 | 44.682 | 135.251 | 27.35 | 5413 | 17.5
2010-08-12 | 14:11:25 | 44.849 | 135.621 | 27.32 | 5411 | 17.5
2010-08-12 | 14:11:35 | 45.016 | 136.739 | 27.23 | 5407 | 17.5
2010-08-12 | 14:11:45 | 45.182 | 136.241 | 27.27 | 5394 | 17.5
2010-08-12 | 14:11:55 | 45.349 | 137.116 | 27.18 | 5397 | 17.6
2010-08-12 | 14:12:05 | 45.516 | 136.741 | 27.21 | 5388 | 17.6
2010-08-12 | 14:12:15 | 45.683 | 136.242 | 27.25 | 5380 | 17.6
2010-08-12 | 14:12:25 | 45.85 | 136.491 | 27.23 | 5387 | 17.6
2010-08-12 | 14:12:35 | 46.017 | 136.991 | 27.19 | 5383 | 17.6
2010-08-12 | 14:12:45 | 46.184 | 138.249 | 27.09 | 5367 | 17.6
2010-08-12 | 14:12:55 | 46.351 | 137.367 | 27.16 | 5363 | 17.6
2010-08-12 | 14:13:05 | 46.518 | 139.138 | 27.02 | 5362 | 17.6
2010-08-12 | 14:13:15 | 46.685 | 137.492 | 27.15 | 5353 | 17.6
2010-08-12 | 14:13:25 | 46.852 | 137.87 | 27.12 | 5343 | 17.6
2010-08-12 | 14:13:35 | 47.018 | 137.87 | 27.12 | 5330 | 17.6
2010-08-12 | 14:13:46 | 47.185 | 137.744 | 27.13 | 5346 | 17.6
2010-08-12 | 14:13:56 | 47.352 | 140.035 | 26.95 | 5342 | 17.6
2010-08-12 | 14:14:06 | 47.519 | 140.679 | 26.9 | 5331 | 17.6
2010-08-12 | 14:14:16 | 47.686 | 139.778 | 26.97 | 5317 | 17.6
2010-08-12 | 14:14:26 | 47.853 | 139.138 | 27.02 | 5319 | 17.6
2010-08-12 | 14:14:36 | 48.02 | 138.757 | 27.05 | 5323 | 17.6
2010-08-12 | 14:14:46 | 48.187 | 138.759 | 27.03 | 5307 | 17.7
2010-08-12 | 14:14:56 | 48.354 | 140.553 | 26.89 | 5304 | 17.7
2010-08-12 | 14:15:06 | 48.521 | 140.553 | 26.89 | 5279 | 17.7
2010-08-12 | 14:15:16 | 48.688 | 141.331 | 26.83 | 5298 | 17.7
2010-08-12 | 14:15:26 | 48.855 | 140.424 | 26.9 | 5273 | 17.7
2010-08-12 | 14:15:36 | 49.021 | 141.331 | 26.83 | 5287 | 17.7
2010-08-12 | 14:15:46 | 49.188 | 141.461 | 26.82 | 5273 | 17.7
2010-08-12 | 14:15:56 | 49.355 | 142.244 | 26.76 | 5256 | 17.7
2010-08-12 | 14:16:06 | 49.522 | 141.591 | 26.81 | 5262 | 17.7
2010-08-12 | 14:16:16 | 49.689 | 142.502 | 26.76 | 5256 | 17.6
2010-08-12 | 14:16:26 | 49.856 | 142.24 | 26.78 | 5250 | 17.6
2010-08-12 | 14:16:36 | 50.023 | 141.587 | 26.83 | 5241 | 17.6
2010-08-12 | 14:16:46 | 50.19 | 141.848 | 26.81 | 5241 | 17.6
2010-08-12 | 14:16:56 | 50.357 | 143.292 | 26.7 | 5236 | 17.6
2010-08-12 | 14:17:06 | 50.524 | 141.587 | 26.83 | 5236 | 17.6
2010-08-12 | 14:17:16 | 50.691 | 142.633 | 26.75 | 5230 | 17.6