## convers.+compens.
file name: PSt1_Eq4_Sal |  |  | 2008-03-29
author: Dr. Christian Huber
programm for temperature compensated oxygen calculation of membrane PSt1
Fit of the calibration curve is optimized between 0 and 250 % air-saturation (0 and 388 Torr)
enter your values here
input of F0,T0 | 58.87 | phase angle of cal 0
input of F100,T100 | 30.69 | phase angle of cal 100 (water-vapor saturated air)
input of Fm,Tm | 15 | measured phase angle
input of T0 | 17.7 | temperature of cal 0
input of T100 | 17.8 | temperatur of cal100
input of Tm | 17.7 | temperature at measurement
air pressure patm | 1013 | air pressure
Salinity | 32.2 | Salinity
Results of calculation + compensation
[%] air saturation | 493.4732 | %
[%] O2 | 103.1359 | %
pO2 [hPa] | 1026.2727 | hPa
pO2 [Torr] | 769.7699 | Torr
cO2 [mg/L] | 38.7213 | mg/L
cO2 [ppm] | 38.7213 | ppm
cO2 [µmol/L] | 1210.0413 | µmol/L
Internal calculated parameter (DO NOT CHANGE !!)
f1 | 0.801
DF/K | -0.08
DKSV/K | 0.000383
m | 22.9
tan(F0,T100) | 1.6552
tan(F0,Tm) | 1.6558
tan(F100,T100) | 0.59352
tan(Fm,Tm) | 0.26795
KSV,T100 | 0.033331
KSV,Tm | 0.033293
A | 156.5815
B | 14.0252
C | -0.64143
a | 7.8328e-06
b | -0.0021668
c | -0.83817
Chlorinity | 17.8227

## Recalculate Fibox values
file name: PSt1_Eq4_Sal |  |  |  | 2008-03-29
author: Dr. Christian Huber
programm for temperature compensated oxygen calculation of membrane PSt1
Fit of the calibration curve is optimized between 0 and 250 % air-saturation (0 and 388 Torr)
Insert your calibration values
cal0 | 58.87 | ° | T0 | 17.7 | °C
cal100 | 30.69 | ° | T100 | 17.8 | °C
air pressure | 1013 | mbar
Salinity | 32.2 | °/°°
Internal calculated parameter (DO NOT CHANGE !!)
f1 | 0.801 | tan(F0,T100) | 1.6552 | A | 156.5815 | KSV,T100 | 0.033331 | a | 1.7072e-05
DF/K | -0.08 | tan(F0,Tm) | 1.6563 | B | 14.0252 | KSV,Tm | 0.033254 | b | 0.0044886 |  |  |  |  | INSERT YOUR RESPECTIVE DATA
DKSV/K | 0.000383 | tan(F100,T100) | 0.59352 | C | -0.64143 |  |  | c | -0.64648 |  |  |  |  | chamber volume [L] | FW alga [g] | mg Chl a Thallus -1
m | 22.9 | tan(Fm,Tm) | 0.58552 |  |  |  |  | Chlorinity | 17.8227 |  |  |  |  | 0.026 | 0.0827 | 0.038671
Insert the data you get from the Ascii File recorded with the OxyView TX3-V5.31.exe software | [7 empty] | The oxygen contents are calculated automatically from the raw data
date | time/hh:mm:ss | logtime/min | oxygen/% airsatur. | phase/° | amp | temp/°C |  | % air-sat. | % oxygen | pO2 (hPa) | pO2 (Torr) | cO2 [mg/L] | cO2 [µmol/L] | cO2 [µmol/L 10 min] | cO2 [µmol/L h] | cO2 [µmol/ h] | cO2 [µmol/ h gFW] | cO2 [µmol/ h mg Chl a]
2010-08-12 | 17:31:49 | 0 | 308.08 | 30.35 | 6241 | 17.6 |  | 103.3803 | 21.6065 | 215.0268 | 161.2838 | 8.1273 | 253.978 | 22.428 | 134.568 | 3.4988 | 42.3067 | 90.4741
2010-08-12 | 17:32:00 | 0.184 | 307.276 | 30.38 | 6235 | 17.6 |  | 103.1105 | 21.5501 | 214.4656 | 160.8629 | 8.1061 | 253.3151
2010-08-12 | 17:32:10 | 0.351 | 305.941 | 30.43 | 6235 | 17.6 |  | 102.6625 | 21.4565 | 213.5339 | 160.1641 | 8.0709 | 252.2146 | change data input according to column N
2010-08-12 | 17:32:20 | 0.517 | 306.474 | 30.41 | 6252 | 17.6 |  | 102.8414 | 21.4939 | 213.9061 | 160.4432 | 8.0849 | 252.6542
2010-08-12 | 17:32:30 | 0.684 | 307.008 | 30.39 | 6250 | 17.6 |  | 103.0207 | 21.5313 | 214.2789 | 160.7228 | 8.099 | 253.0946
2010-08-12 | 17:32:40 | 0.851 | 306.081 | 30.38 | 6248 | 17.7 |  | 102.9211 | 21.5105 | 214.0443 | 160.5468 | 8.0759 | 252.3717
2010-08-12 | 17:32:50 | 1.018 | 303.692 | 30.47 | 6236 | 17.7 |  | 102.1178 | 21.3426 | 212.3737 | 159.2938 | 8.0129 | 250.4021
2010-08-12 | 17:33:00 | 1.185 | 305.815 | 30.39 | 6240 | 17.7 |  | 102.8315 | 21.4918 | 213.8579 | 160.4071 | 8.0689 | 252.152
2010-08-12 | 17:33:10 | 1.352 | 304.486 | 30.44 | 6238 | 17.7 |  | 102.3848 | 21.3984 | 212.9289 | 159.7103 | 8.0338 | 251.0567
2010-08-12 | 17:33:20 | 1.519 | 306.348 | 30.37 | 6249 | 17.7 |  | 103.0108 | 21.5293 | 214.2308 | 160.6868 | 8.0829 | 252.5917
2010-08-12 | 17:33:30 | 1.686 | 302.637 | 30.51 | 6251 | 17.7 |  | 101.7631 | 21.2685 | 211.636 | 158.7405 | 7.985 | 249.5323
2010-08-12 | 17:33:41 | 1.853 | 308.222 | 30.3 | 6246 | 17.7 |  | 103.6411 | 21.661 | 215.5417 | 161.67 | 8.1324 | 254.1373
2010-08-12 | 17:33:51 | 2.02 | 305.282 | 30.41 | 6239 | 17.7 |  | 102.6526 | 21.4544 | 213.4858 | 160.1279 | 8.0548 | 251.7132
2010-08-12 | 17:34:01 | 2.186 | 304.221 | 30.45 | 6241 | 17.7 |  | 102.2957 | 21.3798 | 212.7437 | 159.5713 | 8.0268 | 250.8383
2010-08-12 | 17:34:11 | 2.353 | 303.428 | 30.48 | 6245 | 17.7 |  | 102.029 | 21.3241 | 212.189 | 159.1553 | 8.0059 | 250.1843
2010-08-12 | 17:34:21 | 2.52 | 304.221 | 30.45 | 6248 | 17.7 |  | 102.2957 | 21.3798 | 212.7437 | 159.5713 | 8.0268 | 250.8383
2010-08-12 | 17:34:31 | 2.687 | 306.348 | 30.37 | 6239 | 17.7 |  | 103.0108 | 21.5293 | 214.2308 | 160.6868 | 8.0829 | 252.5917
2010-08-12 | 17:34:41 | 2.854 | 303.957 | 30.46 | 6253 | 17.7 |  | 102.2067 | 21.3612 | 212.5586 | 159.4325 | 8.0198 | 250.6201
2010-08-12 | 17:34:51 | 3.021 | 304.751 | 30.43 | 6239 | 17.7 |  | 102.474 | 21.4171 | 213.1144 | 159.8494 | 8.0408 | 251.2753
2010-08-12 | 17:35:01 | 3.188 | 306.882 | 30.35 | 6245 | 17.7 |  | 103.1904 | 21.5668 | 214.6044 | 160.967 | 8.097 | 253.0322
2010-08-12 | 17:35:11 | 3.355 | 304.751 | 30.43 | 6250 | 17.7 |  | 102.474 | 21.4171 | 213.1144 | 159.8494 | 8.0408 | 251.2753
2010-08-12 | 17:35:21 | 3.522 | 305.941 | 30.43 | 6247 | 17.6 |  | 102.6625 | 21.4565 | 213.5339 | 160.1641 | 8.0709 | 252.2146
2010-08-12 | 17:35:31 | 3.689 | 307.008 | 30.39 | 6251 | 17.6 |  | 103.0207 | 21.5313 | 214.2789 | 160.7228 | 8.099 | 253.0946 |  | regression formula | time
2010-08-12 | 17:35:41 | 3.856 | 305.941 | 30.43 | 6244 | 17.6 |  | 102.6625 | 21.4565 | 213.5339 | 160.1641 | 8.0709 | 252.2146 | value for T=26 min. | 295.834 | T11
2010-08-12 | 17:35:51 | 4.023 | 306.474 | 30.41 | 6249 | 17.6 |  | 102.8414 | 21.4939 | 213.9061 | 160.4432 | 8.0849 | 252.6542 | value for T=1 min. | 273.406 | T1
2010-08-12 | 17:36:01 | 4.19 | 302.238 | 30.57 | 6254 | 17.6 |  | 101.42 | 21.1968 | 210.9495 | 158.2256 | 7.9732 | 249.1621 | difference between T25 and T1 | 22.428 | 10minutes
2010-08-12 | 17:36:11 | 4.356 | 304.083 | 30.5 | 6260 | 17.6 |  | 102.0391 | 21.3262 | 212.2373 | 159.1915 | 8.0219 | 250.6831 | calculate from regression curve values for 10 minutes of photosynthesis or respiration
2010-08-12 | 17:36:21 | 4.523 | 305.941 | 30.43 | 6252 | 17.6 |  | 102.6625 | 21.4565 | 213.5339 | 160.1641 | 8.0709 | 252.2146
2010-08-12 | 17:36:31 | 4.69 | 303.291 | 30.53 | 6256 | 17.6 |  | 101.7733 | 21.2706 | 211.6843 | 158.7767 | 8.001 | 250.0299
2010-08-12 | 17:36:41 | 4.857 | 306.07 | 30.47 | 6253 | 17.5 |  | 102.4942 | 21.4213 | 213.211 | 159.9219 | 8.0729 | 252.2787
2010-08-12 | 17:36:51 | 5.024 | 305.804 | 30.48 | 6259 | 17.5 |  | 102.405 | 21.4027 | 213.0256 | 159.7828 | 8.0659 | 252.0593
2010-08-12 | 17:37:01 | 5.191 | 305.538 | 30.49 | 6245 | 17.5 |  | 102.316 | 21.384 | 212.8404 | 159.6439 | 8.0589 | 251.8401
2010-08-12 | 17:37:11 | 5.358 | 306.336 | 30.46 | 6255 | 17.5 |  | 102.5834 | 21.4399 | 213.3966 | 160.0611 | 8.0799 | 252.4983
2010-08-12 | 17:37:21 | 5.525 | 306.603 | 30.45 | 6264 | 17.5 |  | 102.6727 | 21.4586 | 213.5824 | 160.2004 | 8.087 | 252.7181
2010-08-12 | 17:37:31 | 5.691 | 304.742 | 30.52 | 6268 | 17.5 |  | 102.0494 | 21.3283 | 212.2858 | 159.2279 | 8.0379 | 251.1839
2010-08-12 | 17:37:41 | 5.858 | 305.804 | 30.48 | 6242 | 17.5 |  | 102.405 | 21.4027 | 213.0256 | 159.7828 | 8.0659 | 252.0593
2010-08-12 | 17:37:51 | 6.025 | 307.137 | 30.43 | 6260 | 17.5 |  | 102.8516 | 21.496 | 213.9545 | 160.4795 | 8.1011 | 253.1583
2010-08-12 | 17:38:01 | 6.192 | 307.137 | 30.43 | 6254 | 17.5 |  | 102.8516 | 21.496 | 213.9545 | 160.4795 | 8.1011 | 253.1583
2010-08-12 | 17:38:11 | 6.359 | 306.336 | 30.46 | 6257 | 17.5 |  | 102.5834 | 21.4399 | 213.3966 | 160.0611 | 8.0799 | 252.4983
2010-08-12 | 17:38:21 | 6.526 | 309.284 | 30.35 | 6255 | 17.5 |  | 103.5706 | 21.6463 | 215.4502 | 161.6014 | 8.1577 | 254.9282
2010-08-12 | 17:38:31 | 6.693 | 307.137 | 30.43 | 6266 | 17.5 |  | 102.8516 | 21.496 | 213.9545 | 160.4795 | 8.1011 | 253.1583
2010-08-12 | 17:38:41 | 6.86 | 307.137 | 30.43 | 6260 | 17.5 |  | 102.8516 | 21.496 | 213.9545 | 160.4795 | 8.1011 | 253.1583
2010-08-12 | 17:38:51 | 7.027 | 307.94 | 30.4 | 6266 | 17.5 |  | 103.1205 | 21.5522 | 214.514 | 160.8992 | 8.1223 | 253.8204
2010-08-12 | 17:39:01 | 7.194 | 304.742 | 30.52 | 6270 | 17.5 |  | 102.0494 | 21.3283 | 212.2858 | 159.2279 | 8.0379 | 251.1839
2010-08-12 | 17:39:11 | 7.361 | 304.212 | 30.54 | 6257 | 17.5 |  | 101.8721 | 21.2913 | 211.917 | 158.9513 | 8.0239 | 250.7475
2010-08-12 | 17:39:21 | 7.527 | 306.603 | 30.45 | 6264 | 17.5 |  | 102.6727 | 21.4586 | 213.5824 | 160.2004 | 8.087 | 252.7181
2010-08-12 | 17:39:31 | 7.694 | 303.156 | 30.58 | 6265 | 17.5 |  | 101.5186 | 21.2174 | 211.1816 | 158.3996 | 7.9961 | 249.8773
2010-08-12 | 17:39:41 | 7.861 | 305.804 | 30.48 | 6271 | 17.5 |  | 102.405 | 21.4027 | 213.0256 | 159.7828 | 8.0659 | 252.0593
2010-08-12 | 17:39:51 | 8.028 | 306.336 | 30.46 | 6262 | 17.5 |  | 102.5834 | 21.4399 | 213.3966 | 160.0611 | 8.0799 | 252.4983
2010-08-12 | 17:40:01 | 8.195 | 305.007 | 30.51 | 6273 | 17.5 |  | 102.1382 | 21.3469 | 212.4705 | 159.3664 | 8.0449 | 251.4024
2010-08-12 | 17:40:11 | 8.362 | 303.948 | 30.55 | 6257 | 17.5 |  | 101.7836 | 21.2728 | 211.7329 | 158.8132 | 8.0169 | 250.5297
2010-08-12 | 17:40:21 | 8.529 | 303.948 | 30.55 | 6275 | 17.5 |  | 101.7836 | 21.2728 | 211.7329 | 158.8132 | 8.0169 | 250.5297
2010-08-12 | 17:40:31 | 8.696 | 305.409 | 30.45 | 6266 | 17.6 |  | 102.484 | 21.4191 | 213.1625 | 159.8855 | 8.0568 | 251.776
2010-08-12 | 17:40:40 | 8.846 | 303.555 | 30.52 | 6273 | 17.6 |  | 101.8618 | 21.2891 | 211.8684 | 158.9148 | 8.0079 | 250.2474
2010-08-12 | 17:40:50 | 9.013 | 302.764 | 30.55 | 6265 | 17.6 |  | 101.5965 | 21.2337 | 211.3166 | 158.5009 | 7.9871 | 249.5956
2010-08-12 | 17:41:00 | 9.18 | 304.348 | 30.49 | 6261 | 17.6 |  | 102.1279 | 21.3447 | 212.422 | 159.33 | 8.0288 | 250.9012
2010-08-12 | 17:41:10 | 9.347 | 303.027 | 30.54 | 6283 | 17.6 |  | 101.6848 | 21.2521 | 211.5003 | 158.6387 | 7.994 | 249.8127
2010-08-12 | 17:41:20 | 9.514 | 304.083 | 30.5 | 6270 | 17.6 |  | 102.0391 | 21.3262 | 212.2373 | 159.1915 | 8.0219 | 250.6831
2010-08-12 | 17:41:30 | 9.681 | 302.501 | 30.56 | 6281 | 17.6 |  | 101.5082 | 21.2152 | 211.133 | 158.3632 | 7.9801 | 249.3787
2010-08-12 | 17:41:40 | 9.847 | 303.819 | 30.51 | 6277 | 17.6 |  | 101.9504 | 21.3076 | 212.0528 | 159.0531 | 8.0149 | 250.4652
2010-08-12 | 17:41:50 | 10.014 | 303.555 | 30.52 | 6276 | 17.6 |  | 101.8618 | 21.2891 | 211.8684 | 158.9148 | 8.0079 | 250.2474
2010-08-12 | 17:42:00 | 10.181 | 304.878 | 30.47 | 6273 | 17.6 |  | 102.3058 | 21.3819 | 212.7919 | 159.6075 | 8.0428 | 251.3382
2010-08-12 | 17:42:10 | 10.348 | 302.764 | 30.55 | 6260 | 17.6 |  | 101.5965 | 21.2337 | 211.3166 | 158.5009 | 7.9871 | 249.5956
2010-08-12 | 17:42:20 | 10.515 | 304.083 | 30.5 | 6274 | 17.6 |  | 102.0391 | 21.3262 | 212.2373 | 159.1915 | 8.0219 | 250.6831
2010-08-12 | 17:42:30 | 10.682 | 305.675 | 30.44 | 6253 | 17.6 |  | 102.5732 | 21.4378 | 213.3481 | 160.0247 | 8.0638 | 251.9952
2010-08-12 | 17:42:40 | 10.849 | 306.474 | 30.41 | 6264 | 17.6 |  | 102.8414 | 21.4939 | 213.9061 | 160.4432 | 8.0849 | 252.6542
2010-08-12 | 17:42:50 | 11.016 | 306.87 | 30.44 | 6254 | 17.5 |  | 102.7621 | 21.4773 | 213.7683 | 160.3399 | 8.094 | 252.9381
2010-08-12 | 17:43:00 | 11.183 | 306.87 | 30.44 | 6258 | 17.5 |  | 102.7621 | 21.4773 | 213.7683 | 160.3399 | 8.094 | 252.9381
2010-08-12 | 17:43:10 | 11.35 | 309.554 | 30.34 | 6257 | 17.5 |  | 103.6609 | 21.6651 | 215.638 | 161.7423 | 8.1648 | 255.1504
2010-08-12 | 17:43:20 | 11.517 | 307.672 | 30.41 | 6247 | 17.5 |  | 103.0308 | 21.5334 | 214.3273 | 160.7591 | 8.1152 | 253.5995
2010-08-12 | 17:43:30 | 11.683 | 307.94 | 30.4 | 6263 | 17.5 |  | 103.1205 | 21.5522 | 214.514 | 160.8992 | 8.1223 | 253.8204
2010-08-12 | 17:43:40 | 11.851 | 304.477 | 30.53 | 6263 | 17.5 |  | 101.9607 | 21.3098 | 212.1013 | 159.0895 | 8.0309 | 250.9656
2010-08-12 | 17:43:50 | 12.017 | 306.336 | 30.46 | 6244 | 17.5 |  | 102.5834 | 21.4399 | 213.3966 | 160.0611 | 8.0799 | 252.4983
2010-08-12 | 17:44:00 | 12.184 | 307.672 | 30.41 | 6226 | 17.5 |  | 103.0308 | 21.5334 | 214.3273 | 160.7591 | 8.1152 | 253.5995
2010-08-12 | 17:44:10 | 12.351 | 310.495 | 30.35 | 6253 | 17.4 |  | 103.7614 | 21.6861 | 215.8746 | 161.9197 | 8.1883 | 255.8828
2010-08-12 | 17:44:20 | 12.518 | 308.339 | 30.43 | 6236 | 17.4 |  | 103.0411 | 21.5356 | 214.376 | 160.7957 | 8.1314 | 254.1065
2010-08-12 | 17:44:30 | 12.685 | 309.415 | 30.39 | 6235 | 17.4 |  | 103.4005 | 21.6107 | 215.1238 | 161.3566 | 8.1598 | 254.9929
2010-08-12 | 17:44:40 | 12.852 | 307.268 | 30.47 | 6233 | 17.4 |  | 102.6831 | 21.4608 | 213.6311 | 160.237 | 8.1032 | 253.2236
2010-08-12 | 17:44:50 | 13.019 | 308.608 | 30.42 | 6231 | 17.4 |  | 103.1308 | 21.5543 | 214.5627 | 160.9357 | 8.1385 | 254.3278
2010-08-12 | 17:45:00 | 13.186 | 311.307 | 30.32 | 6233 | 17.4 |  | 104.033 | 21.7429 | 216.4396 | 162.3435 | 8.2097 | 256.5526
2010-08-12 | 17:45:11 | 13.353 | 312.94 | 30.26 | 6239 | 17.4 |  | 104.5786 | 21.8569 | 217.5748 | 163.195 | 8.2527 | 257.8981
2010-08-12 | 17:45:21 | 13.52 | 310.765 | 30.34 | 6224 | 17.4 |  | 103.8519 | 21.705 | 216.0628 | 162.0608 | 8.1954 | 256.1058
2010-08-12 | 17:45:31 | 13.686 | 308.877 | 30.41 | 6227 | 17.4 |  | 103.2206 | 21.5731 | 214.7495 | 161.0758 | 8.1456 | 254.5492
2010-08-12 | 17:45:41 | 13.853 | 309.824 | 30.33 | 6219 | 17.5 |  | 103.7512 | 21.684 | 215.826 | 161.8833 | 8.1719 | 255.3728
2010-08-12 | 17:45:51 | 14.02 | 308.746 | 30.37 | 6227 | 17.5 |  | 103.3903 | 21.6086 | 215.0752 | 161.3201 | 8.1435 | 254.4844
2010-08-12 | 17:46:01 | 14.187 | 310.094 | 30.32 | 6220 | 17.5 |  | 103.8417 | 21.7029 | 216.0142 | 162.0244 | 8.1791 | 255.5955
2010-08-12 | 17:46:11 | 14.354 | 310.364 | 30.31 | 6215 | 17.5 |  | 103.9323 | 21.7218 | 216.2026 | 162.1657 | 8.1862 | 255.8183
2010-08-12 | 17:46:21 | 14.521 | 311.992 | 30.25 | 6201 | 17.5 |  | 104.4774 | 21.8358 | 217.3366 | 163.0163 | 8.2291 | 257.1602
2010-08-12 | 17:46:31 | 14.688 | 312.81 | 30.22 | 6211 | 17.5 |  | 104.7512 | 21.893 | 217.9062 | 163.4436 | 8.2507 | 257.8342
2010-08-12 | 17:46:41 | 14.855 | 311.992 | 30.25 | 6193 | 17.5 |  | 104.4774 | 21.8358 | 217.3366 | 163.0163 | 8.2291 | 257.1602
2010-08-12 | 17:46:51 | 15.021 | 313.356 | 30.2 | 6194 | 17.5 |  | 104.9342 | 21.9313 | 218.2869 | 163.7291 | 8.2651 | 258.2846
2010-08-12 | 17:47:01 | 15.189 | 313.63 | 30.19 | 6191 | 17.5 |  | 105.0259 | 21.9504 | 218.4775 | 163.8721 | 8.2723 | 258.5101
2010-08-12 | 17:47:11 | 15.355 | 313.63 | 30.19 | 6200 | 17.5 |  | 105.0259 | 21.9504 | 218.4775 | 163.8721 | 8.2723 | 258.5101
2010-08-12 | 17:47:21 | 15.522 | 315.553 | 30.12 | 6194 | 17.5 |  | 105.6699 | 22.085 | 219.8172 | 164.8769 | 8.323 | 260.0953
2010-08-12 | 17:47:31 | 15.689 | 311.992 | 30.25 | 6191 | 17.5 |  | 104.4774 | 21.8358 | 217.3366 | 163.0163 | 8.2291 | 257.1602
2010-08-12 | 17:47:41 | 15.856 | 313.083 | 30.21 | 6180 | 17.5 |  | 104.8427 | 21.9121 | 218.0965 | 163.5863 | 8.2579 | 258.0593
2010-08-12 | 17:47:51 | 16.023 | 312.81 | 30.22 | 6175 | 17.5 |  | 104.7512 | 21.893 | 217.9062 | 163.4436 | 8.2507 | 257.8342
2010-08-12 | 17:48:01 | 16.19 | 312.409 | 30.19 | 6174 | 17.6 |  | 104.833 | 21.9101 | 218.0484 | 163.5502 | 8.2415 | 257.5468
2010-08-12 | 17:48:11 | 16.357 | 314.325 | 30.12 | 6174 | 17.6 |  | 105.4758 | 22.0444 | 219.3854 | 164.5531 | 8.292 | 259.1261
2010-08-12 | 17:48:21 | 16.524 | 312.409 | 30.19 | 6175 | 17.6 |  | 104.833 | 21.9101 | 218.0484 | 163.5502 | 8.2415 | 257.5468
2010-08-12 | 17:48:31 | 16.691 | 315.701 | 30.07 | 6157 | 17.6 |  | 105.9377 | 22.141 | 220.3462 | 165.2737 | 8.3284 | 260.261
2010-08-12 | 17:48:41 | 16.858 | 313.228 | 30.16 | 6169 | 17.6 |  | 105.1079 | 21.9676 | 218.6203 | 163.9791 | 8.2631 | 258.2223
2010-08-12 | 17:48:51 | 17.024 | 317.362 | 30.01 | 6160 | 17.6 |  | 106.4951 | 22.2575 | 221.5055 | 166.1433 | 8.3722 | 261.6303
2010-08-12 | 17:49:01 | 17.191 | 316.254 | 30.05 | 6158 | 17.6 |  | 106.1232 | 22.1797 | 220.7319 | 165.563 | 8.3429 | 260.7165
2010-08-12 | 17:49:11 | 17.358 | 315.425 | 30.08 | 6156 | 17.6 |  | 105.8452 | 22.1216 | 220.1537 | 165.1293 | 8.3211 | 260.0335
2010-08-12 | 17:49:21 | 17.525 | 315.277 | 30.13 | 6135 | 17.5 |  | 105.5776 | 22.0657 | 219.6252 | 164.7329 | 8.3158 | 259.8681
2010-08-12 | 17:49:31 | 17.692 | 317.49 | 30.05 | 6134 | 17.5 |  | 106.3184 | 22.2205 | 221.1662 | 165.8887 | 8.3741 | 261.6915
2010-08-12 | 17:49:41 | 17.859 | 319.719 | 29.97 | 6126 | 17.5 |  | 107.0651 | 22.3766 | 222.7195 | 167.0538 | 8.4329 | 263.5295
2010-08-12 | 17:49:51 | 18.026 | 319.16 | 29.99 | 6129 | 17.5 |  | 106.8778 | 22.3375 | 222.33 | 166.7617 | 8.4182 | 263.0686
2010-08-12 | 17:50:01 | 18.193 | 321.685 | 29.9 | 6128 | 17.5 |  | 107.7234 | 22.5142 | 224.089 | 168.081 | 8.4848 | 265.1498
2010-08-12 | 17:50:11 | 18.36 | 321.967 | 29.89 | 6132 | 17.5 |  | 107.8178 | 22.5339 | 224.2854 | 168.2283 | 8.4922 | 265.3822
2010-08-12 | 17:50:21 | 18.527 | 319.44 | 29.98 | 6113 | 17.5 |  | 106.9714 | 22.357 | 222.5247 | 166.9077 | 8.4256 | 263.2989
2010-08-12 | 17:50:31 | 18.694 | 322.944 | 29.9 | 6116 | 17.4 |  | 107.9216 | 22.5556 | 224.5298 | 168.4117 | 8.5165 | 266.1421
2010-08-12 | 17:50:41 | 18.86 | 322.378 | 29.92 | 6122 | 17.4 |  | 107.7327 | 22.5161 | 224.1368 | 168.1169 | 8.5016 | 265.6763
2010-08-12 | 17:50:51 | 19.028 | 321.251 | 29.96 | 6119 | 17.4 |  | 107.356 | 22.4374 | 223.3532 | 167.5291 | 8.4719 | 264.7474
2010-08-12 | 17:51:01 | 19.194 | 324.078 | 29.86 | 6105 | 17.4 |  | 108.3005 | 22.6348 | 225.3182 | 169.003 | 8.5465 | 267.0766
2010-08-12 | 17:51:11 | 19.361 | 322.944 | 29.9 | 6095 | 17.4 |  | 107.9216 | 22.5556 | 224.5298 | 168.4117 | 8.5165 | 266.1421
2010-08-12 | 17:51:21 | 19.528 | 327.219 | 29.75 | 6098 | 17.4 |  | 109.3505 | 22.8543 | 227.5026 | 170.6415 | 8.6293 | 269.6658
2010-08-12 | 17:51:31 | 19.695 | 324.362 | 29.85 | 6097 | 17.4 |  | 108.3955 | 22.6547 | 225.5158 | 169.1512 | 8.5539 | 267.3108
2010-08-12 | 17:51:41 | 19.862 | 324.646 | 29.84 | 6084 | 17.4 |  | 108.4906 | 22.6745 | 225.7135 | 169.2996 | 8.5614 | 267.5452
2010-08-12 | 17:51:51 | 20.029 | 326.073 | 29.79 | 6081 | 17.4 |  | 108.9673 | 22.7742 | 226.7055 | 170.0436 | 8.5991 | 268.721
2010-08-12 | 17:52:01 | 20.196 | 325.216 | 29.82 | 6083 | 17.4 |  | 108.681 | 22.7143 | 226.1097 | 169.5967 | 8.5765 | 268.0148
2010-08-12 | 17:52:10 | 20.346 | 325.787 | 29.8 | 6078 | 17.4 |  | 108.8718 | 22.7542 | 226.5067 | 169.8945 | 8.5915 | 268.4853
2010-08-12 | 17:52:20 | 20.513 | 322.25 | 29.88 | 6071 | 17.5 |  | 107.9123 | 22.5537 | 224.482 | 168.3758 | 8.4997 | 265.6149
2010-08-12 | 17:52:30 | 20.68 | 326.804 | 29.72 | 6073 | 17.5 |  | 109.4376 | 22.8725 | 227.6549 | 170.7557 | 8.6198 | 269.3692
2010-08-12 | 17:52:40 | 20.847 | 327.667 | 29.69 | 6074 | 17.5 |  | 109.7263 | 22.9328 | 228.2555 | 171.2062 | 8.6426 | 270.0799
2010-08-12 | 17:52:50 | 21.014 | 326.804 | 29.72 | 6077 | 17.5 |  | 109.4376 | 22.8725 | 227.6549 | 170.7557 | 8.6198 | 269.3692
2010-08-12 | 17:53:00 | 21.181 | 327.667 | 29.69 | 6062 | 17.5 |  | 109.7263 | 22.9328 | 228.2555 | 171.2062 | 8.6426 | 270.0799
2010-08-12 | 17:53:10 | 21.348 | 329.689 | 29.62 | 6055 | 17.5 |  | 110.4035 | 23.0743 | 229.6642 | 172.2628 | 8.6959 | 271.7466
2010-08-12 | 17:53:20 | 21.514 | 326.517 | 29.73 | 6044 | 17.5 |  | 109.3415 | 22.8524 | 227.4551 | 170.6058 | 8.6122 | 269.1327
2010-08-12 | 17:53:30 | 21.681 | 327.955 | 29.68 | 6054 | 17.5 |  | 109.8228 | 22.953 | 228.4562 | 171.3567 | 8.6502 | 270.3172
2010-08-12 | 17:53:40 | 21.848 | 326.231 | 29.74 | 6041 | 17.5 |  | 109.2456 | 22.8323 | 227.2555 | 170.4561 | 8.6047 | 268.8966
2010-08-12 | 17:53:50 | 22.015 | 327.667 | 29.69 | 6050 | 17.5 |  | 109.7263 | 22.9328 | 228.2555 | 171.2062 | 8.6426 | 270.0799
2010-08-12 | 17:54:00 | 22.182 | 327.667 | 29.69 | 6038 | 17.5 |  | 109.7263 | 22.9328 | 228.2555 | 171.2062 | 8.6426 | 270.0799
2010-08-12 | 17:54:10 | 22.349 | 327.379 | 29.7 | 6034 | 17.5 |  | 109.63 | 22.9127 | 228.0551 | 171.0559 | 8.635 | 269.8427
2010-08-12 | 17:54:20 | 22.516 | 329.109 | 29.64 | 6033 | 17.5 |  | 110.2095 | 23.0338 | 229.2607 | 171.9601 | 8.6806 | 271.2692
2010-08-12 | 17:54:30 | 22.683 | 327.541 | 29.65 | 6023 | 17.6 |  | 109.9107 | 22.9713 | 228.6099 | 171.472 | 8.6407 | 270.0216
2010-08-12 | 17:54:40 | 22.849 | 327.829 | 29.64 | 6016 | 17.6 |  | 110.0074 | 22.9915 | 228.811 | 171.6228 | 8.6483 | 270.2591
2010-08-12 | 17:54:50 | 23.016 | 330.435 | 29.55 | 6020 | 17.6 |  | 110.8818 | 23.1743 | 230.6296 | 172.9869 | 8.717 | 272.4071
2010-08-12 | 17:55:00 | 23.183 | 329.564 | 29.58 | 6007 | 17.6 |  | 110.5894 | 23.1132 | 230.0216 | 172.5308 | 8.694 | 271.6889
2010-08-12 | 17:55:10 | 23.35 | 333.652 | 29.44 | 5999 | 17.6 |  | 111.9613 | 23.3999 | 232.8751 | 174.6712 | 8.8019 | 275.0594
2010-08-12 | 17:55:20 | 23.517 | 333.064 | 29.46 | 6011 | 17.6 |  | 111.7642 | 23.3587 | 232.465 | 174.3635 | 8.7864 | 274.5749
2010-08-12 | 17:55:30 | 23.684 | 331.893 | 29.5 | 5998 | 17.6 |  | 111.371 | 23.2765 | 231.6472 | 173.7502 | 8.7555 | 273.609
2010-08-12 | 17:55:40 | 23.851 | 330.435 | 29.55 | 5994 | 17.6 |  | 110.8818 | 23.1743 | 230.6296 | 172.9869 | 8.717 | 272.4071
2010-08-12 | 17:55:50 | 24.018 | 333.358 | 29.45 | 5985 | 17.6 |  | 111.8627 | 23.3793 | 232.6699 | 174.5173 | 8.7941 | 274.817
2010-08-12 | 17:56:00 | 24.185 | 334.831 | 29.4 | 5967 | 17.6 |  | 112.3569 | 23.4826 | 233.6979 | 175.2883 | 8.833 | 276.0312
2010-08-12 | 17:56:10 | 24.352 | 332.478 | 29.48 | 5969 | 17.6 |  | 111.5674 | 23.3176 | 232.0557 | 174.0565 | 8.7709 | 274.0915
2010-08-12 | 17:56:20 | 24.519 | 336.732 | 29.38 | 5969 | 17.5 |  | 112.762 | 23.5673 | 234.5704 | 175.9427 | 8.8817 | 277.5518
2010-08-12 | 17:56:30 | 24.685 | 336.138 | 29.4 | 5961 | 17.5 |  | 112.5632 | 23.5257 | 234.1569 | 175.6326 | 8.866 | 277.0626
2010-08-12 | 17:56:40 | 24.852 | 338.519 | 29.32 | 5952 | 17.5 |  | 113.3607 | 23.6924 | 235.8158 | 176.8769 | 8.9288 | 279.0254
2010-08-12 | 17:56:51 | 25.019 | 335.842 | 29.41 | 5957 | 17.5 |  | 112.464 | 23.505 | 233.9505 | 175.4778 | 8.8582 | 276.8183
2010-08-12 | 17:57:01 | 25.186 | 338.519 | 29.32 | 5955 | 17.5 |  | 113.3607 | 23.6924 | 235.8158 | 176.8769 | 8.9288 | 279.0254
2010-08-12 | 17:57:11 | 25.353 | 337.624 | 29.35 | 5947 | 17.5 |  | 113.0609 | 23.6297 | 235.1921 | 176.4091 | 8.9052 | 278.2874
2010-08-12 | 17:57:21 | 25.52 | 340.018 | 29.27 | 5946 | 17.5 |  | 113.8624 | 23.7972 | 236.8595 | 177.6597 | 8.9683 | 280.2603
2010-08-12 | 17:57:31 | 25.687 | 338.818 | 29.31 | 5945 | 17.5 |  | 113.4608 | 23.7133 | 236.0241 | 177.0331 | 8.9367 | 279.2719
2010-08-12 | 17:57:41 | 25.854 | 341.648 | 29.26 | 5943 | 17.4 |  | 114.1724 | 23.862 | 237.5345 | 178.166 | 9.0098 | 281.557
2010-08-12 | 17:57:51 | 26.021 | 339.843 | 29.32 | 5928 | 17.4 |  | 113.5689 | 23.7359 | 236.279 | 177.2243 | 8.9622 | 280.0688
2010-08-12 | 17:58:01 | 26.188 | 342.555 | 29.23 | 5939 | 17.4 |  | 114.4755 | 23.9254 | 238.1651 | 178.639 | 9.0337 | 282.3045
2010-08-12 | 17:58:11 | 26.355 | 342.858 | 29.22 | 5938 | 17.4 |  | 114.5767 | 23.9465 | 238.3758 | 178.797 | 9.0417 | 282.5542
2010-08-12 | 17:58:21 | 26.521 | 342.253 | 29.24 | 5922 | 17.4 |  | 114.3744 | 23.9042 | 237.9547 | 178.4812 | 9.0258 | 282.055
2010-08-12 | 17:58:31 | 26.688 | 342.858 | 29.22 | 5921 | 17.4 |  | 114.5767 | 23.9465 | 238.3758 | 178.797 | 9.0417 | 282.5542
2010-08-12 | 17:58:41 | 26.855 | 340.318 | 29.26 | 5913 | 17.5 |  | 113.963 | 23.8183 | 237.0689 | 177.8168 | 8.9763 | 280.5081
2010-08-12 | 17:58:51 | 27.022 | 338.818 | 29.31 | 5910 | 17.5 |  | 113.4608 | 23.7133 | 236.0241 | 177.0331 | 8.9367 | 279.2719
2010-08-12 | 17:59:01 | 27.189 | 341.826 | 29.21 | 5896 | 17.5 |  | 114.4679 | 23.9238 | 238.119 | 178.6044 | 9.016 | 281.7506
2010-08-12 | 17:59:11 | 27.356 | 342.734 | 29.18 | 5907 | 17.5 |  | 114.772 | 23.9873 | 238.7517 | 179.079 | 9.04 | 282.4992
2010-08-12 | 17:59:21 | 27.523 | 339.417 | 29.29 | 5904 | 17.5 |  | 113.6614 | 23.7552 | 236.4414 | 177.3461 | 8.9525 | 279.7656
2010-08-12 | 17:59:31 | 27.69 | 342.128 | 29.2 | 5903 | 17.5 |  | 114.5691 | 23.945 | 238.3297 | 178.7625 | 9.024 | 281.9999
2010-08-12 | 17:59:41 | 27.857 | 347.01 | 29.04 | 5888 | 17.5 |  | 116.2038 | 24.2866 | 241.7301 | 181.313 | 9.1528 | 286.0234
2010-08-12 | 17:59:51 | 28.024 | 343.645 | 29.15 | 5879 | 17.5 |  | 115.0771 | 24.0511 | 239.3863 | 179.555 | 9.064 | 283.2502
2010-08-12 | 18:00:01 | 28.191 | 346.088 | 29.07 | 5871 | 17.5 |  | 115.8952 | 24.2221 | 241.0883 | 180.8316 | 9.1284 | 285.264
2010-08-12 | 18:00:11 | 28.358 | 347.626 | 29.02 | 5876 | 17.5 |  | 116.41 | 24.3297 | 242.1591 | 181.6348 | 9.169 | 286.5311
2010-08-12 | 18:00:21 | 28.525 | 343.949 | 29.14 | 5876 | 17.5 |  | 115.179 | 24.0724 | 239.5983 | 179.714 | 9.072 | 283.501
2010-08-12 | 18:00:31 | 28.691 | 346.702 | 29.05 | 5879 | 17.5 |  | 116.1008 | 24.2651 | 241.516 | 181.1524 | 9.1446 | 285.77
2010-08-12 | 18:00:41 | 28.858 | 346.395 | 29.06 | 5873 | 17.5 |  | 115.998 | 24.2436 | 241.302 | 180.9919 | 9.1365 | 285.5169
2010-08-12 | 18:00:51 | 29.025 | 346.395 | 29.06 | 5857 | 17.5 |  | 115.998 | 24.2436 | 241.302 | 180.9919 | 9.1365 | 285.5169
2010-08-12 | 18:01:01 | 29.192 | 350.102 | 28.94 | 5865 | 17.5 |  | 117.2392 | 24.503 | 243.8841 | 182.9286 | 9.2343 | 288.5721
2010-08-12 | 18:01:11 | 29.359 | 347.505 | 28.98 | 5859 | 17.6 |  | 116.6099 | 24.3715 | 242.5439 | 181.9234 | 9.1673 | 286.4797
2010-08-12 | 18:01:21 | 29.526 | 348.432 | 28.95 | 5852 | 17.6 |  | 116.9208 | 24.4364 | 243.1906 | 182.4084 | 9.1918 | 287.2434
2010-08-12 | 18:01:31 | 29.693 | 349.982 | 28.9 | 5846 | 17.6 |  | 117.4411 | 24.5452 | 244.2727 | 183.2201 | 9.2327 | 288.5216
2010-08-12 | 18:01:41 | 29.86 | 349.671 | 28.91 | 5830 | 17.6 |  | 117.3368 | 24.5234 | 244.0559 | 183.0575 | 9.2245 | 288.2655
2010-08-12 | 18:01:51 | 30.027 | 349.361 | 28.92 | 5835 | 17.6 |  | 117.2326 | 24.5016 | 243.8392 | 182.8949 | 9.2163 | 288.0096
2010-08-12 | 18:02:01 | 30.194 | 348.432 | 28.95 | 5827 | 17.6 |  | 116.9208 | 24.4364 | 243.1906 | 182.4084 | 9.1918 | 287.2434
2010-08-12 | 18:02:11 | 30.36 | 352.166 | 28.83 | 5816 | 17.6 |  | 118.174 | 24.6984 | 245.7973 | 184.3636 | 9.2903 | 290.3223
2010-08-12 | 18:02:21 | 30.527 | 350.293 | 28.89 | 5824 | 17.6 |  | 117.5455 | 24.567 | 244.4899 | 183.383 | 9.2409 | 288.7781
2010-08-12 | 18:02:31 | 30.694 | 350.916 | 28.87 | 5818 | 17.6 |  | 117.7546 | 24.6107 | 244.9247 | 183.7092 | 9.2573 | 289.2918
2010-08-12 | 18:02:41 | 30.861 | 350.916 | 28.87 | 5815 | 17.6 |  | 117.7546 | 24.6107 | 244.9247 | 183.7092 | 9.2573 | 289.2918
2010-08-12 | 18:02:51 | 31.028 | 353.736 | 28.78 | 5808 | 17.6 |  | 118.7008 | 24.8085 | 246.893 | 185.1855 | 9.3317 | 291.6165
2010-08-12 | 18:03:01 | 31.195 | 354.051 | 28.77 | 5797 | 17.6 |  | 118.8065 | 24.8306 | 247.1128 | 185.3504 | 9.34 | 291.8762
2010-08-12 | 18:03:11 | 31.362 | 353.107 | 28.8 | 5802 | 17.6 |  | 118.4898 | 24.7644 | 246.454 | 184.8562 | 9.3151 | 291.098
2010-08-12 | 18:03:21 | 31.528 | 352.166 | 28.83 | 5805 | 17.6 |  | 118.174 | 24.6984 | 245.7973 | 184.3636 | 9.2903 | 290.3223
2010-08-12 | 18:03:31 | 31.696 | 355.948 | 28.71 | 5778 | 17.6 |  | 119.443 | 24.9636 | 248.4367 | 186.3433 | 9.3901 | 293.4398
2010-08-12 | 18:03:41 | 31.862 | 350.605 | 28.88 | 5796 | 17.6 |  | 117.65 | 24.5888 | 244.7072 | 183.546 | 9.2491 | 289.0348
2010-08-12 | 18:03:50 | 32.013 | 355.948 | 28.71 | 5772 | 17.6 |  | 119.443 | 24.9636 | 248.4367 | 186.3433 | 9.3901 | 293.4398
2010-08-12 | 18:04:00 | 32.18 | 356.583 | 28.69 | 5783 | 17.6 |  | 119.656 | 25.0081 | 248.8798 | 186.6757 | 9.4068 | 293.9632
2010-08-12 | 18:04:10 | 32.347 | 355.314 | 28.73 | 5771 | 17.6 |  | 119.2304 | 24.9192 | 247.9945 | 186.0117 | 9.3734 | 292.9175
2010-08-12 | 18:04:20 | 32.513 | 353.736 | 28.78 | 5770 | 17.6 |  | 118.7008 | 24.8085 | 246.893 | 185.1855 | 9.3317 | 291.6165
2010-08-12 | 18:04:30 | 32.68 | 356.901 | 28.68 | 5765 | 17.6 |  | 119.7627 | 25.0304 | 249.1017 | 186.8421 | 9.4152 | 294.2253
2010-08-12 | 18:04:40 | 32.847 | 360.536 | 28.61 | 5760 | 17.5 |  | 120.7335 | 25.2333 | 251.1529 | 188.3807 | 9.5095 | 297.1728
2010-08-12 | 18:04:50 | 33.014 | 359.252 | 28.65 | 5758 | 17.5 |  | 120.3035 | 25.1434 | 250.2584 | 187.7098 | 9.4757 | 296.1145
2010-08-12 | 18:05:00 | 33.181 | 360.536 | 28.61 | 5756 | 17.5 |  | 120.7335 | 25.2333 | 251.1529 | 188.3807 | 9.5095 | 297.1728
2010-08-12 | 18:05:10 | 33.348 | 357.336 | 28.71 | 5761 | 17.5 |  | 119.6619 | 25.0093 | 248.9238 | 186.7087 | 9.4251 | 294.5352
2010-08-12 | 18:05:20 | 33.515 | 359.573 | 28.64 | 5733 | 17.5 |  | 120.4108 | 25.1659 | 250.4817 | 187.8773 | 9.4841 | 296.3786
2010-08-12 | 18:05:30 | 33.682 | 363.12 | 28.53 | 5758 | 17.5 |  | 121.5988 | 25.4142 | 252.953 | 189.7309 | 9.5777 | 299.3028
2010-08-12 | 18:05:40 | 33.849 | 363.77 | 28.51 | 5739 | 17.5 |  | 121.8163 | 25.4596 | 253.4055 | 190.0702 | 9.5948 | 299.8381
2010-08-12 | 18:05:50 | 34.016 | 360.215 | 28.62 | 5725 | 17.5 |  | 120.6258 | 25.2108 | 250.9289 | 188.2127 | 9.501 | 296.9078
2010-08-12 | 18:06:00 | 34.182 | 365.726 | 28.45 | 5734 | 17.5 |  | 122.4715 | 25.5965 | 254.7684 | 191.0925 | 9.6464 | 301.4508
2010-08-12 | 18:06:10 | 34.349 | 360.215 | 28.62 | 5741 | 17.5 |  | 120.6258 | 25.2108 | 250.9289 | 188.2127 | 9.501 | 296.9078
2010-08-12 | 18:06:20 | 34.516 | 365.399 | 28.46 | 5720 | 17.5 |  | 122.362 | 25.5737 | 254.5406 | 190.9217 | 9.6378 | 301.1813
2010-08-12 | 18:06:30 | 34.683 | 364.747 | 28.48 | 5716 | 17.5 |  | 122.1434 | 25.528 | 254.0858 | 190.5806 | 9.6206 | 300.6432