## convers.+compens.
file name: PSt1_Eq4_Sal |  |  | 2008-03-29
author: Dr. Christian Huber
programm for temperature compensated oxygen calculation of membrane PSt1
Fit of the calibration curve is optimized between 0 and 250 % air-saturation (0 and 388 Torr)
enter your values here
input of F0,T0 | 58.2 | phase angle of cal 0
input of F100,T100 | 28.64 | phase angle of cal 100 (water-vapor saturated air)
input of Fm,Tm | 15 | measured phase angle
input of T0 | 17.6 | temperature of cal 0
input of T100 | 17.7 | temperatur of cal100
input of Tm | 17.4 | temperature at measurement
air pressure patm | 1013 | air pressure
Salinity | 32.2 | Salinity
Results of calculation + compensation
[%] air saturation | 417.4737 | %
[%] O2 | 87.252 | %
pO2 [hPa] | 868.5499 | hPa
pO2 [Torr] | 651.4678 | Torr
cO2 [mg/L] | 32.9446 | mg/L
cO2 [ppm] | 32.9446 | ppm
cO2 [µmol/L] | 1029.5196 | µmol/L
Internal calculated parameter (DO NOT CHANGE !!)
f1 | 0.801
DF/K | -0.08
DKSV/K | 0.000383
m | 22.9
tan(F0,T100) | 1.6123
tan(F0,Tm) | 1.6138
tan(F100,T100) | 0.54612
tan(Fm,Tm) | 0.26795
KSV,T100 | 0.037719
KSV,Tm | 0.037604
A | 147.9112
B | 11.953
C | -0.66128
a | 1.0252e-05
b | -0.0022824
c | -0.83397
Chlorinity | 17.8227

## Recalculate Fibox values
file name: PSt1_Eq4_Sal |  |  |  | 2008-03-29
author: Dr. Christian Huber
programm for temperature compensated oxygen calculation of membrane PSt1
Fit of the calibration curve is optimized between 0 and 250 % air-saturation (0 and 388 Torr)
Insert your calibration values
cal0 | 58.2 | ° | T0 | 17.6 | °C
cal100 | 28.64 | ° | T100 | 17.7 | °C
air pressure | 1013 | mbar
Salinity | 32.2 | °/°°
Internal calculated parameter (DO NOT CHANGE !!)
f1 | 0.801 | tan(F0,T100) | 1.6123 | A | 147.9112 | KSV,T100 | 0.037719 | a | 2.0731e-05
DF/K | -0.08 | tan(F0,Tm) | 1.6138 | B | 11.953 | KSV,Tm | 0.037604 | b | 0.0043778 |  |  |  |  | INSERT YOUR RESPECTIVE DATA
DKSV/K | 0.000383 | tan(F100,T100) | 0.54612 | C | -0.66128 |  |  | c | -0.66426 |  |  |  |  | chamber volume [L] | FW alga [g] | Chl a Thallus-1
m | 22.9 | tan(Fm,Tm) | 0.54183 |  |  |  |  | Chlorinity | 17.8227 |  |  |  |  | 0.026 | 0.1365 | 0.033162
Insert the data you get from the Ascii File recorded with the OxyView TX3-V5.31.exe software | [7 empty] | The oxygen contents are calculated automatically from the raw data
date | time/hh:mm:ss | logtime/min | oxygen/% airsatur. | phase/° | amp | temp/°C |  | % air-sat. | % oxygen | pO2 (hPa) | pO2 (Torr) | cO2 [mg/L] | cO2 [µmol/L] | cO2 [µmol/L 10 min] | cO2 [µmol/L h] | cO2 [µmol/ h] | cO2 [µmol/ h gFW] | cO2 [µmol/ h g Chl a]
2010-08-12 | 15:10:34 | 0 | 307.163 | 28.45 | 2878 | 17.4 |  | 102.2363 | 21.3674 | 212.7016 | 159.5398 | 8.0679 | 252.1218 | 27.042 | 162.252 | 4.2186 | 30.9051 | 127.2093
2010-08-12 | 15:10:46 | 0.201 | 303.6 | 28.58 | 2879 | 17.4 |  | 101.0503 | 21.1195 | 210.2342 | 157.689 | 7.9743 | 249.197
2010-08-12 | 15:10:56 | 0.368 | 303.872 | 28.57 | 2872 | 17.4 |  | 101.141 | 21.1385 | 210.4228 | 157.8305 | 7.9815 | 249.4206 | change data input according to column N
2010-08-12 | 15:11:06 | 0.535 | 303.6 | 28.58 | 2877 | 17.4 |  | 101.0503 | 21.1195 | 210.2342 | 157.689 | 7.9743 | 249.197
2010-08-12 | 15:11:16 | 0.702 | 299.872 | 28.76 | 2873 | 17.3 |  | 99.6031 | 20.817 | 207.2494 | 155.4503 | 7.875 | 246.0951
2010-08-12 | 15:11:26 | 0.869 | 303.328 | 28.59 | 2875 | 17.4 |  | 100.9597 | 21.1006 | 210.0458 | 157.5477 | 7.9672 | 248.9737
2010-08-12 | 15:11:36 | 1.036 | 301.431 | 28.66 | 2878 | 17.4 |  | 100.3285 | 20.9686 | 208.7324 | 156.5626 | 7.9173 | 247.4169
2010-08-12 | 15:11:46 | 1.203 | 300.354 | 28.7 | 2874 | 17.4 |  | 99.9698 | 20.8937 | 207.9862 | 156.0029 | 7.889 | 246.5324
2010-08-12 | 15:11:56 | 1.37 | 302.785 | 28.61 | 2881 | 17.4 |  | 100.7789 | 21.0628 | 209.6695 | 157.2655 | 7.9529 | 248.5277
2010-08-12 | 15:12:06 | 1.537 | 299.817 | 28.72 | 2873 | 17.4 |  | 99.791 | 20.8563 | 207.6143 | 155.7239 | 7.8749 | 246.0916
2010-08-12 | 15:12:16 | 1.703 | 298.212 | 28.78 | 2883 | 17.4 |  | 99.257 | 20.7447 | 206.5031 | 154.8905 | 7.8328 | 244.7745
2010-08-12 | 15:12:26 | 1.87 | 300.892 | 28.68 | 2880 | 17.4 |  | 100.1489 | 20.9311 | 208.3589 | 156.2824 | 7.9032 | 246.9742
2010-08-12 | 15:12:36 | 2.037 | 300.623 | 28.69 | 2886 | 17.4 |  | 100.0593 | 20.9124 | 208.1724 | 156.1426 | 7.8961 | 246.7532
2010-08-12 | 15:12:46 | 2.204 | 299.548 | 28.73 | 2884 | 17.4 |  | 99.7018 | 20.8377 | 207.4286 | 155.5847 | 7.8679 | 245.8715
2010-08-12 | 15:12:56 | 2.371 | 299.817 | 28.72 | 2884 | 17.4 |  | 99.791 | 20.8563 | 207.6143 | 155.7239 | 7.8749 | 246.0916
2010-08-12 | 15:13:06 | 2.538 | 299.817 | 28.72 | 2882 | 17.4 |  | 99.791 | 20.8563 | 207.6143 | 155.7239 | 7.8749 | 246.0916
2010-08-12 | 15:13:16 | 2.705 | 297.148 | 28.82 | 2887 | 17.4 |  | 98.9028 | 20.6707 | 205.7662 | 154.3378 | 7.8048 | 243.901
2010-08-12 | 15:13:26 | 2.872 | 297.413 | 28.81 | 2888 | 17.4 |  | 98.9912 | 20.6892 | 205.9501 | 154.4757 | 7.8118 | 244.119
2010-08-12 | 15:13:36 | 3.039 | 298.212 | 28.78 | 2883 | 17.4 |  | 99.257 | 20.7447 | 206.5031 | 154.8905 | 7.8328 | 244.7745
2010-08-12 | 15:13:46 | 3.206 | 300.892 | 28.68 | 2894 | 17.4 |  | 100.1489 | 20.9311 | 208.3589 | 156.2824 | 7.9032 | 246.9742
2010-08-12 | 15:13:56 | 3.373 | 298.479 | 28.77 | 2891 | 17.4 |  | 99.3457 | 20.7633 | 206.6878 | 155.029 | 7.8398 | 244.9934
2010-08-12 | 15:14:06 | 3.54 | 297.413 | 28.81 | 2889 | 17.4 |  | 98.9912 | 20.6892 | 205.9501 | 154.4757 | 7.8118 | 244.119
2010-08-12 | 15:14:16 | 3.706 | 301.161 | 28.67 | 2890 | 17.4 |  | 100.2387 | 20.9499 | 208.5455 | 156.4224 | 7.9103 | 247.1954 |  | regression formula | time
2010-08-12 | 15:14:26 | 3.873 | 296.617 | 28.84 | 2888 | 17.4 |  | 98.7262 | 20.6338 | 205.3989 | 154.0623 | 7.7909 | 243.4656 | value for T=26 min. | 279.446 | T11
2010-08-12 | 15:14:36 | 4.04 | 295.824 | 28.87 | 2887 | 17.4 |  | 98.4621 | 20.5786 | 204.8494 | 153.6501 | 7.7701 | 242.8143 | value for T=1 min. | 252.404 | T1
2010-08-12 | 15:14:46 | 4.207 | 298.746 | 28.76 | 2889 | 17.4 |  | 99.4346 | 20.7818 | 206.8727 | 155.1677 | 7.8468 | 245.2126 | difference between T25 and T1 | 27.042 | 10minutes
2010-08-12 | 15:14:56 | 4.374 | 302.242 | 28.63 | 2885 | 17.4 |  | 100.5984 | 21.0251 | 209.2941 | 156.9839 | 7.9386 | 248.0827 | calculate from regression curve values for 10 minutes of photosynthesis or respiration
2010-08-12 | 15:15:06 | 4.541 | 299.013 | 28.75 | 2888 | 17.4 |  | 99.5236 | 20.8004 | 207.0578 | 155.3066 | 7.8538 | 245.432
2010-08-12 | 15:15:16 | 4.708 | 297.679 | 28.8 | 2891 | 17.4 |  | 99.0797 | 20.7077 | 206.1343 | 154.6138 | 7.8188 | 244.3373
2010-08-12 | 15:15:26 | 4.875 | 296.353 | 28.85 | 2893 | 17.4 |  | 98.6381 | 20.6154 | 205.2155 | 153.9247 | 7.7839 | 243.2483
2010-08-12 | 15:15:36 | 5.042 | 295.824 | 28.87 | 2893 | 17.4 |  | 98.4621 | 20.5786 | 204.8494 | 153.6501 | 7.7701 | 242.8143
2010-08-12 | 15:15:47 | 5.209 | 294.77 | 28.91 | 2892 | 17.4 |  | 98.1112 | 20.5052 | 204.1193 | 153.1025 | 7.7424 | 241.9489
2010-08-12 | 15:15:57 | 5.376 | 293.982 | 28.94 | 2893 | 17.4 |  | 97.8489 | 20.4504 | 203.5738 | 152.6933 | 7.7217 | 241.3022
2010-08-12 | 15:16:07 | 5.542 | 294.714 | 28.87 | 2897 | 17.5 |  | 98.2956 | 20.5438 | 204.477 | 153.3708 | 7.7422 | 241.9442
2010-08-12 | 15:16:17 | 5.709 | 297.625 | 28.76 | 2895 | 17.5 |  | 99.2665 | 20.7467 | 206.4968 | 154.8857 | 7.8187 | 244.3341
2010-08-12 | 15:16:27 | 5.876 | 293.14 | 28.93 | 2903 | 17.5 |  | 97.7706 | 20.4341 | 203.385 | 152.5517 | 7.7009 | 240.6521
2010-08-12 | 15:16:37 | 6.043 | 295.769 | 28.83 | 2900 | 17.5 |  | 98.6474 | 20.6173 | 205.2088 | 153.9197 | 7.7699 | 242.8101
2010-08-12 | 15:16:47 | 6.21 | 293.402 | 28.92 | 2893 | 17.5 |  | 97.8579 | 20.4523 | 203.5665 | 152.6878 | 7.7077 | 240.8669
2010-08-12 | 15:16:57 | 6.377 | 295.241 | 28.85 | 2904 | 17.5 |  | 98.4713 | 20.5805 | 204.8425 | 153.6449 | 7.7561 | 242.3767
2010-08-12 | 15:17:06 | 6.527 | 296.353 | 28.85 | 2910 | 17.4 |  | 98.6381 | 20.6154 | 205.2155 | 153.9247 | 7.7839 | 243.2483
2010-08-12 | 15:17:16 | 6.694 | 296.617 | 28.84 | 2904 | 17.4 |  | 98.7262 | 20.6338 | 205.3989 | 154.0623 | 7.7909 | 243.4656
2010-08-12 | 15:17:26 | 6.861 | 294.244 | 28.93 | 2903 | 17.4 |  | 97.9363 | 20.4687 | 203.7554 | 152.8296 | 7.7286 | 241.5176
2010-08-12 | 15:17:36 | 7.028 | 297.946 | 28.79 | 2912 | 17.4 |  | 99.1683 | 20.7262 | 206.3186 | 154.7521 | 7.8258 | 244.5558
2010-08-12 | 15:17:46 | 7.195 | 293.982 | 28.94 | 2911 | 17.4 |  | 97.8489 | 20.4504 | 203.5738 | 152.6933 | 7.7217 | 241.3022
2010-08-12 | 15:17:56 | 7.362 | 296.088 | 28.86 | 2911 | 17.4 |  | 98.55 | 20.597 | 205.0323 | 153.7873 | 7.777 | 243.0312
2010-08-12 | 15:18:06 | 7.529 | 295.56 | 28.88 | 2910 | 17.4 |  | 98.3742 | 20.5602 | 204.6666 | 153.513 | 7.7631 | 242.5976
2010-08-12 | 15:18:16 | 7.696 | 293.982 | 28.94 | 2911 | 17.4 |  | 97.8489 | 20.4504 | 203.5738 | 152.6933 | 7.7217 | 241.3022
2010-08-12 | 15:18:26 | 7.862 | 299.013 | 28.75 | 2917 | 17.4 |  | 99.5236 | 20.8004 | 207.0578 | 155.3066 | 7.8538 | 245.432
2010-08-12 | 15:18:36 | 8.029 | 293.196 | 28.97 | 2918 | 17.4 |  | 97.5875 | 20.3958 | 203.0299 | 152.2854 | 7.701 | 240.6576
2010-08-12 | 15:18:46 | 8.196 | 299.281 | 28.74 | 2919 | 17.4 |  | 99.6126 | 20.819 | 207.2431 | 155.4455 | 7.8609 | 245.6516
2010-08-12 | 15:18:56 | 8.363 | 293.982 | 28.94 | 2919 | 17.4 |  | 97.8489 | 20.4504 | 203.5738 | 152.6933 | 7.7217 | 241.3022
2010-08-12 | 15:19:06 | 8.53 | 293.982 | 28.94 | 2928 | 17.4 |  | 97.8489 | 20.4504 | 203.5738 | 152.6933 | 7.7217 | 241.3022
2010-08-12 | 15:19:16 | 8.697 | 293.196 | 28.97 | 2917 | 17.4 |  | 97.5875 | 20.3958 | 203.0299 | 152.2854 | 7.701 | 240.6576
2010-08-12 | 15:19:26 | 8.864 | 292.413 | 29 | 2918 | 17.4 |  | 97.3269 | 20.3413 | 202.4877 | 151.8787 | 7.6805 | 240.0149
2010-08-12 | 15:19:36 | 9.031 | 295.033 | 28.9 | 2927 | 17.4 |  | 98.1988 | 20.5235 | 204.3015 | 153.2392 | 7.7493 | 242.1649
2010-08-12 | 15:19:46 | 9.198 | 297.413 | 28.81 | 2921 | 17.4 |  | 98.9912 | 20.6892 | 205.9501 | 154.4757 | 7.8118 | 244.119
2010-08-12 | 15:19:56 | 9.365 | 295.824 | 28.87 | 2923 | 17.4 |  | 98.4621 | 20.5786 | 204.8494 | 153.6501 | 7.7701 | 242.8143
2010-08-12 | 15:20:06 | 9.531 | 297.679 | 28.8 | 2926 | 17.4 |  | 99.0797 | 20.7077 | 206.1343 | 154.6138 | 7.8188 | 244.3373
2010-08-12 | 15:20:16 | 9.698 | 293.458 | 28.96 | 2919 | 17.4 |  | 97.6746 | 20.414 | 203.211 | 152.4212 | 7.7079 | 240.8723
2010-08-12 | 15:20:26 | 9.865 | 293.72 | 28.95 | 2923 | 17.4 |  | 97.7617 | 20.4322 | 203.3923 | 152.5572 | 7.7148 | 241.0871
2010-08-12 | 15:20:36 | 10.032 | 292.935 | 28.98 | 2929 | 17.4 |  | 97.5006 | 20.3776 | 202.849 | 152.1497 | 7.6942 | 240.4432
2010-08-12 | 15:20:46 | 10.199 | 291.893 | 29.02 | 2931 | 17.4 |  | 97.1537 | 20.3051 | 202.1272 | 151.6083 | 7.6668 | 239.5876
2010-08-12 | 15:20:56 | 10.366 | 290.337 | 29.08 | 2923 | 17.4 |  | 96.6359 | 20.1969 | 201.0501 | 150.8004 | 7.6259 | 238.3109
2010-08-12 | 15:21:06 | 10.533 | 296.088 | 28.86 | 2931 | 17.4 |  | 98.55 | 20.597 | 205.0323 | 153.7873 | 7.777 | 243.0312
2010-08-12 | 15:21:16 | 10.7 | 295.824 | 28.87 | 2930 | 17.4 |  | 98.4621 | 20.5786 | 204.8494 | 153.6501 | 7.7701 | 242.8143
2010-08-12 | 15:21:26 | 10.867 | 294.77 | 28.91 | 2930 | 17.4 |  | 98.1112 | 20.5052 | 204.1193 | 153.1025 | 7.7424 | 241.9489
2010-08-12 | 15:21:36 | 11.034 | 295.033 | 28.9 | 2926 | 17.4 |  | 98.1988 | 20.5235 | 204.3015 | 153.2392 | 7.7493 | 242.1649
2010-08-12 | 15:21:46 | 11.2 | 294.507 | 28.92 | 2930 | 17.4 |  | 98.0237 | 20.4869 | 203.9373 | 152.966 | 7.7355 | 241.7331
2010-08-12 | 15:21:56 | 11.368 | 291.633 | 29.03 | 2929 | 17.4 |  | 97.0671 | 20.287 | 201.9472 | 151.4733 | 7.66 | 239.3743
2010-08-12 | 15:22:06 | 11.534 | 292.413 | 29 | 2923 | 17.4 |  | 97.3269 | 20.3413 | 202.4877 | 151.8787 | 7.6805 | 240.0149
2010-08-12 | 15:22:16 | 11.701 | 293.458 | 28.96 | 2924 | 17.4 |  | 97.6746 | 20.414 | 203.211 | 152.4212 | 7.7079 | 240.8723
2010-08-12 | 15:22:26 | 11.868 | 292.674 | 28.99 | 2927 | 17.4 |  | 97.4137 | 20.3595 | 202.6683 | 152.0141 | 7.6873 | 240.2289
2010-08-12 | 15:22:36 | 12.035 | 292.879 | 28.94 | 2926 | 17.5 |  | 97.6834 | 20.4158 | 203.2036 | 152.4157 | 7.694 | 240.4375
2010-08-12 | 15:22:46 | 12.202 | 291.576 | 28.99 | 2920 | 17.5 |  | 97.2489 | 20.325 | 202.2997 | 151.7377 | 7.6598 | 239.368
2010-08-12 | 15:22:56 | 12.369 | 290.797 | 29.02 | 2924 | 17.5 |  | 96.9893 | 20.2708 | 201.7596 | 151.3326 | 7.6393 | 238.729
2010-08-12 | 15:23:06 | 12.536 | 295.241 | 28.85 | 2930 | 17.5 |  | 98.4713 | 20.5805 | 204.8425 | 153.6449 | 7.7561 | 242.3767
2010-08-12 | 15:23:16 | 12.703 | 293.926 | 28.9 | 2926 | 17.5 |  | 98.0327 | 20.4888 | 203.9301 | 152.9606 | 7.7215 | 241.2972
2010-08-12 | 15:23:26 | 12.87 | 292.096 | 28.97 | 2922 | 17.5 |  | 97.4225 | 20.3613 | 202.6607 | 152.0085 | 7.6734 | 239.7952
2010-08-12 | 15:23:36 | 13.037 | 291.836 | 28.98 | 2927 | 17.5 |  | 97.3356 | 20.3432 | 202.4801 | 151.873 | 7.6666 | 239.5815
2010-08-12 | 15:23:46 | 13.203 | 295.824 | 28.87 | 2921 | 17.4 |  | 98.4621 | 20.5786 | 204.8494 | 153.6501 | 7.7701 | 242.8143
2010-08-12 | 15:23:56 | 13.37 | 292.935 | 28.98 | 2921 | 17.4 |  | 97.5006 | 20.3776 | 202.849 | 152.1497 | 7.6942 | 240.4432
2010-08-12 | 15:24:06 | 13.537 | 295.033 | 28.9 | 2927 | 17.4 |  | 98.1988 | 20.5235 | 204.3015 | 153.2392 | 7.7493 | 242.1649
2010-08-12 | 15:24:16 | 13.704 | 293.982 | 28.94 | 2923 | 17.4 |  | 97.8489 | 20.4504 | 203.5738 | 152.6933 | 7.7217 | 241.3022
2010-08-12 | 15:24:26 | 13.871 | 293.72 | 28.95 | 2922 | 17.4 |  | 97.7617 | 20.4322 | 203.3923 | 152.5572 | 7.7148 | 241.0871
2010-08-12 | 15:24:36 | 14.038 | 295.033 | 28.9 | 2923 | 17.4 |  | 98.1988 | 20.5235 | 204.3015 | 153.2392 | 7.7493 | 242.1649
2010-08-12 | 15:24:46 | 14.205 | 290.596 | 29.07 | 2922 | 17.4 |  | 96.722 | 20.2149 | 201.2292 | 150.9347 | 7.6327 | 238.5232
2010-08-12 | 15:24:56 | 14.372 | 292.674 | 28.99 | 2921 | 17.4 |  | 97.4137 | 20.3595 | 202.6683 | 152.0141 | 7.6873 | 240.2289
2010-08-12 | 15:25:06 | 14.539 | 292.153 | 29.01 | 2920 | 17.4 |  | 97.2403 | 20.3232 | 202.3074 | 151.7434 | 7.6736 | 239.8012
2010-08-12 | 15:25:16 | 14.706 | 292.413 | 29 | 2924 | 17.4 |  | 97.3269 | 20.3413 | 202.4877 | 151.8787 | 7.6805 | 240.0149
2010-08-12 | 15:25:26 | 14.873 | 293.196 | 28.97 | 2924 | 17.4 |  | 97.5875 | 20.3958 | 203.0299 | 152.2854 | 7.701 | 240.6576
2010-08-12 | 15:25:36 | 15.039 | 291.893 | 29.02 | 2927 | 17.4 |  | 97.1537 | 20.3051 | 202.1272 | 151.6083 | 7.6668 | 239.5876
2010-08-12 | 15:25:46 | 15.206 | 288.535 | 29.15 | 2924 | 17.4 |  | 96.036 | 20.0715 | 199.8019 | 149.8642 | 7.5786 | 236.8314
2010-08-12 | 15:25:56 | 15.373 | 289.563 | 29.11 | 2926 | 17.4 |  | 96.3783 | 20.1431 | 200.5141 | 150.3983 | 7.6056 | 237.6755
2010-08-12 | 15:26:06 | 15.54 | 292.413 | 29 | 2924 | 17.4 |  | 97.3269 | 20.3413 | 202.4877 | 151.8787 | 7.6805 | 240.0149
2010-08-12 | 15:26:16 | 15.707 | 290.855 | 29.06 | 2926 | 17.4 |  | 96.8082 | 20.2329 | 201.4084 | 151.0692 | 7.6395 | 238.7356
2010-08-12 | 15:26:26 | 15.874 | 293.982 | 28.94 | 2920 | 17.4 |  | 97.8489 | 20.4504 | 203.5738 | 152.6933 | 7.7217 | 241.3022
2010-08-12 | 15:26:36 | 16.041 | 290.337 | 29.08 | 2925 | 17.4 |  | 96.6359 | 20.1969 | 201.0501 | 150.8004 | 7.6259 | 238.3109
2010-08-12 | 15:26:46 | 16.208 | 292.153 | 29.01 | 2926 | 17.4 |  | 97.2403 | 20.3232 | 202.3074 | 151.7434 | 7.6736 | 239.8012
2010-08-12 | 15:26:56 | 16.375 | 291.114 | 29.05 | 2919 | 17.4 |  | 96.8944 | 20.2509 | 201.5878 | 151.2037 | 7.6463 | 238.9483
2010-08-12 | 15:27:06 | 16.541 | 295.033 | 28.9 | 2916 | 17.4 |  | 98.1988 | 20.5235 | 204.3015 | 153.2392 | 7.7493 | 242.1649
2010-08-12 | 15:27:16 | 16.708 | 295.296 | 28.89 | 2914 | 17.4 |  | 98.2864 | 20.5419 | 204.484 | 153.376 | 7.7562 | 242.3811
2010-08-12 | 15:27:26 | 16.875 | 295.824 | 28.87 | 2916 | 17.4 |  | 98.4621 | 20.5786 | 204.8494 | 153.6501 | 7.7701 | 242.8143
2010-08-12 | 15:27:37 | 17.042 | 295.296 | 28.89 | 2919 | 17.4 |  | 98.2864 | 20.5419 | 204.484 | 153.376 | 7.7562 | 242.3811
2010-08-12 | 15:27:47 | 17.209 | 295.296 | 28.89 | 2916 | 17.4 |  | 98.2864 | 20.5419 | 204.484 | 153.376 | 7.7562 | 242.3811
2010-08-12 | 15:27:57 | 17.376 | 296.088 | 28.86 | 2914 | 17.4 |  | 98.55 | 20.597 | 205.0323 | 153.7873 | 7.777 | 243.0312
2010-08-12 | 15:28:07 | 17.543 | 294.507 | 28.92 | 2912 | 17.4 |  | 98.0237 | 20.4869 | 203.9373 | 152.966 | 7.7355 | 241.7331
2010-08-12 | 15:28:17 | 17.71 | 292.674 | 28.99 | 2914 | 17.4 |  | 97.4137 | 20.3595 | 202.6683 | 152.0141 | 7.6873 | 240.2289
2010-08-12 | 15:28:27 | 17.877 | 294.507 | 28.92 | 2915 | 17.4 |  | 98.0237 | 20.4869 | 203.9373 | 152.966 | 7.7355 | 241.7331
2010-08-12 | 15:28:37 | 18.044 | 294.507 | 28.92 | 2913 | 17.4 |  | 98.0237 | 20.4869 | 203.9373 | 152.966 | 7.7355 | 241.7331
2010-08-12 | 15:28:46 | 18.194 | 297.148 | 28.82 | 2912 | 17.4 |  | 98.9028 | 20.6707 | 205.7662 | 154.3378 | 7.8048 | 243.901
2010-08-12 | 15:28:56 | 18.361 | 291.893 | 29.02 | 2910 | 17.4 |  | 97.1537 | 20.3051 | 202.1272 | 151.6083 | 7.6668 | 239.5876
2010-08-12 | 15:29:06 | 18.528 | 294.714 | 28.87 | 2911 | 17.5 |  | 98.2956 | 20.5438 | 204.477 | 153.3708 | 7.7422 | 241.9442
2010-08-12 | 15:29:16 | 18.695 | 297.625 | 28.76 | 2910 | 17.5 |  | 99.2665 | 20.7467 | 206.4968 | 154.8857 | 7.8187 | 244.3341
2010-08-12 | 15:29:26 | 18.861 | 298.425 | 28.73 | 2912 | 17.5 |  | 99.5333 | 20.8025 | 207.0516 | 155.3019 | 7.8397 | 244.9907
2010-08-12 | 15:29:36 | 19.029 | 298.158 | 28.74 | 2909 | 17.5 |  | 99.4442 | 20.7838 | 206.8665 | 155.1631 | 7.8327 | 244.7716
2010-08-12 | 15:29:46 | 19.195 | 294.188 | 28.89 | 2906 | 17.5 |  | 98.1202 | 20.5071 | 204.1122 | 153.0972 | 7.7284 | 241.5126
2010-08-12 | 15:29:56 | 19.362 | 296.033 | 28.82 | 2912 | 17.5 |  | 98.7355 | 20.6357 | 205.3922 | 154.0572 | 7.7769 | 243.0271
2010-08-12 | 15:30:06 | 19.529 | 296.828 | 28.79 | 2908 | 17.5 |  | 99.0006 | 20.6911 | 205.9436 | 154.4708 | 7.7977 | 243.6796
2010-08-12 | 15:30:16 | 19.696 | 297.625 | 28.76 | 2910 | 17.5 |  | 99.2665 | 20.7467 | 206.4968 | 154.8857 | 7.8187 | 244.3341
2010-08-12 | 15:30:26 | 19.863 | 301.109 | 28.63 | 2908 | 17.5 |  | 100.4285 | 20.9895 | 208.9139 | 156.6987 | 7.9102 | 247.1941
2010-08-12 | 15:30:36 | 20.03 | 295.769 | 28.83 | 2905 | 17.5 |  | 98.6474 | 20.6173 | 205.2088 | 153.9197 | 7.7699 | 242.8101
2010-08-12 | 15:30:46 | 20.197 | 299.763 | 28.68 | 2904 | 17.5 |  | 99.9797 | 20.8958 | 207.9803 | 155.9985 | 7.8749 | 246.0895
2010-08-12 | 15:30:56 | 20.364 | 297.359 | 28.77 | 2899 | 17.5 |  | 99.1778 | 20.7282 | 206.3122 | 154.7473 | 7.8117 | 244.1157
2010-08-12 | 15:31:06 | 20.531 | 298.158 | 28.74 | 2900 | 17.5 |  | 99.4442 | 20.7838 | 206.8665 | 155.1631 | 7.8327 | 244.7716
2010-08-12 | 15:31:16 | 20.698 | 302.733 | 28.57 | 2903 | 17.5 |  | 100.9701 | 21.1027 | 210.0406 | 157.5438 | 7.9529 | 248.5273
2010-08-12 | 15:31:26 | 20.864 | 299.013 | 28.75 | 2903 | 17.4 |  | 99.5236 | 20.8004 | 207.0578 | 155.3066 | 7.8538 | 245.432
2010-08-12 | 15:31:36 | 21.031 | 302.513 | 28.62 | 2902 | 17.4 |  | 100.6886 | 21.0439 | 209.4817 | 157.1246 | 7.9458 | 248.3051
2010-08-12 | 15:31:46 | 21.198 | 304.418 | 28.55 | 2893 | 17.4 |  | 101.3226 | 21.1764 | 210.8006 | 158.1139 | 7.9958 | 249.8684
2010-08-12 | 15:31:56 | 21.365 | 302.513 | 28.62 | 2901 | 17.4 |  | 100.6886 | 21.0439 | 209.4817 | 157.1246 | 7.9458 | 248.3051
2010-08-12 | 15:32:06 | 21.532 | 300.892 | 28.68 | 2894 | 17.4 |  | 100.1489 | 20.9311 | 208.3589 | 156.2824 | 7.9032 | 246.9742
2010-08-12 | 15:32:16 | 21.699 | 302.513 | 28.62 | 2896 | 17.4 |  | 100.6886 | 21.0439 | 209.4817 | 157.1246 | 7.9458 | 248.3051
2010-08-12 | 15:32:26 | 21.866 | 302.785 | 28.61 | 2893 | 17.4 |  | 100.7789 | 21.0628 | 209.6695 | 157.2655 | 7.9529 | 248.5277
2010-08-12 | 15:32:36 | 22.033 | 306.664 | 28.51 | 2893 | 17.3 |  | 101.8591 | 21.2885 | 211.9436 | 158.9712 | 8.0534 | 251.6691
2010-08-12 | 15:32:46 | 22.2 | 305.291 | 28.56 | 2889 | 17.3 |  | 101.4031 | 21.1933 | 210.9949 | 158.2596 | 8.0174 | 250.5426
2010-08-12 | 15:32:56 | 22.367 | 304.471 | 28.59 | 2889 | 17.3 |  | 101.1307 | 21.1363 | 210.4281 | 157.8345 | 7.9958 | 249.8695
2010-08-12 | 15:33:06 | 22.534 | 308.044 | 28.46 | 2889 | 17.3 |  | 102.3174 | 21.3843 | 212.8973 | 159.6865 | 8.0896 | 252.8015
2010-08-12 | 15:33:16 | 22.7 | 304.145 | 28.56 | 2886 | 17.4 |  | 101.2317 | 21.1574 | 210.6116 | 157.9721 | 7.9886 | 249.6444
2010-08-12 | 15:33:26 | 22.867 | 306.062 | 28.49 | 2882 | 17.4 |  | 101.8697 | 21.2908 | 211.9388 | 158.9676 | 8.039 | 251.2176
2010-08-12 | 15:33:36 | 23.034 | 305.513 | 28.51 | 2880 | 17.4 |  | 101.6869 | 21.2526 | 211.5586 | 158.6824 | 8.0245 | 250.7669
2010-08-12 | 15:33:46 | 23.201 | 305.787 | 28.5 | 2880 | 17.4 |  | 101.7782 | 21.2717 | 211.7486 | 158.825 | 8.0317 | 250.9922
2010-08-12 | 15:33:56 | 23.368 | 306.612 | 28.47 | 2875 | 17.4 |  | 102.0528 | 21.329 | 212.3198 | 159.2534 | 8.0534 | 251.6692
2010-08-12 | 15:34:06 | 23.535 | 309.659 | 28.36 | 2874 | 17.4 |  | 103.0669 | 21.541 | 214.4298 | 160.836 | 8.1334 | 254.1702
2010-08-12 | 15:34:16 | 23.702 | 306.062 | 28.49 | 2868 | 17.4 |  | 101.8697 | 21.2908 | 211.9388 | 158.9676 | 8.039 | 251.2176
2010-08-12 | 15:34:26 | 23.869 | 309.938 | 28.35 | 2863 | 17.4 |  | 103.1597 | 21.5604 | 214.6228 | 160.9808 | 8.1408 | 254.399
2010-08-12 | 15:34:36 | 24.036 | 309.659 | 28.36 | 2863 | 17.4 |  | 103.0669 | 21.541 | 214.4298 | 160.836 | 8.1334 | 254.1702
2010-08-12 | 15:34:46 | 24.203 | 309.938 | 28.35 | 2871 | 17.4 |  | 103.1597 | 21.5604 | 214.6228 | 160.9808 | 8.1408 | 254.399
2010-08-12 | 15:34:56 | 24.369 | 309.38 | 28.37 | 2862 | 17.4 |  | 102.9743 | 21.5216 | 214.2369 | 160.6914 | 8.1261 | 253.9416
2010-08-12 | 15:35:06 | 24.536 | 308.824 | 28.39 | 2863 | 17.4 |  | 102.7892 | 21.4829 | 213.8519 | 160.4026 | 8.1115 | 253.4852
2010-08-12 | 15:35:16 | 24.703 | 310.776 | 28.32 | 2857 | 17.4 |  | 103.4387 | 21.6187 | 215.2031 | 161.416 | 8.1628 | 255.0869
2010-08-12 | 15:35:26 | 24.87 | 312.46 | 28.26 | 2859 | 17.4 |  | 103.9992 | 21.7358 | 216.3693 | 162.2907 | 8.207 | 256.4692
2010-08-12 | 15:35:36 | 25.037 | 314.722 | 28.18 | 2850 | 17.4 |  | 104.7521 | 21.8932 | 217.9357 | 163.4657 | 8.2664 | 258.326
2010-08-12 | 15:35:46 | 25.204 | 313.024 | 28.24 | 2863 | 17.4 |  | 104.1868 | 21.775 | 216.7596 | 162.5836 | 8.2218 | 256.9319
2010-08-12 | 15:35:56 | 25.371 | 313.588 | 28.22 | 2846 | 17.4 |  | 104.3749 | 21.8143 | 217.1508 | 162.877 | 8.2367 | 257.3956
2010-08-12 | 15:36:06 | 25.538 | 315.29 | 28.16 | 2846 | 17.4 |  | 104.9414 | 21.9327 | 218.3294 | 163.761 | 8.2814 | 258.7926
2010-08-12 | 15:36:16 | 25.705 | 314.722 | 28.18 | 2843 | 17.4 |  | 104.7521 | 21.8932 | 217.9357 | 163.4657 | 8.2664 | 258.326
2010-08-12 | 15:36:26 | 25.872 | 316.146 | 28.13 | 2838 | 17.4 |  | 105.226 | 21.9922 | 218.9216 | 164.2051 | 8.3038 | 259.4945
2010-08-12 | 15:36:36 | 26.038 | 312.46 | 28.26 | 2834 | 17.4 |  | 103.9992 | 21.7358 | 216.3693 | 162.2907 | 8.207 | 256.4692
2010-08-12 | 15:36:46 | 26.206 | 316.431 | 28.12 | 2833 | 17.4 |  | 105.321 | 22.0121 | 219.1194 | 164.3535 | 8.3113 | 259.7289
2010-08-12 | 15:36:56 | 26.372 | 317.864 | 28.07 | 2828 | 17.4 |  | 105.7979 | 22.1118 | 220.1115 | 165.0977 | 8.349 | 260.905
2010-08-12 | 15:37:06 | 26.539 | 313.825 | 28.17 | 2827 | 17.5 |  | 104.6697 | 21.876 | 217.7367 | 163.3164 | 8.2443 | 257.6336
2010-08-12 | 15:37:16 | 26.706 | 315.815 | 28.1 | 2820 | 17.5 |  | 105.3335 | 22.0147 | 219.1174 | 164.352 | 8.2966 | 259.2672
2010-08-12 | 15:37:26 | 26.873 | 316.387 | 28.08 | 2818 | 17.5 |  | 105.524 | 22.0545 | 219.5137 | 164.6493 | 8.3116 | 259.7362
2010-08-12 | 15:37:36 | 27.04 | 317.533 | 28.04 | 2819 | 17.5 |  | 105.9063 | 22.1344 | 220.309 | 165.2458 | 8.3417 | 260.6772
2010-08-12 | 15:37:46 | 27.207 | 316.673 | 28.07 | 2814 | 17.5 |  | 105.6194 | 22.0745 | 219.7122 | 164.7982 | 8.3191 | 259.9711
2010-08-12 | 15:37:56 | 27.374 | 319.84 | 27.96 | 2811 | 17.5 |  | 106.6758 | 22.2952 | 221.9098 | 166.4465 | 8.4023 | 262.5713
2010-08-12 | 15:38:06 | 27.541 | 318.684 | 28 | 2806 | 17.5 |  | 106.2902 | 22.2147 | 221.1077 | 165.8449 | 8.3719 | 261.6223
2010-08-12 | 15:38:16 | 27.708 | 320.13 | 27.95 | 2809 | 17.5 |  | 106.7725 | 22.3154 | 222.1109 | 166.5973 | 8.4099 | 262.8093
2010-08-12 | 15:38:26 | 27.875 | 323.632 | 27.83 | 2799 | 17.5 |  | 107.9404 | 22.5596 | 224.5405 | 168.4197 | 8.5019 | 265.6841
2010-08-12 | 15:38:36 | 28.041 | 320.71 | 27.93 | 2799 | 17.5 |  | 106.9661 | 22.3559 | 222.5136 | 166.8994 | 8.4251 | 263.2858
2010-08-12 | 15:38:46 | 28.208 | 323.632 | 27.83 | 2795 | 17.5 |  | 107.9404 | 22.5596 | 224.5405 | 168.4197 | 8.5019 | 265.6841
2010-08-12 | 15:38:56 | 28.375 | 322.459 | 27.87 | 2794 | 17.5 |  | 107.5495 | 22.4778 | 223.7271 | 167.8096 | 8.4711 | 264.7217
2010-08-12 | 15:39:07 | 28.542 | 325.104 | 27.78 | 2793 | 17.5 |  | 108.4316 | 22.6622 | 225.5621 | 169.186 | 8.5406 | 266.8929
2010-08-12 | 15:39:17 | 28.709 | 321.292 | 27.91 | 2788 | 17.5 |  | 107.1601 | 22.3965 | 222.9173 | 167.2022 | 8.4404 | 263.7634
2010-08-12 | 15:39:27 | 28.876 | 327.217 | 27.75 | 2784 | 17.4 |  | 108.9111 | 22.7624 | 226.5884 | 169.9558 | 8.5946 | 268.5822
2010-08-12 | 15:39:37 | 29.043 | 322.793 | 27.9 | 2785 | 17.4 |  | 107.4385 | 22.4546 | 223.5247 | 167.6578 | 8.4784 | 264.9507
2010-08-12 | 15:39:47 | 29.21 | 327.813 | 27.73 | 2780 | 17.4 |  | 109.1092 | 22.8038 | 227.0007 | 170.265 | 8.6103 | 269.0709
2010-08-12 | 15:39:57 | 29.377 | 323.966 | 27.86 | 2781 | 17.4 |  | 107.8289 | 22.5362 | 224.3369 | 168.267 | 8.5092 | 265.9134
2010-08-12 | 15:40:07 | 29.544 | 329.307 | 27.68 | 2774 | 17.4 |  | 109.6065 | 22.9078 | 228.0352 | 171.0409 | 8.6495 | 270.2971
2010-08-12 | 15:40:17 | 29.711 | 326.327 | 27.78 | 2767 | 17.4 |  | 108.6147 | 22.7005 | 225.9717 | 169.4932 | 8.5712 | 267.8512
2010-08-12 | 15:40:26 | 29.861 | 328.708 | 27.7 | 2767 | 17.4 |  | 109.4072 | 22.8661 | 227.6207 | 170.73 | 8.6338 | 269.8058
2010-08-12 | 15:40:36 | 30.028 | 330.546 | 27.68 | 2761 | 17.3 |  | 109.7916 | 22.9464 | 228.4492 | 171.3515 | 8.6806 | 271.2684
2010-08-12 | 15:40:46 | 30.195 | 329.047 | 27.73 | 2759 | 17.3 |  | 109.2936 | 22.8424 | 227.4129 | 170.5742 | 8.6412 | 270.0378
2010-08-12 | 15:40:56 | 30.361 | 327.555 | 27.78 | 2760 | 17.3 |  | 108.7982 | 22.7388 | 226.3821 | 169.801 | 8.602 | 268.8139
2010-08-12 | 15:41:06 | 30.528 | 331.148 | 27.66 | 2759 | 17.3 |  | 109.9916 | 22.9882 | 228.8653 | 171.6636 | 8.6964 | 271.7625
2010-08-12 | 15:41:16 | 30.695 | 330.847 | 27.67 | 2755 | 17.3 |  | 109.8915 | 22.9673 | 228.6572 | 171.5075 | 8.6885 | 271.5153
2010-08-12 | 15:41:26 | 30.862 | 332.962 | 27.6 | 2751 | 17.3 |  | 110.5941 | 23.1142 | 230.119 | 172.604 | 8.744 | 273.2512
2010-08-12 | 15:41:36 | 31.029 | 332.054 | 27.63 | 2751 | 17.3 |  | 110.2924 | 23.0511 | 229.4912 | 172.133 | 8.7202 | 272.5057
2010-08-12 | 15:41:46 | 31.196 | 336.284 | 27.45 | 2741 | 17.4 |  | 111.9287 | 23.3931 | 232.8667 | 174.6648 | 8.8328 | 276.024
2010-08-12 | 15:41:56 | 31.363 | 334.142 | 27.52 | 2740 | 17.4 |  | 111.2158 | 23.2441 | 231.3834 | 173.5523 | 8.7765 | 274.2659
2010-08-12 | 15:42:06 | 31.53 | 335.364 | 27.48 | 2744 | 17.4 |  | 111.6225 | 23.3291 | 232.2296 | 174.187 | 8.8086 | 275.2689
2010-08-12 | 15:42:16 | 31.697 | 334.142 | 27.52 | 2739 | 17.4 |  | 111.2158 | 23.2441 | 231.3834 | 173.5523 | 8.7765 | 274.2659
2010-08-12 | 15:42:26 | 31.863 | 336.591 | 27.44 | 2738 | 17.4 |  | 112.031 | 23.4145 | 233.0795 | 174.8245 | 8.8408 | 276.2763
2010-08-12 | 15:42:36 | 32.03 | 338.752 | 27.37 | 2730 | 17.4 |  | 112.7502 | 23.5648 | 234.5757 | 175.9467 | 8.8976 | 278.0498
2010-08-12 | 15:42:46 | 32.198 | 337.207 | 27.42 | 2732 | 17.4 |  | 112.236 | 23.4573 | 233.5058 | 175.1442 | 8.857 | 276.7816
2010-08-12 | 15:42:56 | 32.364 | 335.977 | 27.46 | 2723 | 17.4 |  | 111.8266 | 23.3718 | 232.6541 | 174.5054 | 8.8247 | 275.772
2010-08-12 | 15:43:06 | 32.531 | 338.133 | 27.39 | 2718 | 17.4 |  | 112.5442 | 23.5217 | 234.147 | 175.6252 | 8.8813 | 277.5417
2010-08-12 | 15:43:16 | 32.698 | 340.617 | 27.31 | 2720 | 17.4 |  | 113.371 | 23.6945 | 235.8672 | 176.9155 | 8.9466 | 279.5807
2010-08-12 | 15:43:26 | 32.865 | 339.683 | 27.34 | 2719 | 17.4 |  | 113.0601 | 23.6296 | 235.2204 | 176.4303 | 8.922 | 278.814
2010-08-12 | 15:43:36 | 33.032 | 338.752 | 27.37 | 2715 | 17.4 |  | 112.7502 | 23.5648 | 234.5757 | 175.9467 | 8.8976 | 278.0498
2010-08-12 | 15:43:46 | 33.199 | 339.062 | 27.36 | 2705 | 17.4 |  | 112.8534 | 23.5864 | 234.7903 | 176.1077 | 8.9057 | 278.3042
2010-08-12 | 15:43:56 | 33.366 | 341.554 | 27.28 | 2711 | 17.4 |  | 113.6829 | 23.7597 | 236.5162 | 177.4022 | 8.9712 | 280.3499
2010-08-12 | 15:44:06 | 33.533 | 341.867 | 27.27 | 2710 | 17.4 |  | 113.7871 | 23.7815 | 236.733 | 177.5648 | 8.9794 | 280.6069
2010-08-12 | 15:44:16 | 33.699 | 342.808 | 27.24 | 2705 | 17.4 |  | 114.1004 | 23.847 | 237.3848 | 178.0537 | 9.0041 | 281.3795
2010-08-12 | 15:44:26 | 33.866 | 342.494 | 27.25 | 2700 | 17.4 |  | 113.9959 | 23.8251 | 237.1673 | 177.8906 | 8.9959 | 281.1217
2010-08-12 | 15:44:36 | 34.033 | 342.808 | 27.24 | 2705 | 17.4 |  | 114.1004 | 23.847 | 237.3848 | 178.0537 | 9.0041 | 281.3795
2010-08-12 | 15:44:46 | 34.2 | 347.562 | 27.09 | 2695 | 17.4 |  | 115.6825 | 24.1776 | 240.6762 | 180.5225 | 9.129 | 285.281
2010-08-12 | 15:44:56 | 34.367 | 346.605 | 27.12 | 2695 | 17.4 |  | 115.364 | 24.1111 | 240.0136 | 180.0255 | 9.1039 | 284.4955
2010-08-12 | 15:45:06 | 34.534 | 342.494 | 27.25 | 2691 | 17.4 |  | 113.9959 | 23.8251 | 237.1673 | 177.8906 | 8.9959 | 281.1217
2010-08-12 | 15:45:16 | 34.701 | 352.069 | 26.95 | 2692 | 17.4 |  | 117.1828 | 24.4912 | 243.7976 | 182.8638 | 9.2474 | 288.9808
2010-08-12 | 15:45:26 | 34.868 | 345.625 | 27.11 | 2686 | 17.5 |  | 115.2757 | 24.0926 | 239.7996 | 179.865 | 9.0797 | 283.7391
2010-08-12 | 15:45:36 | 35.035 | 347.537 | 27.05 | 2687 | 17.5 |  | 115.9134 | 24.2259 | 241.1261 | 180.86 | 9.1299 | 285.3088
2010-08-12 | 15:45:46 | 35.202 | 344.041 | 27.16 | 2681 | 17.5 |  | 114.7475 | 23.9822 | 238.7007 | 179.0408 | 9.038 | 282.439
2010-08-12 | 15:45:56 | 35.369 | 344.357 | 27.15 | 2674 | 17.5 |  | 114.8529 | 24.0043 | 238.92 | 179.2053 | 9.0463 | 282.6984
2010-08-12 | 15:46:06 | 35.536 | 348.818 | 27.01 | 2674 | 17.5 |  | 116.3409 | 24.3153 | 242.0154 | 181.527 | 9.1636 | 286.361
2010-08-12 | 15:46:16 | 35.702 | 349.783 | 26.98 | 2668 | 17.5 |  | 116.6628 | 24.3825 | 242.6849 | 182.0292 | 9.1889 | 287.1532
2010-08-12 | 15:46:26 | 35.869 | 345.942 | 27.1 | 2669 | 17.5 |  | 115.3817 | 24.1148 | 240.02 | 180.0303 | 9.088 | 284
2010-08-12 | 15:46:36 | 36.036 | 348.497 | 27.02 | 2667 | 17.5 |  | 116.2339 | 24.2929 | 241.7927 | 181.3599 | 9.1551 | 286.0975
2010-08-12 | 15:46:46 | 36.203 | 349.461 | 26.99 | 2666 | 17.5 |  | 116.5554 | 24.3601 | 242.4615 | 181.8616 | 9.1804 | 286.8888
2010-08-12 | 15:46:56 | 36.37 | 350.752 | 26.95 | 2661 | 17.5 |  | 116.9857 | 24.45 | 243.3567 | 182.533 | 9.2143 | 287.948
2010-08-12 | 15:47:06 | 36.537 | 350.106 | 26.97 | 2664 | 17.5 |  | 116.7703 | 24.405 | 242.9086 | 182.1969 | 9.1974 | 287.4178
2010-08-12 | 15:47:16 | 36.704 | 351.399 | 26.93 | 2658 | 17.5 |  | 117.2016 | 24.4951 | 243.8057 | 182.8699 | 9.2313 | 288.4794
2010-08-12 | 15:47:26 | 36.871 | 355.313 | 26.81 | 2655 | 17.5 |  | 118.5069 | 24.7679 | 246.5212 | 184.9066 | 9.3342 | 291.6924
2010-08-12 | 15:47:36 | 37.038 | 347.217 | 27.06 | 2706 | 17.5 |  | 115.8069 | 24.2036 | 240.9044 | 180.6937 | 9.1215 | 285.0464
2010-08-12 | 15:47:46 | 37.205 | 300.301 | 28.66 | 2900 | 17.5 |  | 100.1589 | 20.9332 | 208.3531 | 156.2781 | 7.889 | 246.5306
2010-08-12 | 15:47:56 | 37.372 | 289.306 | 29.12 | 2966 | 17.4 |  | 96.2926 | 20.1252 | 200.3358 | 150.2646 | 7.5989 | 237.4642
2010-08-12 | 15:48:06 | 37.538 | 276.517 | 29.63 | 3030 | 17.4 |  | 92.0359 | 19.2355 | 191.4798 | 143.6221 | 7.2629 | 226.9669
2010-08-12 | 15:48:16 | 37.705 | 271.44 | 29.84 | 3053 | 17.4 |  | 90.3463 | 18.8824 | 187.9645 | 140.9854 | 7.1296 | 222.8002
2010-08-12 | 15:48:26 | 37.872 | 270.247 | 29.89 | 3048 | 17.4 |  | 89.9492 | 18.7994 | 187.1385 | 140.3658 | 7.0983 | 221.821
2010-08-12 | 15:48:36 | 38.039 | 279.219 | 29.52 | 3022 | 17.4 |  | 92.9354 | 19.4235 | 193.3511 | 145.0257 | 7.3339 | 229.1851
2010-08-12 | 15:48:46 | 38.206 | 278.972 | 29.53 | 3030 | 17.4 |  | 92.8532 | 19.4063 | 193.1802 | 144.8974 | 7.3274 | 228.9824
2010-08-12 | 15:48:56 | 38.373 | 271.44 | 29.84 | 3066 | 17.4 |  | 90.3463 | 18.8824 | 187.9645 | 140.9854 | 7.1296 | 222.8002
2010-08-12 | 15:49:06 | 38.54 | 269.535 | 29.92 | 3080 | 17.4 |  | 89.712 | 18.7498 | 186.6448 | 139.9955 | 7.0795 | 221.2359
2010-08-12 | 15:49:16 | 38.707 | 269.535 | 29.92 | 3071 | 17.4 |  | 89.712 | 18.7498 | 186.6448 | 139.9955 | 7.0795 | 221.2359
2010-08-12 | 15:49:26 | 38.874 | 269.061 | 29.94 | 3066 | 17.4 |  | 89.5542 | 18.7168 | 186.3166 | 139.7493 | 7.0671 | 220.8468
2010-08-12 | 15:49:36 | 39.04 | 272.88 | 29.78 | 3057 | 17.4 |  | 90.8254 | 18.9825 | 188.9613 | 141.733 | 7.1674 | 223.9817
2010-08-12 | 15:49:46 | 39.207 | 283.204 | 29.36 | 3007 | 17.4 |  | 94.2618 | 19.7007 | 196.1107 | 147.0956 | 7.4386 | 232.4561
2010-08-12 | 15:49:56 | 39.374 | 278.726 | 29.54 | 2996 | 17.4 |  | 92.7711 | 19.3892 | 193.0093 | 144.7693 | 7.321 | 228.7799
2010-08-12 | 15:50:06 | 39.541 | 280.209 | 29.48 | 3020 | 17.4 |  | 93.265 | 19.4924 | 194.0368 | 145.54 | 7.3599 | 229.9978
2010-08-12 | 15:50:16 | 39.708 | 273.362 | 29.76 | 3069 | 17.4 |  | 90.9857 | 19.016 | 189.2949 | 141.9833 | 7.1801 | 224.3771
2010-08-12 | 15:50:26 | 39.875 | 317.003 | 28.1 | 2865 | 17.4 |  | 105.5115 | 22.0519 | 219.5156 | 164.6507 | 8.3264 | 260.1986
2010-08-12 | 15:50:36 | 40.042 | 329.007 | 27.69 | 2789 | 17.4 |  | 109.5068 | 22.8869 | 227.8278 | 170.8854 | 8.6416 | 270.0513
2010-08-12 | 15:50:47 | 40.209 | 339.062 | 27.36 | 2753 | 17.4 |  | 112.8534 | 23.5864 | 234.7903 | 176.1077 | 8.9057 | 278.3042
2010-08-12 | 15:50:57 | 40.376 | 342.808 | 27.24 | 2742 | 17.4 |  | 114.1004 | 23.847 | 237.3848 | 178.0537 | 9.0041 | 281.3795
2010-08-12 | 15:51:07 | 40.543 | 342.808 | 27.24 | 2734 | 17.4 |  | 114.1004 | 23.847 | 237.3848 | 178.0537 | 9.0041 | 281.3795
2010-08-12 | 15:51:17 | 40.71 | 345.968 | 27.14 | 2732 | 17.4 |  | 115.1522 | 24.0668 | 239.5731 | 179.6951 | 9.0871 | 283.9733
2010-08-12 | 15:51:27 | 40.876 | 345.651 | 27.15 | 2727 | 17.4 |  | 115.0465 | 24.0447 | 239.3532 | 179.5301 | 9.0788 | 283.7127
2010-08-12 | 15:51:37 | 41.043 | 346.286 | 27.13 | 2720 | 17.4 |  | 115.258 | 24.0889 | 239.7932 | 179.8602 | 9.0955 | 284.2343
2010-08-12 | 15:51:47 | 41.21 | 350.129 | 27.01 | 2719 | 17.4 |  | 116.537 | 24.3562 | 242.454 | 181.8559 | 9.1964 | 287.3882
2010-08-12 | 15:51:56 | 41.361 | 346.605 | 27.12 | 2726 | 17.4 |  | 115.364 | 24.1111 | 240.0136 | 180.0255 | 9.1039 | 284.4955
2010-08-12 | 15:52:06 | 41.527 | 348.522 | 27.06 | 2723 | 17.4 |  | 116.002 | 24.2444 | 241.3411 | 181.0212 | 9.1542 | 286.069
2010-08-12 | 15:52:16 | 41.694 | 345.017 | 27.17 | 2718 | 17.4 |  | 114.8355 | 24.0006 | 238.9141 | 179.2008 | 9.0622 | 283.1922
2010-08-12 | 15:52:26 | 41.861 | 347.217 | 27.06 | 2714 | 17.5 |  | 115.8069 | 24.2036 | 240.9044 | 180.6937 | 9.1215 | 285.0464
2010-08-12 | 15:52:36 | 42.028 | 347.217 | 27.06 | 2721 | 17.5 |  | 115.8069 | 24.2036 | 240.9044 | 180.6937 | 9.1215 | 285.0464
2010-08-12 | 15:52:46 | 42.195 | 345.942 | 27.1 | 2717 | 17.5 |  | 115.3817 | 24.1148 | 240.02 | 180.0303 | 9.088 | 284
2010-08-12 | 15:52:56 | 42.362 | 346.261 | 27.09 | 2713 | 17.5 |  | 115.4878 | 24.137 | 240.2408 | 180.1959 | 9.0964 | 284.2612
2010-08-12 | 15:53:06 | 42.529 | 348.497 | 27.02 | 2715 | 17.5 |  | 116.2339 | 24.2929 | 241.7927 | 181.3599 | 9.1551 | 286.0975
2010-08-12 | 15:53:16 | 42.696 | 347.857 | 27.04 | 2717 | 17.5 |  | 116.0201 | 24.2482 | 241.3481 | 181.0264 | 9.1383 | 285.5714
2010-08-12 | 15:53:26 | 42.863 | 348.497 | 27.02 | 2714 | 17.5 |  | 116.2339 | 24.2929 | 241.7927 | 181.3599 | 9.1551 | 286.0975
2010-08-12 | 15:53:36 | 43.03 | 349.461 | 26.99 | 2715 | 17.5 |  | 116.5554 | 24.3601 | 242.4615 | 181.8616 | 9.1804 | 286.8888
2010-08-12 | 15:53:46 | 43.197 | 349.783 | 26.98 | 2717 | 17.5 |  | 116.6628 | 24.3825 | 242.6849 | 182.0292 | 9.1889 | 287.1532
2010-08-12 | 15:53:56 | 43.364 | 349.14 | 27 | 2712 | 17.5 |  | 116.4481 | 24.3377 | 242.2383 | 181.6942 | 9.172 | 286.6248
2010-08-12 | 15:54:06 | 43.53 | 348.818 | 27.01 | 2710 | 17.5 |  | 116.3409 | 24.3153 | 242.0154 | 181.527 | 9.1636 | 286.361
2010-08-12 | 15:54:16 | 43.697 | 346.579 | 27.08 | 2709 | 17.5 |  | 115.5941 | 24.1592 | 240.4617 | 180.3616 | 9.1047 | 284.5227
2010-08-12 | 15:54:26 | 43.864 | 354.329 | 26.84 | 2717 | 17.5 |  | 118.179 | 24.6994 | 245.8389 | 184.3949 | 9.3083 | 290.8851
2010-08-12 | 15:54:36 | 44.031 | 348.497 | 27.02 | 2715 | 17.5 |  | 116.2339 | 24.2929 | 241.7927 | 181.3599 | 9.1551 | 286.0975
2010-08-12 | 15:54:46 | 44.198 | 346.898 | 27.07 | 2711 | 17.5 |  | 115.7004 | 24.1814 | 240.683 | 180.5276 | 9.1131 | 284.7844
2010-08-12 | 15:54:56 | 44.365 | 349.461 | 26.99 | 2708 | 17.5 |  | 116.5554 | 24.3601 | 242.4615 | 181.8616 | 9.1804 | 286.8888
2010-08-12 | 15:55:06 | 44.532 | 348.497 | 27.02 | 2707 | 17.5 |  | 116.2339 | 24.2929 | 241.7927 | 181.3599 | 9.1551 | 286.0975