## convers.+compens.
file name: PSt1_Eq4_Sal |  |  | 2008-03-29
author: Dr. Christian Huber
programm for temperature compensated oxygen calculation of membrane PSt1
Fit of the calibration curve is optimized between 0 and 250 % air-saturation (0 and 388 Torr)
enter your values here
input of F0,T0 | 59.41 | phase angle of cal 0
input of F100,T100 | 31.06 | phase angle of cal 100 (water-vapor saturated air)
input of Fm,Tm | 30.4 | measured phase angle
input of T0 | 18 | temperature of cal 0
input of T100 | 17.8 | temperatur of cal100
input of Tm | 18.2 | temperature at measurement
air pressure patm | 1013 | air pressure
Salinity | 29.8 | Salinity
Results of calculation + compensation
[%] air saturation | 105.0999 | %
[%] O2 | 21.9659 | %
pO2 [hPa] | 218.4326 | hPa
pO2 [Torr] | 163.8384 | Torr
cO2 [mg/L] | 8.2965 | mg/L
cO2 [ppm] | 8.2965 | ppm
cO2 [µmol/L] | 259.2647 | µmol/L
Internal calculated parameter (DO NOT CHANGE !!)
f1 | 0.801
DF/K | -0.08
DKSV/K | 0.000383
m | 22.9
tan(F0,T100) | 1.6927
tan(F0,Tm) | 1.6905
tan(F100,T100) | 0.60229
tan(Fm,Tm) | 0.5867
KSV,T100 | 0.033895
KSV,Tm | 0.034048
A | 155.3811
B | 13.7383
C | -0.64418
a | 1.7569e-05
b | 0.0043661
c | -0.65295
Chlorinity | 16.4931

## Recalculate Fibox values
file name: PSt1_Eq4_Sal |  |  |  | 2008-03-29
author: Dr. Christian Huber
programm for temperature compensated oxygen calculation of membrane PSt1
Fit of the calibration curve is optimized between 0 and 250 % air-saturation (0 and 388 Torr)
Insert your calibration values
cal0 | 59.41 | ° | T0 | 18 | °C
cal100 | 31.06 | ° | T100 | 17.8 | °C
air pressure | 1020 | mbar
Salinity | 29.8 | °/°°
Internal calculated parameter (DO NOT CHANGE !!)
f1 | 0.801 | tan(F0,T100) | 1.6927 | A | 155.3811 | KSV,T100 | 0.033895 | a | 1.759e-05
DF/K | -0.08 | tan(F0,Tm) | 1.6905 | B | 13.7383 | KSV,Tm | 0.034048 | b | 0.0043809 | [5 empty] | INSERT YOUR RESPECTIVE DATA
DKSV/K | 0.000383 | tan(F100,T100) | 0.60229 | C | -0.64418 |  |  | c | -0.65253 |  |  |  |  | Blank (chamber 3) cO2 [µmol/ h] | chamber volume [L] | Fresh weight [g]
m | 22.9 | tan(Fm,Tm) | 0.5874 |  |  |  |  | Chlorinity | 16.4931 |  |  |  |  | -0.41107 | 0.02607 | 0.055
Insert the data you get from the Ascii File recorded with the OxyView TX3-V5.31.exe software | [7 empty] | The oxygen contents are calculated automatically from the raw data
date | time/hh:mm:ss | logtime/min | oxygen/% airsatur. | phase/° | amp | temp/°C |  | % air-sat. | % oxygen | pO2 (hPa) | pO2 (Torr) | cO2 [mg/L] | cO2 [µmol/L] |  | cO2 [µmol/L 10 min] | cO2 [µmol/L h] | cO2 [µmol/ h] | cO2 [µmol/ h gFW]
2010-07-28 | 14:39:07 | 0 | 308.574 | 30.43 | 5906 | 18.2 |  | 104.8268 | 21.9088 | 219.4023 | 164.5657 | 8.3333 | 260.4155 |  | 37.374 | 224.244 | 6.257 | 113.7644
2010-07-28 | 14:39:18 | 0.185 | 311.533 | 30.32 | 5906 | 18.2 |  | 105.8321 | 22.1189 | 221.5064 | 166.1439 | 8.4132 | 262.9128
2010-07-28 | 14:39:28 | 0.352 | 313.161 | 30.26 | 5906 | 18.2 |  | 106.385 | 22.2345 | 222.6638 | 167.012 | 8.4572 | 264.2865 |  | change data input according to column N
2010-07-28 | 14:39:38 | 0.519 | 312.618 | 30.28 | 5905 | 18.2 |  | 106.2003 | 22.1959 | 222.2772 | 166.7221 | 8.4425 | 263.8277
2010-07-28 | 14:39:48 | 0.686 | 311.804 | 30.31 | 5898 | 18.2 |  | 105.924 | 22.1381 | 221.6988 | 166.2882 | 8.4205 | 263.1412
2010-07-28 | 14:39:58 | 0.853 | 309.915 | 30.38 | 5907 | 18.2 |  | 105.2824 | 22.004 | 220.3559 | 165.281 | 8.3695 | 261.5473
2010-07-28 | 14:40:08 | 1.02 | 311.533 | 30.32 | 5916 | 18.2 |  | 105.8321 | 22.1189 | 221.5064 | 166.1439 | 8.4132 | 262.9128
2010-07-28 | 14:40:18 | 1.187 | 310.454 | 30.36 | 5904 | 18.2 |  | 105.4652 | 22.0422 | 220.7386 | 165.5681 | 8.3841 | 262.0016
2010-07-28 | 14:40:28 | 1.354 | 313.161 | 30.26 | 5914 | 18.2 |  | 106.385 | 22.2345 | 222.6638 | 167.012 | 8.4572 | 264.2865
2010-07-28 | 14:40:38 | 1.521 | 309.647 | 30.39 | 5906 | 18.2 |  | 105.1911 | 21.9849 | 220.1648 | 165.1376 | 8.3623 | 261.3205
2010-07-28 | 14:40:48 | 1.687 | 309.378 | 30.4 | 5919 | 18.2 |  | 105.0999 | 21.9659 | 219.9739 | 164.9945 | 8.355 | 261.0939
2010-07-28 | 14:40:58 | 1.855 | 308.307 | 30.44 | 5915 | 18.2 |  | 104.7359 | 21.8898 | 219.2122 | 164.4231 | 8.3261 | 260.1898
2010-07-28 | 14:41:08 | 2.021 | 309.915 | 30.38 | 5912 | 18.2 |  | 105.2824 | 22.004 | 220.3559 | 165.281 | 8.3695 | 261.5473
2010-07-28 | 14:41:18 | 2.188 | 307.24 | 30.48 | 5909 | 18.2 |  | 104.3734 | 21.814 | 218.4535 | 163.854 | 8.2973 | 259.2892
2010-07-28 | 14:41:28 | 2.355 | 308.574 | 30.43 | 5921 | 18.2 |  | 104.8268 | 21.9088 | 219.4023 | 164.5657 | 8.3333 | 260.4155
2010-07-28 | 14:41:38 | 2.522 | 306.442 | 30.51 | 5920 | 18.2 |  | 104.1025 | 21.7574 | 217.8864 | 163.4287 | 8.2757 | 258.6162
2010-07-28 | 14:41:48 | 2.689 | 309.647 | 30.39 | 5920 | 18.2 |  | 105.1911 | 21.9849 | 220.1648 | 165.1376 | 8.3623 | 261.3205
2010-07-28 | 14:41:58 | 2.856 | 307.773 | 30.46 | 5919 | 18.2 |  | 104.5545 | 21.8519 | 218.8325 | 164.1383 | 8.3117 | 259.7391
2010-07-28 | 14:42:08 | 3.023 | 309.647 | 30.39 | 5913 | 18.2 |  | 105.1911 | 21.9849 | 220.1648 | 165.1376 | 8.3623 | 261.3205
2010-07-28 | 14:42:18 | 3.19 | 309.915 | 30.38 | 5929 | 18.2 |  | 105.2824 | 22.004 | 220.3559 | 165.281 | 8.3695 | 261.5473
2010-07-28 | 14:42:28 | 3.357 | 309.11 | 30.41 | 5917 | 18.2 |  | 105.0087 | 21.9468 | 219.7832 | 164.8514 | 8.3478 | 260.8675
2010-07-28 | 14:42:38 | 3.524 | 308.842 | 30.42 | 5927 | 18.2 |  | 104.9177 | 21.9278 | 219.5927 | 164.7085 | 8.3405 | 260.6414
2010-07-28 | 14:42:48 | 3.69 | 308.842 | 30.42 | 5934 | 18.2 |  | 104.9177 | 21.9278 | 219.5927 | 164.7085 | 8.3405 | 260.6414 |  |  | regression formula | time
2010-07-28 | 14:42:58 | 3.857 | 309.11 | 30.41 | 5929 | 18.2 |  | 105.0087 | 21.9468 | 219.7832 | 164.8514 | 8.3478 | 260.8675 |  | value for T=26 min. | 305.972 | T11
2010-07-28 | 14:43:08 | 4.024 | 309.647 | 30.39 | 5932 | 18.2 |  | 105.1911 | 21.9849 | 220.1648 | 165.1376 | 8.3623 | 261.3205 |  | value for T=1 min. | 268.598 | T1
2010-07-28 | 14:43:18 | 4.191 | 305.647 | 30.54 | 5922 | 18.2 |  | 103.8324 | 21.701 | 217.321 | 163.0046 | 8.2542 | 257.9451 |  | difference between T25 and T1 | 37.374 | 10minutes
2010-07-28 | 14:43:28 | 4.358 | 307.773 | 30.46 | 5937 | 18.2 |  | 104.5545 | 21.8519 | 218.8325 | 164.1383 | 8.3117 | 259.7391 |  | calculate from regression curve values for 10 minutes of photosynthesis or respiration
2010-07-28 | 14:43:38 | 4.525 | 307.24 | 30.48 | 5930 | 18.2 |  | 104.3734 | 21.814 | 218.4535 | 163.854 | 8.2973 | 259.2892
2010-07-28 | 14:43:48 | 4.692 | 306.708 | 30.5 | 5935 | 18.2 |  | 104.1927 | 21.7763 | 218.0753 | 163.5703 | 8.2829 | 258.8403
2010-07-28 | 14:43:58 | 4.859 | 304.59 | 30.58 | 5949 | 18.2 |  | 103.4734 | 21.6259 | 216.5697 | 162.4411 | 8.2257 | 257.0534
2010-07-28 | 14:44:08 | 5.025 | 305.912 | 30.53 | 5923 | 18.2 |  | 103.9223 | 21.7198 | 217.5093 | 163.1458 | 8.2614 | 258.1686
2010-07-28 | 14:44:18 | 5.192 | 308.457 | 30.39 | 5941 | 18.3 |  | 105.0003 | 21.9451 | 219.7365 | 164.8164 | 8.3313 | 260.3528
2010-07-28 | 14:44:28 | 5.359 | 306.059 | 30.48 | 5936 | 18.3 |  | 104.1841 | 21.7745 | 218.0284 | 163.5352 | 8.2665 | 258.329
2010-07-28 | 14:44:38 | 5.526 | 305.794 | 30.49 | 5933 | 18.3 |  | 104.0938 | 21.7556 | 217.8395 | 163.3935 | 8.2594 | 258.1052
2010-07-28 | 14:44:48 | 5.693 | 306.059 | 30.48 | 5933 | 18.3 |  | 104.1841 | 21.7745 | 218.0284 | 163.5352 | 8.2665 | 258.329
2010-07-28 | 14:44:58 | 5.86 | 304.472 | 30.54 | 5939 | 18.3 |  | 103.644 | 21.6616 | 216.8981 | 162.6874 | 8.2237 | 256.9897
2010-07-28 | 14:45:08 | 6.027 | 307.655 | 30.42 | 5939 | 18.3 |  | 104.7274 | 21.888 | 219.1654 | 164.3881 | 8.3096 | 259.6762
2010-07-28 | 14:45:18 | 6.194 | 305.794 | 30.49 | 5940 | 18.3 |  | 104.0938 | 21.7556 | 217.8395 | 163.3935 | 8.2594 | 258.1052
2010-07-28 | 14:45:28 | 6.361 | 305.529 | 30.5 | 5946 | 18.3 |  | 104.0037 | 21.7368 | 217.6509 | 163.252 | 8.2522 | 257.8817
2010-07-28 | 14:45:38 | 6.528 | 305.529 | 30.5 | 5945 | 18.3 |  | 104.0037 | 21.7368 | 217.6509 | 163.252 | 8.2522 | 257.8817
2010-07-28 | 14:45:48 | 6.695 | 305.529 | 30.5 | 5946 | 18.3 |  | 104.0037 | 21.7368 | 217.6509 | 163.252 | 8.2522 | 257.8817
2010-07-28 | 14:45:58 | 6.862 | 305.529 | 30.5 | 5945 | 18.3 |  | 104.0037 | 21.7368 | 217.6509 | 163.252 | 8.2522 | 257.8817
2010-07-28 | 14:46:08 | 7.028 | 305.529 | 30.5 | 5946 | 18.3 |  | 104.0037 | 21.7368 | 217.6509 | 163.252 | 8.2522 | 257.8817
2010-07-28 | 14:46:18 | 7.195 | 303.945 | 30.56 | 5953 | 18.3 |  | 103.4646 | 21.6241 | 216.5228 | 162.4059 | 8.2094 | 256.545
2010-07-28 | 14:46:28 | 7.362 | 305.264 | 30.51 | 5952 | 18.3 |  | 103.9136 | 21.7179 | 217.4624 | 163.1107 | 8.2451 | 257.6583
2010-07-28 | 14:46:38 | 7.529 | 304.472 | 30.54 | 5952 | 18.3 |  | 103.644 | 21.6616 | 216.8981 | 162.6874 | 8.2237 | 256.9897
2010-07-28 | 14:46:47 | 7.679 | 304.472 | 30.54 | 5950 | 18.3 |  | 103.644 | 21.6616 | 216.8981 | 162.6874 | 8.2237 | 256.9897
2010-07-28 | 14:46:57 | 7.846 | 305 | 30.52 | 5956 | 18.3 |  | 103.8236 | 21.6991 | 217.2741 | 162.9694 | 8.2379 | 257.4352
2010-07-28 | 14:47:07 | 8.013 | 300.925 | 30.72 | 5953 | 18.2 |  | 102.2281 | 21.3657 | 213.9634 | 160.4862 | 8.1267 | 253.9598
2010-07-28 | 14:47:17 | 8.18 | 309.378 | 30.4 | 5949 | 18.2 |  | 105.0999 | 21.9659 | 219.9739 | 164.9945 | 8.355 | 261.0939
2010-07-28 | 14:47:27 | 8.347 | 306.442 | 30.51 | 5964 | 18.2 |  | 104.1025 | 21.7574 | 217.8864 | 163.4287 | 8.2757 | 258.6162
2010-07-28 | 14:47:37 | 8.514 | 306.177 | 30.52 | 5964 | 18.2 |  | 104.0124 | 21.7386 | 217.6978 | 163.2872 | 8.2686 | 258.3923
2010-07-28 | 14:47:47 | 8.681 | 308.307 | 30.44 | 5954 | 18.2 |  | 104.7359 | 21.8898 | 219.2122 | 164.4231 | 8.3261 | 260.1898
2010-07-28 | 14:47:57 | 8.848 | 305.382 | 30.55 | 5964 | 18.2 |  | 103.7425 | 21.6822 | 217.1329 | 162.8635 | 8.2471 | 257.7218
2010-07-28 | 14:48:07 | 9.015 | 305.118 | 30.56 | 5954 | 18.2 |  | 103.6527 | 21.6634 | 216.945 | 162.7226 | 8.24 | 257.4988
2010-07-28 | 14:48:17 | 9.182 | 304.327 | 30.59 | 5974 | 18.2 |  | 103.3839 | 21.6072 | 216.3824 | 162.3006 | 8.2186 | 256.831
2010-07-28 | 14:48:27 | 9.348 | 304.854 | 30.57 | 5959 | 18.2 |  | 103.563 | 21.6447 | 216.7573 | 162.5818 | 8.2328 | 257.276
2010-07-28 | 14:48:37 | 9.515 | 305.912 | 30.53 | 5965 | 18.2 |  | 103.9223 | 21.7198 | 217.5093 | 163.1458 | 8.2614 | 258.1686
2010-07-28 | 14:48:47 | 9.682 | 304.184 | 30.64 | 5972 | 18.1 |  | 103.1249 | 21.5531 | 215.8687 | 161.9153 | 8.2136 | 256.674
2010-07-28 | 14:48:57 | 9.849 | 304.975 | 30.61 | 5965 | 18.1 |  | 103.3929 | 21.6091 | 216.4296 | 162.336 | 8.2349 | 257.341
2010-07-28 | 14:49:07 | 10.016 | 304.975 | 30.61 | 5981 | 18.1 |  | 103.3929 | 21.6091 | 216.4296 | 162.336 | 8.2349 | 257.341
2010-07-28 | 14:49:18 | 10.183 | 306.563 | 30.55 | 5982 | 18.1 |  | 103.9312 | 21.7216 | 217.5565 | 163.1812 | 8.2778 | 258.6808
2010-07-28 | 14:49:28 | 10.35 | 306.032 | 30.57 | 5976 | 18.1 |  | 103.7514 | 21.684 | 217.1801 | 162.8989 | 8.2635 | 258.2333
2010-07-28 | 14:49:38 | 10.517 | 302.61 | 30.7 | 5979 | 18.1 |  | 102.5913 | 21.4416 | 214.7518 | 161.0775 | 8.1711 | 255.346
2010-07-28 | 14:49:48 | 10.684 | 303.013 | 30.64 | 5979 | 18.2 |  | 102.9376 | 21.514 | 215.4483 | 161.6 | 8.1831 | 255.7223
2010-07-28 | 14:49:58 | 10.851 | 304.59 | 30.58 | 5982 | 18.2 |  | 103.4734 | 21.6259 | 216.5697 | 162.4411 | 8.2257 | 257.0534
2010-07-28 | 14:50:08 | 11.018 | 303.538 | 30.62 | 5982 | 18.2 |  | 103.1159 | 21.5512 | 215.8214 | 161.8798 | 8.1973 | 256.1651
2010-07-28 | 14:50:18 | 11.185 | 303.801 | 30.61 | 5986 | 18.2 |  | 103.2051 | 21.5699 | 216.0082 | 162.0199 | 8.2044 | 256.3869
2010-07-28 | 14:50:28 | 11.351 | 300.665 | 30.73 | 5987 | 18.2 |  | 102.1398 | 21.3472 | 213.7786 | 160.3476 | 8.1197 | 253.7404
2010-07-28 | 14:50:38 | 11.518 | 305.118 | 30.56 | 5985 | 18.2 |  | 103.6527 | 21.6634 | 216.945 | 162.7226 | 8.24 | 257.4988
2010-07-28 | 14:50:48 | 11.685 | 302.751 | 30.65 | 5984 | 18.2 |  | 102.8486 | 21.4954 | 215.2621 | 161.4603 | 8.176 | 255.5013
2010-07-28 | 14:50:58 | 11.852 | 301.706 | 30.69 | 5987 | 18.2 |  | 102.4935 | 21.4211 | 214.5189 | 160.9028 | 8.1478 | 254.6191
2010-07-28 | 14:51:08 | 12.019 | 303.013 | 30.64 | 5990 | 18.2 |  | 102.9376 | 21.514 | 215.4483 | 161.6 | 8.1831 | 255.7223
2010-07-28 | 14:51:18 | 12.186 | 300.925 | 30.72 | 5990 | 18.2 |  | 102.2281 | 21.3657 | 213.9634 | 160.4862 | 8.1267 | 253.9598
2010-07-28 | 14:51:28 | 12.353 | 301.185 | 30.71 | 5996 | 18.2 |  | 102.3165 | 21.3842 | 214.1484 | 160.6249 | 8.1337 | 254.1793
2010-07-28 | 14:51:38 | 12.52 | 300.146 | 30.75 | 5997 | 18.2 |  | 101.9635 | 21.3104 | 213.4095 | 160.0707 | 8.1057 | 253.3024
2010-07-28 | 14:51:48 | 12.687 | 304.064 | 30.6 | 5990 | 18.2 |  | 103.2945 | 21.5885 | 216.1952 | 162.1602 | 8.2115 | 256.6088
2010-07-28 | 14:51:58 | 12.853 | 302.228 | 30.67 | 5991 | 18.2 |  | 102.6709 | 21.4582 | 214.8901 | 161.1813 | 8.1619 | 255.0598
2010-07-28 | 14:52:08 | 13.02 | 304.064 | 30.6 | 5996 | 18.2 |  | 103.2945 | 21.5885 | 216.1952 | 162.1602 | 8.2115 | 256.6088
2010-07-28 | 14:52:18 | 13.187 | 301.706 | 30.69 | 5996 | 18.2 |  | 102.4935 | 21.4211 | 214.5189 | 160.9028 | 8.1478 | 254.6191
2010-07-28 | 14:52:28 | 13.354 | 302.633 | 30.61 | 5989 | 18.3 |  | 103.0178 | 21.5307 | 215.5877 | 161.7045 | 8.174 | 255.4372
2010-07-28 | 14:52:38 | 13.521 | 300.806 | 30.68 | 6002 | 18.3 |  | 102.396 | 21.4008 | 214.2864 | 160.7285 | 8.1246 | 253.8953
2010-07-28 | 14:52:48 | 13.688 | 302.109 | 30.63 | 6002 | 18.3 |  | 102.8397 | 21.4935 | 215.215 | 161.425 | 8.1599 | 254.9955
2010-07-28 | 14:52:58 | 13.855 | 300.546 | 30.69 | 5996 | 18.3 |  | 102.3075 | 21.3823 | 214.1012 | 160.5896 | 8.1176 | 253.6759
2010-07-28 | 14:53:08 | 14.022 | 303.157 | 30.59 | 5996 | 18.3 |  | 103.1963 | 21.568 | 215.9612 | 161.9847 | 8.1881 | 255.8796
2010-07-28 | 14:53:18 | 14.189 | 297.96 | 30.79 | 6007 | 18.3 |  | 101.4274 | 21.1983 | 212.2593 | 159.208 | 8.0478 | 251.4935
2010-07-28 | 14:53:28 | 14.356 | 301.587 | 30.65 | 6012 | 18.3 |  | 102.662 | 21.4563 | 214.843 | 161.146 | 8.1458 | 254.5548
2010-07-28 | 14:53:38 | 14.523 | 298.733 | 30.76 | 5998 | 18.3 |  | 101.6905 | 21.2533 | 212.81 | 159.6211 | 8.0687 | 252.146
2010-07-28 | 14:53:48 | 14.689 | 300.546 | 30.69 | 6018 | 18.3 |  | 102.3075 | 21.3823 | 214.1012 | 160.5896 | 8.1176 | 253.6759
2010-07-28 | 14:53:58 | 14.856 | 300.027 | 30.71 | 5999 | 18.3 |  | 102.1308 | 21.3453 | 213.7314 | 160.3122 | 8.1036 | 253.2377
2010-07-28 | 14:54:08 | 15.023 | 299.768 | 30.72 | 6002 | 18.3 |  | 102.0425 | 21.3269 | 213.5468 | 160.1737 | 8.0966 | 253.0189
2010-07-28 | 14:54:18 | 15.19 | 296.933 | 30.83 | 6003 | 18.3 |  | 101.0777 | 21.1252 | 211.5276 | 158.6592 | 8.02 | 250.6265
2010-07-28 | 14:54:28 | 15.357 | 301.066 | 30.67 | 6008 | 18.3 |  | 102.4846 | 21.4193 | 214.4717 | 160.8675 | 8.1317 | 254.1149
2010-07-28 | 14:54:38 | 15.524 | 299.111 | 30.79 | 6010 | 18.2 |  | 101.6119 | 21.2369 | 212.6735 | 159.5187 | 8.0777 | 252.4288
2010-07-28 | 14:54:48 | 15.691 | 301.706 | 30.69 | 6007 | 18.2 |  | 102.4935 | 21.4211 | 214.5189 | 160.9028 | 8.1478 | 254.6191
2010-07-28 | 14:54:58 | 15.858 | 301.445 | 30.7 | 6000 | 18.2 |  | 102.405 | 21.4026 | 214.3335 | 160.7638 | 8.1408 | 254.3991
2010-07-28 | 14:55:08 | 16.025 | 298.595 | 30.81 | 6007 | 18.2 |  | 101.4366 | 21.2002 | 212.3066 | 159.2435 | 8.0638 | 251.9933
2010-07-28 | 14:55:18 | 16.192 | 297.566 | 30.85 | 6006 | 18.2 |  | 101.087 | 21.1272 | 211.5749 | 158.6947 | 8.036 | 251.1249
2010-07-28 | 14:55:28 | 16.359 | 300.146 | 30.75 | 6019 | 18.2 |  | 101.9635 | 21.3104 | 213.4095 | 160.0707 | 8.1057 | 253.3024
2010-07-28 | 14:55:38 | 16.525 | 299.887 | 30.76 | 6007 | 18.2 |  | 101.8755 | 21.292 | 213.2252 | 159.9325 | 8.0987 | 253.0837
2010-07-28 | 14:55:48 | 16.692 | 298.852 | 30.8 | 6020 | 18.2 |  | 101.5242 | 21.2186 | 212.49 | 159.381 | 8.0708 | 252.211
2010-07-28 | 14:55:58 | 16.859 | 299.111 | 30.79 | 6012 | 18.2 |  | 101.6119 | 21.2369 | 212.6735 | 159.5187 | 8.0777 | 252.4288
2010-07-28 | 14:56:08 | 17.026 | 299.628 | 30.77 | 6018 | 18.2 |  | 101.7875 | 21.2736 | 213.0411 | 159.7944 | 8.0917 | 252.8652
2010-07-28 | 14:56:18 | 17.193 | 298.595 | 30.81 | 6020 | 18.2 |  | 101.4366 | 21.2002 | 212.3066 | 159.2435 | 8.0638 | 251.9933
2010-07-28 | 14:56:28 | 17.36 | 298.08 | 30.83 | 6019 | 18.2 |  | 101.2616 | 21.1637 | 211.9404 | 158.9688 | 8.0499 | 251.5587
2010-07-28 | 14:56:38 | 17.527 | 297.822 | 30.84 | 6031 | 18.2 |  | 101.1742 | 21.1454 | 211.7576 | 158.8317 | 8.0429 | 251.3417
2010-07-28 | 14:56:48 | 17.694 | 298.852 | 30.8 | 6013 | 18.2 |  | 101.5242 | 21.2186 | 212.49 | 159.381 | 8.0708 | 252.211
2010-07-28 | 14:56:58 | 17.861 | 298.08 | 30.83 | 6025 | 18.2 |  | 101.2616 | 21.1637 | 211.9404 | 158.9688 | 8.0499 | 251.5587
2010-07-28 | 14:57:08 | 18.028 | 297.309 | 30.86 | 6033 | 18.2 |  | 100.9998 | 21.109 | 211.3924 | 158.5578 | 8.0291 | 250.9083
2010-07-28 | 14:57:18 | 18.195 | 299.628 | 30.77 | 6031 | 18.2 |  | 101.7875 | 21.2736 | 213.0411 | 159.7944 | 8.0917 | 252.8652
2010-07-28 | 14:57:28 | 18.362 | 298.337 | 30.82 | 6020 | 18.2 |  | 101.349 | 21.1819 | 212.1234 | 159.1061 | 8.0568 | 251.7759
2010-07-28 | 14:57:38 | 18.528 | 299.369 | 30.78 | 6031 | 18.2 |  | 101.6996 | 21.2552 | 212.8572 | 159.6565 | 8.0847 | 252.6469
2010-07-28 | 14:57:48 | 18.695 | 298.337 | 30.82 | 6025 | 18.2 |  | 101.349 | 21.1819 | 212.1234 | 159.1061 | 8.0568 | 251.7759
2010-07-28 | 14:57:58 | 18.862 | 296.796 | 30.88 | 6038 | 18.2 |  | 100.8257 | 21.0726 | 211.028 | 158.2845 | 8.0152 | 250.4757
2010-07-28 | 14:58:08 | 19.029 | 300.665 | 30.73 | 6032 | 18.2 |  | 102.1398 | 21.3472 | 213.7786 | 160.3476 | 8.1197 | 253.7404
2010-07-28 | 14:58:17 | 19.179 | 296.796 | 30.88 | 6034 | 18.2 |  | 100.8257 | 21.0726 | 211.028 | 158.2845 | 8.0152 | 250.4757
2010-07-28 | 14:58:27 | 19.346 | 296.796 | 30.88 | 6028 | 18.2 |  | 100.8257 | 21.0726 | 211.028 | 158.2845 | 8.0152 | 250.4757
2010-07-28 | 14:58:37 | 19.513 | 296.54 | 30.89 | 6026 | 18.2 |  | 100.7387 | 21.0544 | 210.8461 | 158.148 | 8.0083 | 250.2598
2010-07-28 | 14:58:47 | 19.68 | 299.887 | 30.76 | 6044 | 18.2 |  | 101.8755 | 21.292 | 213.2252 | 159.9325 | 8.0987 | 253.0837
2010-07-28 | 14:58:57 | 19.847 | 298.595 | 30.81 | 6050 | 18.2 |  | 101.4366 | 21.2002 | 212.3066 | 159.2435 | 8.0638 | 251.9933
2010-07-28 | 14:59:07 | 20.014 | 298.337 | 30.82 | 6038 | 18.2 |  | 101.349 | 21.1819 | 212.1234 | 159.1061 | 8.0568 | 251.7759
2010-07-28 | 14:59:17 | 20.181 | 295.774 | 30.92 | 6038 | 18.2 |  | 100.4785 | 21 | 210.3013 | 157.7394 | 7.9876 | 249.6132
2010-07-28 | 14:59:27 | 20.348 | 295.519 | 30.93 | 6045 | 18.2 |  | 100.3919 | 20.9819 | 210.1201 | 157.6035 | 7.9807 | 249.3981
2010-07-28 | 14:59:37 | 20.515 | 298.08 | 30.83 | 6046 | 18.2 |  | 101.2616 | 21.1637 | 211.9404 | 158.9688 | 8.0499 | 251.5587
2010-07-28 | 14:59:47 | 20.682 | 299.111 | 30.79 | 6055 | 18.2 |  | 101.6119 | 21.2369 | 212.6735 | 159.5187 | 8.0777 | 252.4288
2010-07-28 | 14:59:57 | 20.848 | 295.4 | 30.89 | 6047 | 18.3 |  | 100.5558 | 21.0162 | 210.4353 | 157.8399 | 7.9786 | 249.3324
2010-07-28 | 15:00:07 | 21.015 | 296.166 | 30.86 | 6035 | 18.3 |  | 100.8164 | 21.0706 | 210.9807 | 158.249 | 7.9993 | 249.9785
2010-07-28 | 15:00:17 | 21.182 | 297.446 | 30.81 | 6042 | 18.3 |  | 101.2524 | 21.1617 | 211.8931 | 158.9333 | 8.0339 | 251.0596
2010-07-28 | 15:00:27 | 21.349 | 295.91 | 30.87 | 6050 | 18.3 |  | 100.7294 | 21.0524 | 210.7987 | 158.1125 | 7.9924 | 249.7629
2010-07-28 | 15:00:37 | 21.516 | 297.96 | 30.79 | 6051 | 18.3 |  | 101.4274 | 21.1983 | 212.2593 | 159.208 | 8.0478 | 251.4935
2010-07-28 | 15:00:47 | 21.683 | 298.992 | 30.75 | 6050 | 18.3 |  | 101.7784 | 21.2717 | 212.9939 | 159.759 | 8.0756 | 252.3639
2010-07-28 | 15:00:58 | 21.85 | 294.637 | 30.92 | 6042 | 18.3 |  | 100.2959 | 20.9619 | 209.8916 | 157.4321 | 7.958 | 248.6881
2010-07-28 | 15:01:08 | 22.017 | 294.129 | 30.94 | 6048 | 18.3 |  | 100.1231 | 20.9257 | 209.53 | 157.1608 | 7.9443 | 248.2597
2010-07-28 | 15:01:18 | 22.183 | 293.37 | 30.97 | 6056 | 18.3 |  | 99.8646 | 20.8717 | 208.9889 | 156.755 | 7.9238 | 247.6186
2010-07-28 | 15:01:28 | 22.35 | 295.145 | 30.9 | 6043 | 18.3 |  | 100.4691 | 20.998 | 210.2539 | 157.7038 | 7.9718 | 249.1174
2010-07-28 | 15:01:38 | 22.517 | 295.91 | 30.87 | 6045 | 18.3 |  | 100.7294 | 21.0524 | 210.7987 | 158.1125 | 7.9924 | 249.7629
2010-07-28 | 15:01:48 | 22.684 | 293.37 | 30.97 | 6058 | 18.3 |  | 99.8646 | 20.8717 | 208.9889 | 156.755 | 7.9238 | 247.6186
2010-07-28 | 15:01:58 | 22.851 | 293.622 | 30.96 | 6047 | 18.3 |  | 99.9507 | 20.8897 | 209.1691 | 156.8901 | 7.9306 | 247.8321
2010-07-28 | 15:02:08 | 23.018 | 296.421 | 30.85 | 6055 | 18.3 |  | 100.9034 | 21.0888 | 211.1628 | 158.3856 | 8.0062 | 250.1943
2010-07-28 | 15:02:18 | 23.185 | 293.995 | 30.99 | 6063 | 18.2 |  | 99.8741 | 20.8737 | 209.0363 | 156.7906 | 7.9396 | 248.1117
2010-07-28 | 15:02:28 | 23.352 | 296.029 | 30.91 | 6055 | 18.2 |  | 100.5651 | 21.0181 | 210.4827 | 157.8755 | 7.9945 | 249.8285
2010-07-28 | 15:02:38 | 23.519 | 298.595 | 30.81 | 6056 | 18.2 |  | 101.4366 | 21.2002 | 212.3066 | 159.2435 | 8.0638 | 251.9933
2010-07-28 | 15:02:48 | 23.686 | 295.774 | 30.92 | 6056 | 18.2 |  | 100.4785 | 21 | 210.3013 | 157.7394 | 7.9876 | 249.6132
2010-07-28 | 15:02:58 | 23.853 | 294.502 | 30.97 | 6059 | 18.2 |  | 100.0463 | 20.9097 | 209.3969 | 157.061 | 7.9533 | 248.5397
2010-07-28 | 15:03:08 | 24.02 | 296.285 | 30.9 | 6058 | 18.2 |  | 100.6519 | 21.0362 | 210.6643 | 158.0117 | 8.0014 | 250.044
2010-07-28 | 15:03:18 | 24.186 | 296.54 | 30.89 | 6052 | 18.2 |  | 100.7387 | 21.0544 | 210.8461 | 158.148 | 8.0083 | 250.2598
2010-07-28 | 15:03:28 | 24.353 | 295.01 | 30.95 | 6067 | 18.2 |  | 100.2189 | 20.9458 | 209.7581 | 157.332 | 7.967 | 248.9685
2010-07-28 | 15:03:38 | 24.52 | 293.489 | 31.01 | 6064 | 18.2 |  | 99.7021 | 20.8377 | 208.6765 | 156.5207 | 7.9259 | 247.6846
2010-07-28 | 15:03:48 | 24.687 | 294.756 | 30.96 | 6070 | 18.2 |  | 100.1326 | 20.9277 | 209.5774 | 157.1964 | 7.9601 | 248.754
2010-07-28 | 15:03:58 | 24.854 | 294.756 | 30.96 | 6067 | 18.2 |  | 100.1326 | 20.9277 | 209.5774 | 157.1964 | 7.9601 | 248.754
2010-07-28 | 15:04:08 | 25.021 | 294.502 | 30.97 | 6058 | 18.2 |  | 100.0463 | 20.9097 | 209.3969 | 157.061 | 7.9533 | 248.5397
2010-07-28 | 15:04:18 | 25.188 | 294.756 | 30.96 | 6071 | 18.2 |  | 100.1326 | 20.9277 | 209.5774 | 157.1964 | 7.9601 | 248.754
2010-07-28 | 15:04:28 | 25.355 | 296.029 | 30.91 | 6070 | 18.2 |  | 100.5651 | 21.0181 | 210.4827 | 157.8755 | 7.9945 | 249.8285
2010-07-28 | 15:04:38 | 25.522 | 297.053 | 30.87 | 6072 | 18.2 |  | 100.9127 | 21.0908 | 211.2101 | 158.4211 | 8.0221 | 250.6919
2010-07-28 | 15:04:48 | 25.689 | 295.265 | 30.94 | 6072 | 18.2 |  | 100.3054 | 20.9638 | 209.939 | 157.4676 | 7.9739 | 249.1832
2010-07-28 | 15:04:58 | 25.856 | 295.01 | 30.95 | 6082 | 18.2 |  | 100.2189 | 20.9458 | 209.7581 | 157.332 | 7.967 | 248.9685
2010-07-28 | 15:05:08 | 26.022 | 292.732 | 31.04 | 6070 | 18.2 |  | 99.4449 | 20.784 | 208.138 | 156.1168 | 7.9055 | 247.0455
2010-07-28 | 15:05:18 | 26.19 | 297.566 | 30.85 | 6077 | 18.2 |  | 101.087 | 21.1272 | 211.5749 | 158.6947 | 8.036 | 251.1249
2010-07-28 | 15:05:28 | 26.356 | 296.029 | 30.91 | 6084 | 18.2 |  | 100.5651 | 21.0181 | 210.4827 | 157.8755 | 7.9945 | 249.8285
2010-07-28 | 15:05:38 | 26.523 | 294.249 | 30.98 | 6082 | 18.2 |  | 99.9602 | 20.8917 | 209.2165 | 156.9257 | 7.9464 | 248.3256
2010-07-28 | 15:05:48 | 26.69 | 291.474 | 31.09 | 6083 | 18.2 |  | 99.0178 | 20.6947 | 207.2441 | 155.4463 | 7.8715 | 245.9844
2010-07-28 | 15:05:58 | 26.857 | 293.236 | 31.02 | 6084 | 18.2 |  | 99.6163 | 20.8198 | 208.4968 | 156.3859 | 7.9191 | 247.4714
2010-07-28 | 15:06:08 | 27.024 | 292.732 | 31.04 | 6086 | 18.2 |  | 99.4449 | 20.784 | 208.138 | 156.1168 | 7.9055 | 247.0455
2010-07-28 | 15:06:18 | 27.191 | 291.474 | 31.09 | 6080 | 18.2 |  | 99.0178 | 20.6947 | 207.2441 | 155.4463 | 7.8715 | 245.9844
2010-07-28 | 15:06:28 | 27.358 | 294.756 | 30.96 | 6086 | 18.2 |  | 100.1326 | 20.9277 | 209.5774 | 157.1964 | 7.9601 | 248.754
2010-07-28 | 15:06:38 | 27.525 | 294.756 | 30.96 | 6081 | 18.2 |  | 100.1326 | 20.9277 | 209.5774 | 157.1964 | 7.9601 | 248.754
2010-07-28 | 15:06:48 | 27.692 | 290.223 | 31.14 | 6089 | 18.2 |  | 98.5927 | 20.6059 | 206.3544 | 154.7789 | 7.8377 | 244.9285
2010-07-28 | 15:06:58 | 27.858 | 295.519 | 30.93 | 6083 | 18.2 |  | 100.3919 | 20.9819 | 210.1201 | 157.6035 | 7.9807 | 249.3981
2010-07-28 | 15:07:08 | 28.025 | 292.48 | 31.05 | 6092 | 18.2 |  | 99.3593 | 20.7661 | 207.9589 | 155.9824 | 7.8987 | 246.8329
2010-07-28 | 15:07:18 | 28.192 | 291.725 | 31.08 | 6089 | 18.2 |  | 99.103 | 20.7125 | 207.4225 | 155.5801 | 7.8783 | 246.1962
2010-07-28 | 15:07:28 | 28.359 | 291.857 | 31.03 | 6093 | 18.3 |  | 99.3497 | 20.7641 | 207.9114 | 155.9468 | 7.8829 | 246.3419
2010-07-28 | 15:07:38 | 28.526 | 291.606 | 31.04 | 6094 | 18.3 |  | 99.2642 | 20.7462 | 207.7324 | 155.8125 | 7.8762 | 246.1298
2010-07-28 | 15:07:48 | 28.693 | 291.104 | 31.06 | 6091 | 18.3 |  | 99.0934 | 20.7105 | 207.375 | 155.5444 | 7.8626 | 245.7063
2010-07-28 | 15:07:58 | 28.86 | 292.864 | 30.99 | 6086 | 18.3 |  | 99.6926 | 20.8358 | 208.629 | 156.4851 | 7.9101 | 247.1922
2010-07-28 | 15:08:08 | 29.027 | 291.355 | 31.05 | 6088 | 18.3 |  | 99.1787 | 20.7284 | 207.5536 | 155.6784 | 7.8694 | 245.918
2010-07-28 | 15:08:18 | 29.194 | 290.853 | 31.07 | 6084 | 18.3 |  | 99.0081 | 20.6927 | 207.1965 | 155.4106 | 7.8558 | 245.4949
2010-07-28 | 15:08:28 | 29.361 | 291.857 | 31.03 | 6093 | 18.3 |  | 99.3497 | 20.7641 | 207.9114 | 155.9468 | 7.8829 | 246.3419
2010-07-28 | 15:08:38 | 29.527 | 290.103 | 31.1 | 6102 | 18.3 |  | 98.7528 | 20.6393 | 206.6622 | 155.0098 | 7.8356 | 244.8618
2010-07-28 | 15:08:48 | 29.694 | 292.108 | 31.02 | 6091 | 18.3 |  | 99.4353 | 20.782 | 208.0905 | 156.0812 | 7.8897 | 246.5542
2010-07-28 | 15:08:58 | 29.861 | 289.605 | 31.12 | 6096 | 18.3 |  | 98.583 | 20.6038 | 206.3068 | 154.7433 | 7.8221 | 244.4407
2010-07-28 | 15:09:08 | 30.028 | 291.857 | 31.03 | 6092 | 18.3 |  | 99.3497 | 20.7641 | 207.9114 | 155.9468 | 7.8829 | 246.3419
2010-07-28 | 15:09:18 | 30.195 | 290.353 | 31.09 | 6097 | 18.3 |  | 98.8378 | 20.6571 | 206.8401 | 155.1433 | 7.8423 | 245.0726
2010-07-28 | 15:09:28 | 30.362 | 291.355 | 31.05 | 6099 | 18.3 |  | 99.1787 | 20.7284 | 207.5536 | 155.6784 | 7.8694 | 245.918
2010-07-28 | 15:09:38 | 30.529 | 291.725 | 31.08 | 6088 | 18.2 |  | 99.103 | 20.7125 | 207.4225 | 155.5801 | 7.8783 | 246.1962
2010-07-28 | 15:09:48 | 30.696 | 294.502 | 30.97 | 6098 | 18.2 |  | 100.0463 | 20.9097 | 209.3969 | 157.061 | 7.9533 | 248.5397
2010-07-28 | 15:09:57 | 30.846 | 292.228 | 31.06 | 6102 | 18.2 |  | 99.2738 | 20.7482 | 207.7799 | 155.8482 | 7.8919 | 246.6205
2010-07-28 | 15:10:07 | 31.013 | 290.473 | 31.13 | 6113 | 18.2 |  | 98.6775 | 20.6236 | 206.532 | 154.9122 | 7.8445 | 245.1392
2010-07-28 | 15:10:17 | 31.18 | 291.474 | 31.09 | 6106 | 18.2 |  | 99.0178 | 20.6947 | 207.2441 | 155.4463 | 7.8715 | 245.9844
2010-07-28 | 15:10:27 | 31.347 | 291.977 | 31.07 | 6094 | 18.2 |  | 99.1884 | 20.7304 | 207.6011 | 155.7141 | 7.8851 | 246.4082
2010-07-28 | 15:10:37 | 31.513 | 291.977 | 31.07 | 6107 | 18.2 |  | 99.1884 | 20.7304 | 207.6011 | 155.7141 | 7.8851 | 246.4082
2010-07-28 | 15:10:47 | 31.681 | 289.974 | 31.15 | 6106 | 18.2 |  | 98.5079 | 20.5882 | 206.177 | 154.6459 | 7.831 | 244.7179
2010-07-28 | 15:10:57 | 31.847 | 291.224 | 31.1 | 6098 | 18.2 |  | 98.9326 | 20.6769 | 207.0658 | 155.3125 | 7.8647 | 245.7728
2010-07-28 | 15:11:07 | 32.014 | 289.475 | 31.17 | 6116 | 18.2 |  | 98.3386 | 20.5528 | 205.8226 | 154.3801 | 7.8175 | 244.2973
2010-07-28 | 15:11:17 | 32.181 | 290.473 | 31.13 | 6107 | 18.2 |  | 98.6775 | 20.6236 | 206.532 | 154.9122 | 7.8445 | 245.1392
2010-07-28 | 15:11:27 | 32.348 | 289.974 | 31.15 | 6108 | 18.2 |  | 98.5079 | 20.5882 | 206.177 | 154.6459 | 7.831 | 244.7179
2010-07-28 | 15:11:37 | 32.515 | 292.228 | 31.06 | 6104 | 18.2 |  | 99.2738 | 20.7482 | 207.7799 | 155.8482 | 7.8919 | 246.6205
2010-07-28 | 15:11:47 | 32.682 | 289.974 | 31.15 | 6117 | 18.2 |  | 98.5079 | 20.5882 | 206.177 | 154.6459 | 7.831 | 244.7179
2010-07-28 | 15:11:57 | 32.849 | 290.473 | 31.13 | 6118 | 18.2 |  | 98.6775 | 20.6236 | 206.532 | 154.9122 | 7.8445 | 245.1392
2010-07-28 | 15:12:07 | 33.016 | 291.474 | 31.09 | 6118 | 18.2 |  | 99.0178 | 20.6947 | 207.2441 | 155.4463 | 7.8715 | 245.9844
2010-07-28 | 15:12:17 | 33.183 | 291.224 | 31.1 | 6114 | 18.2 |  | 98.9326 | 20.6769 | 207.0658 | 155.3125 | 7.8647 | 245.7728
2010-07-28 | 15:12:28 | 33.35 | 289.475 | 31.17 | 6121 | 18.2 |  | 98.3386 | 20.5528 | 205.8226 | 154.3801 | 7.8175 | 244.2973
2010-07-28 | 15:12:38 | 33.517 | 290.723 | 31.12 | 6109 | 18.2 |  | 98.7625 | 20.6414 | 206.7097 | 155.0455 | 7.8512 | 245.3502
2010-07-28 | 15:12:48 | 33.683 | 289.475 | 31.17 | 6132 | 18.2 |  | 98.3386 | 20.5528 | 205.8226 | 154.3801 | 7.8175 | 244.2973
2010-07-28 | 15:12:58 | 33.85 | 290.973 | 31.11 | 6121 | 18.2 |  | 98.8475 | 20.6591 | 206.8877 | 155.1789 | 7.858 | 245.5614
2010-07-28 | 15:13:08 | 34.017 | 287.739 | 31.24 | 6110 | 18.2 |  | 97.7487 | 20.4295 | 204.5879 | 153.454 | 7.7706 | 242.8317
2010-07-28 | 15:13:18 | 34.184 | 289.227 | 31.18 | 6125 | 18.2 |  | 98.2541 | 20.5351 | 205.6457 | 154.2474 | 7.8108 | 244.0873
2010-07-28 | 15:13:28 | 34.351 | 291.474 | 31.09 | 6121 | 18.2 |  | 99.0178 | 20.6947 | 207.2441 | 155.4463 | 7.8715 | 245.9844
2010-07-28 | 15:13:38 | 34.518 | 291.224 | 31.1 | 6116 | 18.2 |  | 98.9326 | 20.6769 | 207.0658 | 155.3125 | 7.8647 | 245.7728
2010-07-28 | 15:13:48 | 34.685 | 290.723 | 31.12 | 6125 | 18.2 |  | 98.7625 | 20.6414 | 206.7097 | 155.0455 | 7.8512 | 245.3502
2010-07-28 | 15:13:58 | 34.852 | 291.474 | 31.09 | 6122 | 18.2 |  | 99.0178 | 20.6947 | 207.2441 | 155.4463 | 7.8715 | 245.9844
2010-07-28 | 15:14:08 | 35.019 | 289.724 | 31.16 | 6119 | 18.2 |  | 98.4232 | 20.5705 | 205.9997 | 154.5129 | 7.8242 | 244.5075
2010-07-28 | 15:14:18 | 35.186 | 289.227 | 31.18 | 6116 | 18.2 |  | 98.2541 | 20.5351 | 205.6457 | 154.2474 | 7.8108 | 244.0873
2010-07-28 | 15:14:28 | 35.353 | 291.725 | 31.08 | 6123 | 18.2 |  | 99.103 | 20.7125 | 207.4225 | 155.5801 | 7.8783 | 246.1962
2010-07-28 | 15:14:38 | 35.52 | 292.732 | 31.04 | 6098 | 18.2 |  | 99.4449 | 20.784 | 208.138 | 156.1168 | 7.9055 | 247.0455
2010-07-28 | 15:14:48 | 35.686 | 292.108 | 31.02 | 6100 | 18.3 |  | 99.4353 | 20.782 | 208.0905 | 156.0812 | 7.8897 | 246.5542
2010-07-28 | 15:14:58 | 35.853 | 292.108 | 31.02 | 6110 | 18.3 |  | 99.4353 | 20.782 | 208.0905 | 156.0812 | 7.8897 | 246.5542
2010-07-28 | 15:15:08 | 36.02 | 292.108 | 31.02 | 6099 | 18.3 |  | 99.4353 | 20.782 | 208.0905 | 156.0812 | 7.8897 | 246.5542
2010-07-28 | 15:15:18 | 36.187 | 292.108 | 31.02 | 6096 | 18.3 |  | 99.4353 | 20.782 | 208.0905 | 156.0812 | 7.8897 | 246.5542
2010-07-28 | 15:15:28 | 36.354 | 292.612 | 31 | 6075 | 18.3 |  | 99.6068 | 20.8178 | 208.4493 | 156.3503 | 7.9033 | 246.9793
2010-07-28 | 15:15:38 | 36.521 | 293.117 | 30.98 | 6077 | 18.3 |  | 99.7786 | 20.8537 | 208.8089 | 156.62 | 7.917 | 247.4053
2010-07-28 | 15:15:48 | 36.688 | 295.145 | 30.9 | 6068 | 18.3 |  | 100.4691 | 20.998 | 210.2539 | 157.7038 | 7.9718 | 249.1174
2010-07-28 | 15:15:58 | 36.855 | 294.383 | 30.93 | 6065 | 18.3 |  | 100.2095 | 20.9438 | 209.7107 | 157.2964 | 7.9512 | 248.4738
2010-07-28 | 15:16:08 | 37.022 | 296.166 | 30.86 | 6039 | 18.3 |  | 100.8164 | 21.0706 | 210.9807 | 158.249 | 7.9993 | 249.9785
2010-07-28 | 15:16:18 | 37.189 | 295.145 | 30.9 | 6040 | 18.3 |  | 100.4691 | 20.998 | 210.2539 | 157.7038 | 7.9718 | 249.1174
2010-07-28 | 15:16:28 | 37.355 | 293.622 | 30.96 | 6056 | 18.3 |  | 99.9507 | 20.8897 | 209.1691 | 156.8901 | 7.9306 | 247.8321
2010-07-28 | 15:16:38 | 37.522 | 298.733 | 30.76 | 6037 | 18.3 |  | 101.6905 | 21.2533 | 212.81 | 159.6211 | 8.0687 | 252.146
2010-07-28 | 15:16:48 | 37.689 | 296.166 | 30.86 | 6024 | 18.3 |  | 100.8164 | 21.0706 | 210.9807 | 158.249 | 7.9993 | 249.9785
2010-07-28 | 15:16:58 | 37.856 | 300.286 | 30.7 | 6016 | 18.3 |  | 102.2191 | 21.3638 | 213.9162 | 160.4508 | 8.1106 | 253.4567
2010-07-28 | 15:17:08 | 38.023 | 299.509 | 30.73 | 6009 | 18.3 |  | 101.9544 | 21.3085 | 213.3623 | 160.0353 | 8.0896 | 252.8004
2010-07-28 | 15:17:18 | 38.19 | 299.25 | 30.74 | 5989 | 18.3 |  | 101.8663 | 21.2901 | 213.178 | 159.8971 | 8.0826 | 252.582
2010-07-28 | 15:17:28 | 38.357 | 301.066 | 30.67 | 5994 | 18.3 |  | 102.4846 | 21.4193 | 214.4717 | 160.8675 | 8.1317 | 254.1149
2010-07-28 | 15:17:38 | 38.524 | 303.157 | 30.59 | 5977 | 18.3 |  | 103.1963 | 21.568 | 215.9612 | 161.9847 | 8.1881 | 255.8796
2010-07-28 | 15:17:48 | 38.69 | 304.472 | 30.54 | 5977 | 18.3 |  | 103.644 | 21.6616 | 216.8981 | 162.6874 | 8.2237 | 256.9897
2010-07-28 | 15:17:58 | 38.858 | 306.856 | 30.45 | 5966 | 18.3 |  | 104.4554 | 21.8312 | 218.5961 | 163.961 | 8.2881 | 259.0016
2010-07-28 | 15:18:08 | 39.024 | 304.472 | 30.54 | 5961 | 18.3 |  | 103.644 | 21.6616 | 216.8981 | 162.6874 | 8.2237 | 256.9897
2010-07-28 | 15:18:18 | 39.191 | 303.157 | 30.59 | 5960 | 18.3 |  | 103.1963 | 21.568 | 215.9612 | 161.9847 | 8.1881 | 255.8796
2010-07-28 | 15:18:28 | 39.358 | 304.472 | 30.54 | 5953 | 18.3 |  | 103.644 | 21.6616 | 216.8981 | 162.6874 | 8.2237 | 256.9897
2010-07-28 | 15:18:38 | 39.525 | 308.04 | 30.45 | 5941 | 18.2 |  | 104.6452 | 21.8708 | 219.0222 | 164.2806 | 8.3189 | 259.9643
2010-07-28 | 15:18:48 | 39.692 | 308.307 | 30.44 | 5937 | 18.2 |  | 104.7359 | 21.8898 | 219.2122 | 164.4231 | 8.3261 | 260.1898
2010-07-28 | 15:18:58 | 39.859 | 309.378 | 30.4 | 5925 | 18.2 |  | 105.0999 | 21.9659 | 219.9739 | 164.9945 | 8.355 | 261.0939
2010-07-28 | 15:19:08 | 40.026 | 309.915 | 30.38 | 5926 | 18.2 |  | 105.2824 | 22.004 | 220.3559 | 165.281 | 8.3695 | 261.5473
2010-07-28 | 15:19:18 | 40.193 | 312.346 | 30.29 | 5902 | 18.2 |  | 106.1081 | 22.1766 | 222.0842 | 166.5773 | 8.4352 | 263.5987
2010-07-28 | 15:19:28 | 40.36 | 311.263 | 30.33 | 5903 | 18.2 |  | 105.7402 | 22.0997 | 221.3142 | 165.9997 | 8.4059 | 262.6847
2010-07-28 | 15:19:38 | 40.527 | 312.346 | 30.29 | 5905 | 18.2 |  | 106.1081 | 22.1766 | 222.0842 | 166.5773 | 8.4352 | 263.5987
2010-07-28 | 15:19:48 | 40.694 | 312.889 | 30.27 | 5887 | 18.2 |  | 106.2926 | 22.2152 | 222.4704 | 166.867 | 8.4498 | 264.057
2010-07-28 | 15:19:58 | 40.86 | 312.618 | 30.28 | 5883 | 18.2 |  | 106.2003 | 22.1959 | 222.2772 | 166.7221 | 8.4425 | 263.8277
2010-07-28 | 15:20:08 | 41.027 | 313.979 | 30.23 | 5885 | 18.2 |  | 106.6628 | 22.2925 | 223.245 | 167.448 | 8.4792 | 264.9765
2010-07-28 | 15:20:18 | 41.194 | 314.252 | 30.22 | 5875 | 18.2 |  | 106.7555 | 22.3119 | 223.4392 | 167.5936 | 8.4866 | 265.2069
2010-07-28 | 15:20:28 | 41.361 | 316.997 | 30.12 | 5867 | 18.2 |  | 107.6882 | 22.5068 | 225.3912 | 169.0578 | 8.5608 | 267.5239
2010-07-28 | 15:20:38 | 41.528 | 312.889 | 30.27 | 5863 | 18.2 |  | 106.2926 | 22.2152 | 222.4704 | 166.867 | 8.4498 | 264.057
2010-07-28 | 15:20:48 | 41.695 | 317.55 | 30.1 | 5856 | 18.2 |  | 107.8758 | 22.546 | 225.784 | 169.3524 | 8.5757 | 267.9901
2010-07-28 | 15:20:58 | 41.862 | 318.38 | 30.07 | 5853 | 18.2 |  | 108.158 | 22.605 | 226.3746 | 169.7954 | 8.5981 | 268.6911
2010-07-28 | 15:21:08 | 42.029 | 315.347 | 30.18 | 5837 | 18.2 |  | 107.1275 | 22.3896 | 224.2177 | 168.1775 | 8.5162 | 266.1309
2010-07-28 | 15:21:18 | 42.196 | 317.157 | 30.07 | 5837 | 18.3 |  | 107.962 | 22.5641 | 225.9346 | 169.4653 | 8.5663 | 267.6965
2010-07-28 | 15:21:27 | 42.346 | 319.098 | 30 | 5818 | 18.3 |  | 108.6226 | 22.7021 | 227.3169 | 170.5022 | 8.6187 | 269.3344
2010-07-28 | 15:21:37 | 42.513 | 321.052 | 29.93 | 5817 | 18.3 |  | 109.2878 | 22.8411 | 228.709 | 171.5463 | 8.6715 | 270.9838
2010-07-28 | 15:21:47 | 42.68 | 323.302 | 29.85 | 5814 | 18.3 |  | 110.0537 | 23.0012 | 230.312 | 172.7487 | 8.7323 | 272.883
2010-07-28 | 15:21:57 | 42.847 | 318.542 | 30.02 | 5805 | 18.3 |  | 108.4334 | 22.6626 | 226.921 | 170.2052 | 8.6037 | 268.8653
2010-07-28 | 15:22:07 | 43.014 | 320.492 | 29.95 | 5788 | 18.3 |  | 109.0972 | 22.8013 | 228.3103 | 171.2473 | 8.6564 | 270.5113
2010-07-28 | 15:22:17 | 43.181 | 323.02 | 29.86 | 5792 | 18.3 |  | 109.9577 | 22.9812 | 230.1109 | 172.5978 | 8.7246 | 272.6448
2010-07-28 | 15:22:27 | 43.348 | 324.434 | 29.81 | 5770 | 18.3 |  | 110.439 | 23.0818 | 231.1183 | 173.3535 | 8.7628 | 273.8384
2010-07-28 | 15:22:37 | 43.514 | 326.426 | 29.74 | 5766 | 18.3 |  | 111.1171 | 23.2235 | 232.5372 | 174.4177 | 8.8166 | 275.5196
2010-07-28 | 15:22:47 | 43.681 | 326.997 | 29.72 | 5766 | 18.3 |  | 111.3117 | 23.2641 | 232.9444 | 174.7232 | 8.8321 | 276.0021
2010-07-28 | 15:22:57 | 43.848 | 326.997 | 29.72 | 5749 | 18.3 |  | 111.3117 | 23.2641 | 232.9444 | 174.7232 | 8.8321 | 276.0021
2010-07-28 | 15:23:07 | 44.015 | 323.867 | 29.83 | 5752 | 18.3 |  | 110.2462 | 23.0415 | 230.7147 | 173.0508 | 8.7475 | 273.3602
2010-07-28 | 15:23:17 | 44.182 | 331.613 | 29.56 | 5754 | 18.3 |  | 112.8827 | 23.5925 | 236.2322 | 177.1892 | 8.9567 | 279.8976
2010-07-28 | 15:23:27 | 44.349 | 325.855 | 29.76 | 5734 | 18.3 |  | 110.9229 | 23.1829 | 232.1308 | 174.1129 | 8.8012 | 275.038
2010-07-28 | 15:23:37 | 44.516 | 328.32 | 29.63 | 5727 | 18.4 |  | 111.9894 | 23.4058 | 234.3315 | 175.7636 | 8.869 | 277.1574
2010-07-28 | 15:23:47 | 44.683 | 330.341 | 29.56 | 5721 | 18.4 |  | 112.6786 | 23.5498 | 235.7738 | 176.8454 | 8.9236 | 278.8632
2010-07-28 | 15:23:58 | 44.85 | 332.085 | 29.5 | 5716 | 18.4 |  | 113.2733 | 23.6741 | 237.0182 | 177.7787 | 8.9707 | 280.335
2010-07-28 | 15:24:07 | 45.016 | 332.085 | 29.5 | 5708 | 18.4 |  | 113.2733 | 23.6741 | 237.0182 | 177.7787 | 8.9707 | 280.335
2010-07-28 | 15:24:18 | 45.184 | 332.668 | 29.48 | 5698 | 18.4 |  | 113.4723 | 23.7157 | 237.4346 | 178.0911 | 8.9865 | 280.8276
2010-07-28 | 15:24:28 | 45.351 | 331.793 | 29.51 | 5686 | 18.4 |  | 113.1739 | 23.6534 | 236.8102 | 177.6228 | 8.9629 | 280.0891
2010-07-28 | 15:24:38 | 45.518 | 331.502 | 29.52 | 5684 | 18.4 |  | 113.0747 | 23.6326 | 236.6025 | 177.467 | 8.955 | 279.8434
2010-07-28 | 15:24:48 | 45.685 | 332.376 | 29.49 | 5682 | 18.4 |  | 113.3728 | 23.6949 | 237.2263 | 177.9348 | 8.9786 | 280.5812
2010-07-28 | 15:24:58 | 45.852 | 332.668 | 29.48 | 5670 | 18.4 |  | 113.4723 | 23.7157 | 237.4346 | 178.0911 | 8.9865 | 280.8276
2010-07-28 | 15:25:08 | 46.019 | 336.49 | 29.35 | 5657 | 18.4 |  | 114.776 | 23.9882 | 240.1625 | 180.1372 | 9.0897 | 284.054
2010-07-28 | 15:25:18 | 46.186 | 336.194 | 29.36 | 5664 | 18.4 |  | 114.6751 | 23.9671 | 239.9514 | 179.9789 | 9.0817 | 283.8043
2010-07-28 | 15:25:28 | 46.353 | 337.676 | 29.31 | 5644 | 18.4 |  | 115.1807 | 24.0728 | 241.0092 | 180.7723 | 9.1218 | 285.0554
2010-07-28 | 15:25:38 | 46.52 | 338.271 | 29.29 | 5639 | 18.4 |  | 115.3836 | 24.1152 | 241.4338 | 181.0908 | 9.1378 | 285.5577
2010-07-28 | 15:25:48 | 46.686 | 339.274 | 29.3 | 5626 | 18.3 |  | 115.4908 | 24.1376 | 241.6902 | 181.283 | 9.1637 | 286.3644
2010-07-28 | 15:25:58 | 46.853 | 338.976 | 29.31 | 5631 | 18.3 |  | 115.3892 | 24.1163 | 241.4775 | 181.1236 | 9.1556 | 286.1125
2010-07-28 | 15:26:08 | 47.02 | 343.182 | 29.17 | 5616 | 18.3 |  | 116.821 | 24.4156 | 244.474 | 183.3711 | 9.2692 | 289.6627
2010-07-28 | 15:26:18 | 47.187 | 341.974 | 29.21 | 5620 | 18.3 |  | 116.4098 | 24.3296 | 243.6134 | 182.7256 | 9.2366 | 288.6432
2010-07-28 | 15:26:28 | 47.354 | 340.471 | 29.26 | 5605 | 18.3 |  | 115.8982 | 24.2227 | 242.5428 | 181.9225 | 9.196 | 287.3746
2010-07-28 | 15:26:38 | 47.521 | 342.879 | 29.18 | 5601 | 18.3 |  | 116.718 | 24.3941 | 244.2585 | 183.2094 | 9.261 | 289.4074
2010-07-28 | 15:26:48 | 47.688 | 345.308 | 29.1 | 5596 | 18.3 |  | 117.5447 | 24.5668 | 245.9884 | 184.507 | 9.3266 | 291.4571
2010-07-28 | 15:26:58 | 47.855 | 346.224 | 29.07 | 5605 | 18.3 |  | 117.8564 | 24.632 | 246.6408 | 184.9963 | 9.3514 | 292.2301
2010-07-28 | 15:27:08 | 48.022 | 344.699 | 29.12 | 5588 | 18.3 |  | 117.3374 | 24.5235 | 245.5546 | 184.1816 | 9.3102 | 290.9431
2010-07-28 | 15:27:18 | 48.189 | 345.003 | 29.11 | 5576 | 18.3 |  | 117.441 | 24.5452 | 245.7714 | 184.3442 | 9.3184 | 291.2
2010-07-28 | 15:27:28 | 48.356 | 347.449 | 29.03 | 5579 | 18.3 |  | 118.2736 | 24.7192 | 247.5138 | 185.6511 | 9.3845 | 293.2644
2010-07-28 | 15:27:38 | 48.522 | 348.064 | 29.01 | 5580 | 18.3 |  | 118.4828 | 24.7629 | 247.9516 | 185.9795 | 9.4011 | 293.7832
2010-07-28 | 15:27:48 | 48.689 | 348.064 | 29.01 | 5562 | 18.3 |  | 118.4828 | 24.7629 | 247.9516 | 185.9795 | 9.4011 | 293.7832
2010-07-28 | 15:27:58 | 48.856 | 349.297 | 28.97 | 5560 | 18.3 |  | 118.9025 | 24.8506 | 248.83 | 186.6384 | 9.4344 | 294.824
2010-07-28 | 15:28:08 | 49.023 | 353.653 | 28.83 | 5547 | 18.3 |  | 120.3854 | 25.1605 | 251.9332 | 188.966 | 9.552 | 298.5008
2010-07-28 | 15:28:18 | 49.19 | 350.845 | 28.92 | 5537 | 18.3 |  | 119.4297 | 24.9608 | 249.9332 | 187.4658 | 9.4762 | 296.131
2010-07-28 | 15:28:28 | 49.357 | 351.778 | 28.89 | 5538 | 18.3 |  | 119.7472 | 25.0272 | 250.5978 | 187.9643 | 9.5014 | 296.9185
2010-07-28 | 15:28:38 | 49.524 | 350.845 | 28.92 | 5528 | 18.3 |  | 119.4297 | 24.9608 | 249.9332 | 187.4658 | 9.4762 | 296.131
2010-07-28 | 15:28:48 | 49.691 | 351.156 | 28.91 | 5523 | 18.3 |  | 119.5354 | 24.9829 | 250.1545 | 187.6318 | 9.4846 | 296.3932
2010-07-28 | 15:28:58 | 49.858 | 354.28 | 28.81 | 5520 | 18.3 |  | 120.599 | 25.2052 | 252.3802 | 189.3013 | 9.569 | 299.0304
2010-07-28 | 15:29:08 | 50.025 | 354.595 | 28.8 | 5515 | 18.3 |  | 120.706 | 25.2275 | 252.6041 | 189.4692 | 9.5775 | 299.2957
2010-07-28 | 15:29:18 | 50.191 | 356.171 | 28.75 | 5508 | 18.3 |  | 121.2425 | 25.3397 | 253.7268 | 190.3113 | 9.62 | 300.6259
2010-07-28 | 15:29:28 | 50.358 | 353.967 | 28.82 | 5489 | 18.3 |  | 120.4921 | 25.1829 | 252.1566 | 189.1335 | 9.5605 | 298.7655
2010-07-28 | 15:29:38 | 50.525 | 353.55 | 28.79 | 5493 | 18.4 |  | 120.595 | 25.2044 | 252.3385 | 189.2699 | 9.5506 | 298.4552
2010-07-28 | 15:29:48 | 50.692 | 354.177 | 28.77 | 5494 | 18.4 |  | 120.8091 | 25.2491 | 252.7865 | 189.606 | 9.5675 | 298.9851
2010-07-28 | 15:29:58 | 50.859 | 357.018 | 28.68 | 5484 | 18.4 |  | 121.7782 | 25.4516 | 254.8141 | 191.1269 | 9.6443 | 301.3833
2010-07-28 | 15:30:08 | 51.026 | 357.971 | 28.65 | 5468 | 18.4 |  | 122.1032 | 25.5196 | 255.4943 | 191.637 | 9.67 | 302.1878
2010-07-28 | 15:30:18 | 51.193 | 358.29 | 28.64 | 5476 | 18.4 |  | 122.2118 | 25.5423 | 255.7215 | 191.8074 | 9.6786 | 302.4565
2010-07-28 | 15:30:28 | 51.36 | 360.848 | 28.56 | 5458 | 18.4 |  | 123.0845 | 25.7247 | 257.5475 | 193.177 | 9.7477 | 304.6162
2010-07-28 | 15:30:38 | 51.527 | 358.608 | 28.63 | 5454 | 18.4 |  | 122.3205 | 25.565 | 255.9489 | 191.978 | 9.6872 | 302.7254
2010-07-28 | 15:30:48 | 51.694 | 362.135 | 28.52 | 5456 | 18.4 |  | 123.5236 | 25.8164 | 258.4662 | 193.8662 | 9.7825 | 305.7029
2010-07-28 | 15:30:58 | 51.86 | 362.781 | 28.5 | 5438 | 18.4 |  | 123.7438 | 25.8625 | 258.9271 | 194.2118 | 9.7999 | 306.2479
2010-07-28 | 15:31:08 | 52.027 | 361.169 | 28.55 | 5432 | 18.4 |  | 123.1941 | 25.7476 | 257.7768 | 193.349 | 9.7564 | 304.8874
2010-07-28 | 15:31:18 | 52.194 | 362.458 | 28.51 | 5432 | 18.4 |  | 123.6336 | 25.8394 | 258.6965 | 194.0389 | 9.7912 | 305.9753
2010-07-28 | 15:31:28 | 52.361 | 366.03 | 28.4 | 5411 | 18.4 |  | 124.8519 | 26.0941 | 261.2458 | 195.951 | 9.8877 | 308.9904
2010-07-28 | 15:31:38 | 52.528 | 365.377 | 28.42 | 5426 | 18.4 |  | 124.6294 | 26.0475 | 260.7801 | 195.6017 | 9.8701 | 308.4396
2010-07-28 | 15:31:48 | 52.695 | 364.076 | 28.46 | 5413 | 18.4 |  | 124.1857 | 25.9548 | 259.8516 | 194.9053 | 9.8349 | 307.3414
2010-07-28 | 15:31:58 | 52.862 | 364.401 | 28.45 | 5400 | 18.4 |  | 124.2964 | 25.9779 | 260.0834 | 195.0791 | 9.8437 | 307.6155
2010-07-28 | 15:32:08 | 53.029 | 366.356 | 28.39 | 5398 | 18.4 |  | 124.9634 | 26.1173 | 261.479 | 196.1259 | 9.8965 | 309.2662
2010-07-28 | 15:32:18 | 53.196 | 367.339 | 28.36 | 5389 | 18.4 |  | 125.2984 | 26.1874 | 262.1801 | 196.6518 | 9.9231 | 310.0954
2010-07-28 | 15:32:28 | 53.363 | 367.995 | 28.34 | 5384 | 18.4 |  | 125.5224 | 26.2342 | 262.6487 | 197.0033 | 9.9408 | 310.6497
2010-07-28 | 15:32:38 | 53.53 | 369.643 | 28.29 | 5375 | 18.4 |  | 126.0844 | 26.3516 | 263.8246 | 197.8853 | 9.9853 | 312.0406
2010-07-28 | 15:32:48 | 53.697 | 372.393 | 28.25 | 5369 | 18.3 |  | 126.7645 | 26.4938 | 265.283 | 198.9792 | 10.0582 | 314.3182
2010-07-28 | 15:32:58 | 53.863 | 369.41 | 28.34 | 5370 | 18.3 |  | 125.7491 | 26.2816 | 263.1579 | 197.3852 | 9.9776 | 311.8002
2010-07-28 | 15:33:07 | 54.014 | 370.732 | 28.3 | 5364 | 18.3 |  | 126.1992 | 26.3756 | 264.0999 | 198.0918 | 10.0133 | 312.9164
2010-07-28 | 15:33:17 | 54.181 | 369.74 | 28.33 | 5361 | 18.3 |  | 125.8614 | 26.305 | 263.393 | 197.5616 | 9.9865 | 312.0788