## convers.+compens.
file name: PSt1_Eq4_Sal |  |  | 2008-03-29
author: Dr. Christian Huber
programm for temperature compensated oxygen calculation of membrane PSt1
Fit of the calibration curve is optimized between 0 and 250 % air-saturation (0 and 388 Torr)
enter your values here
input of F0,T0 | 58.62 | phase angle of cal 0
input of F100,T100 | 31.05 | phase angle of cal 100 (water-vapor saturated air)
input of Fm,Tm | 30.5 | measured phase angle
input of T0 | 23.3 | temperature of cal 0
input of T100 | 17.9 | temperatur of cal100
input of Tm | 17.9 | temperature at measurement
air pressure patm | 1013 | air pressure
Salinity | 32.3 | Salinity
Results of calculation + compensation
[%] air saturation | 104.8779 | %
[%] O2 | 21.9195 | %
pO2 [hPa] | 218.0572 | hPa
pO2 [Torr] | 163.5568 | Torr
cO2 [mg/L] | 8.1931 | mg/L
cO2 [ppm] | 8.1931 | ppm
cO2 [µmol/L] | 256.0334 | µmol/L
Internal calculated parameter (DO NOT CHANGE !!)
f1 | 0.801
DF/K | -0.08
DKSV/K | 0.000383
m | 22.9
tan(F0,T100) | 1.6677
tan(F0,Tm) | 1.6677
tan(F100,T100) | 0.60205
tan(Fm,Tm) | 0.58905
KSV,T100 | 0.032842
KSV,Tm | 0.032842
A | 157.6438
B | 14.279
C | -0.639
a | 1.6637e-05
b | 0.0044223
c | -0.64679
Chlorinity | 17.8781

## Recalculate Fibox values
file name: PSt1_Eq4_Sal |  |  |  | 2008-03-29
author: Dr. Christian Huber
programm for temperature compensated oxygen calculation of membrane PSt1
Fit of the calibration curve is optimized between 0 and 250 % air-saturation (0 and 388 Torr)
Insert your calibration values
cal0 | 58.62 | ° | T0 | 23.3 | °C
cal100 | 31.05 | ° | T100 | 17.9 | °C
air pressure | 1020 | mbar
Salinity | 32.3 | °/°°
Internal calculated parameter (DO NOT CHANGE !!)
f1 | 0.801 | tan(F0,T100) | 1.6677 | A | 157.6438 | KSV,T100 | 0.032842 | a | 1.6696e-05
DF/K | -0.08 | tan(F0,Tm) | 1.6677 | B | 14.279 | KSV,Tm | 0.032842 | b | 0.0044658 |  |  |  |  | INSERT YOUR RESPECTIVE DATA
DKSV/K | 0.000383 | tan(F100,T100) | 0.60205 | C | -0.639 |  |  | c | -0.64552 |  |  |  |  | chamber volume [L]
m | 22.9 | tan(Fm,Tm) | 0.59116 |  |  |  |  | Chlorinity | 17.8781 |  |  |  |  | 0.026
Insert the data you get from the Ascii File recorded with the OxyView TX3-V5.31.exe software | [7 empty] | The oxygen contents are calculated automatically from the raw data
date | time/hh:mm:ss | logtime/min | oxygen/% airsatur. | phase/° | amp | temp/°C |  | % air-sat. | % oxygen | pO2 (hPa) | pO2 (Torr) | cO2 [mg/L] | cO2 [µmol/L] | cO2 [µmol/L 10 min] | cO2 [µmol/L h] | cO2 [µmol/ h]
2010-08-23 | 11:04:36 | 0 | 104.062 | 30.59 | 5933 | 17.9 |  | 104.0616 | 21.7489 | 217.8861 | 163.4285 | 8.1866 | 255.8324 | -0.306 | -1.836 | -0.047736
2010-08-23 | 11:04:47 | 0.184 | 104.152 | 30.58 | 5929 | 17.9 |  | 104.1519 | 21.7678 | 218.0753 | 163.5704 | 8.1937 | 256.0545
2010-08-23 | 11:04:57 | 0.35 | 104.242 | 30.57 | 5932 | 17.9 |  | 104.2424 | 21.7867 | 218.2646 | 163.7124 | 8.2009 | 256.2768 | change data input according to column N
2010-08-23 | 11:05:07 | 0.517 | 103.881 | 30.61 | 5931 | 17.9 |  | 103.8812 | 21.7112 | 217.5084 | 163.1452 | 8.1724 | 255.3888
2010-08-23 | 11:05:17 | 0.684 | 104.605 | 30.53 | 5928 | 17.9 |  | 104.605 | 21.8624 | 219.0239 | 164.2819 | 8.2294 | 257.1682
2010-08-23 | 11:05:27 | 0.851 | 104.696 | 30.52 | 5925 | 17.9 |  | 104.6958 | 21.8814 | 219.2141 | 164.4246 | 8.2365 | 257.3917
2010-08-23 | 11:05:37 | 1.018 | 105.243 | 30.46 | 5928 | 17.9 |  | 105.243 | 21.9958 | 220.3597 | 165.2838 | 8.2796 | 258.7367
2010-08-23 | 11:05:47 | 1.185 | 105.426 | 30.44 | 5925 | 17.9 |  | 105.4261 | 22.034 | 220.7431 | 165.5714 | 8.294 | 259.1869
2010-08-23 | 11:05:57 | 1.352 | 104.605 | 30.53 | 5922 | 17.9 |  | 104.605 | 21.8624 | 219.0239 | 164.2819 | 8.2294 | 257.1682
2010-08-23 | 11:06:07 | 1.519 | 105.61 | 30.42 | 5925 | 17.9 |  | 105.6095 | 22.0724 | 221.1272 | 165.8595 | 8.3084 | 259.6379
2010-08-23 | 11:06:17 | 1.686 | 104.878 | 30.5 | 5931 | 17.9 |  | 104.8779 | 21.9195 | 219.5952 | 164.7104 | 8.2509 | 257.8391
2010-08-23 | 11:06:27 | 1.852 | 103.701 | 30.63 | 5923 | 17.9 |  | 103.7011 | 21.6735 | 217.1314 | 162.8624 | 8.1583 | 254.9462
2010-08-23 | 11:06:37 | 2.02 | 104.969 | 30.49 | 5931 | 17.9 |  | 104.969 | 21.9385 | 219.7861 | 164.8536 | 8.258 | 258.0632
2010-08-23 | 11:06:47 | 2.186 | 104.333 | 30.56 | 5926 | 17.9 |  | 104.3329 | 21.8056 | 218.4542 | 163.8545 | 8.208 | 256.4993
2010-08-23 | 11:06:57 | 2.353 | 104.423 | 30.55 | 5917 | 17.9 |  | 104.4235 | 21.8245 | 218.6439 | 163.9968 | 8.2151 | 256.7221
2010-08-23 | 11:07:07 | 2.52 | 104.152 | 30.58 | 5913 | 17.9 |  | 104.1519 | 21.7678 | 218.0753 | 163.5704 | 8.1937 | 256.0545
2010-08-23 | 11:07:17 | 2.687 | 104.05 | 30.57 | 5924 | 18 |  | 104.0492 | 21.7463 | 217.8319 | 163.3878 | 8.1702 | 255.3201
2010-08-23 | 11:07:27 | 2.854 | 104.141 | 30.56 | 5925 | 18 |  | 104.1396 | 21.7652 | 218.021 | 163.5297 | 8.1773 | 255.5418
2010-08-23 | 11:07:37 | 3.021 | 104.231 | 30.55 | 5926 | 18 |  | 104.23 | 21.7841 | 218.2104 | 163.6717 | 8.1844 | 255.7637
2010-08-23 | 11:07:47 | 3.188 | 103.69 | 30.61 | 5915 | 18 |  | 103.6887 | 21.6709 | 217.0771 | 162.8217 | 8.1419 | 254.4354
2010-08-23 | 11:07:57 | 3.355 | 104.685 | 30.5 | 5930 | 18 |  | 104.6836 | 21.8789 | 219.1599 | 164.3839 | 8.2201 | 256.8767
2010-08-23 | 11:08:07 | 3.522 | 104.231 | 30.55 | 5924 | 18 |  | 104.23 | 21.7841 | 218.2104 | 163.6717 | 8.1844 | 255.7637
2010-08-23 | 11:08:17 | 3.689 | 104.412 | 30.53 | 5928 | 18 |  | 104.4112 | 21.8219 | 218.5897 | 163.9562 | 8.1987 | 256.2082 |  | regression formula | time
2010-08-23 | 11:08:27 | 3.856 | 104.855 | 30.46 | 5913 | 18.1 |  | 104.8536 | 21.9144 | 219.4872 | 164.6294 | 8.2179 | 256.8097 | value for T=26 min. | 256.2989 | T11
2010-08-23 | 11:08:37 | 4.022 | 103.858 | 30.57 | 5917 | 18.1 |  | 103.8565 | 21.706 | 217.4001 | 163.064 | 8.1398 | 254.3678 | value for T=1 min. | 256.6049 | T1
2010-08-23 | 11:08:47 | 4.189 | 103.588 | 30.6 | 5927 | 18.1 |  | 103.5865 | 21.6496 | 216.8349 | 162.64 | 8.1186 | 253.7064 | difference between T25 and T1 | -0.306 | 10minutes
2010-08-23 | 11:08:57 | 4.356 | 104.855 | 30.46 | 5918 | 18.1 |  | 104.8536 | 21.9144 | 219.4872 | 164.6294 | 8.2179 | 256.8097 | calculate from regression curve values for 10 minutes of photosynthesis or respiration
2010-08-23 | 11:09:07 | 4.523 | 103.319 | 30.63 | 5924 | 18.1 |  | 103.3172 | 21.5933 | 216.2712 | 162.2172 | 8.0975 | 253.0469
2010-08-23 | 11:09:17 | 4.69 | 104.039 | 30.55 | 5916 | 18.1 |  | 104.037 | 21.7437 | 217.7779 | 163.3473 | 8.1539 | 254.8098
2010-08-23 | 11:09:27 | 4.857 | 103.858 | 30.57 | 5927 | 18.1 |  | 103.8565 | 21.706 | 217.4001 | 163.064 | 8.1398 | 254.3678
2010-08-23 | 11:09:37 | 5.024 | 103.949 | 30.56 | 5920 | 18.1 |  | 103.9467 | 21.7249 | 217.5889 | 163.2056 | 8.1468 | 254.5887
2010-08-23 | 11:09:47 | 5.191 | 103.69 | 30.61 | 5927 | 18 |  | 103.6887 | 21.6709 | 217.0771 | 162.8217 | 8.1419 | 254.4354
2010-08-23 | 11:09:57 | 5.358 | 104.503 | 30.52 | 5917 | 18 |  | 104.5019 | 21.8409 | 218.7796 | 164.0986 | 8.2058 | 256.4308
2010-08-23 | 11:10:07 | 5.525 | 104.231 | 30.55 | 5936 | 18 |  | 104.23 | 21.7841 | 218.2104 | 163.6717 | 8.1844 | 255.7637
2010-08-23 | 11:10:17 | 5.692 | 104.775 | 30.49 | 5914 | 18 |  | 104.7745 | 21.8979 | 219.3504 | 164.5268 | 8.2272 | 257.0999
2010-08-23 | 11:10:27 | 5.859 | 103.78 | 30.6 | 5923 | 18 |  | 103.7787 | 21.6897 | 217.2655 | 162.963 | 8.149 | 254.6562
2010-08-23 | 11:10:37 | 6.025 | 104.605 | 30.53 | 5923 | 17.9 |  | 104.605 | 21.8624 | 219.0239 | 164.2819 | 8.2294 | 257.1682
2010-08-23 | 11:10:47 | 6.192 | 105.426 | 30.44 | 5926 | 17.9 |  | 105.4261 | 22.034 | 220.7431 | 165.5714 | 8.294 | 259.1869
2010-08-23 | 11:10:57 | 6.359 | 104.605 | 30.53 | 5924 | 17.9 |  | 104.605 | 21.8624 | 219.0239 | 164.2819 | 8.2294 | 257.1682
2010-08-23 | 11:11:07 | 6.526 | 104.707 | 30.54 | 5928 | 17.8 |  | 104.7084 | 21.884 | 219.2686 | 164.4654 | 8.2531 | 257.9087
2010-08-23 | 11:11:17 | 6.693 | 104.345 | 30.58 | 5920 | 17.8 |  | 104.3455 | 21.8082 | 218.5087 | 163.8954 | 8.2245 | 257.0148
2010-08-23 | 11:11:27 | 6.86 | 104.617 | 30.55 | 5929 | 17.8 |  | 104.6175 | 21.8651 | 219.0784 | 164.3227 | 8.2459 | 257.6849
2010-08-23 | 11:11:37 | 7.027 | 104.345 | 30.58 | 5925 | 17.8 |  | 104.3455 | 21.8082 | 218.5087 | 163.8954 | 8.2245 | 257.0148
2010-08-23 | 11:11:47 | 7.194 | 104.707 | 30.54 | 5925 | 17.8 |  | 104.7084 | 21.884 | 219.2686 | 164.4654 | 8.2531 | 257.9087
2010-08-23 | 11:11:57 | 7.36 | 104.798 | 30.53 | 5928 | 17.8 |  | 104.7993 | 21.9031 | 219.4591 | 164.6083 | 8.2602 | 258.1327
2010-08-23 | 11:12:07 | 7.527 | 104.345 | 30.58 | 5926 | 17.8 |  | 104.3455 | 21.8082 | 218.5087 | 163.8954 | 8.2245 | 257.0148
2010-08-23 | 11:12:17 | 7.694 | 105.163 | 30.49 | 5932 | 17.8 |  | 105.164 | 21.9793 | 220.2228 | 165.1811 | 8.289 | 259.031
2010-08-23 | 11:12:27 | 7.861 | 104.526 | 30.56 | 5927 | 17.8 |  | 104.5267 | 21.8461 | 218.8883 | 164.1802 | 8.2388 | 257.4613
2010-08-23 | 11:12:37 | 8.028 | 104.707 | 30.54 | 5927 | 17.8 |  | 104.7084 | 21.884 | 219.2686 | 164.4654 | 8.2531 | 257.9087
2010-08-23 | 11:12:47 | 8.195 | 104.889 | 30.52 | 5928 | 17.8 |  | 104.8903 | 21.9221 | 219.6497 | 164.7513 | 8.2674 | 258.3569
2010-08-23 | 11:12:57 | 8.362 | 103.803 | 30.64 | 5930 | 17.8 |  | 103.8038 | 21.695 | 217.3744 | 163.0446 | 8.1818 | 255.6806
2010-08-23 | 11:13:08 | 8.529 | 104.345 | 30.58 | 5929 | 17.8 |  | 104.3455 | 21.8082 | 218.5087 | 163.8954 | 8.2245 | 257.0148
2010-08-23 | 11:13:18 | 8.696 | 104.423 | 30.55 | 5933 | 17.9 |  | 104.4235 | 21.8245 | 218.6439 | 163.9968 | 8.2151 | 256.7221
2010-08-23 | 11:13:27 | 8.846 | 103.521 | 30.65 | 5931 | 17.9 |  | 103.5214 | 21.636 | 216.7551 | 162.5801 | 8.1441 | 254.5044
2010-08-23 | 11:13:37 | 9.013 | 104.152 | 30.58 | 5938 | 17.9 |  | 104.1519 | 21.7678 | 218.0753 | 163.5704 | 8.1937 | 256.0545
2010-08-23 | 11:13:47 | 9.18 | 104.696 | 30.52 | 5925 | 17.9 |  | 104.6958 | 21.8814 | 219.2141 | 164.4246 | 8.2365 | 257.3917
2010-08-23 | 11:13:57 | 9.347 | 104.242 | 30.57 | 5932 | 17.9 |  | 104.2424 | 21.7867 | 218.2646 | 163.7124 | 8.2009 | 256.2768
2010-08-23 | 11:14:07 | 9.514 | 103.791 | 30.62 | 5933 | 17.9 |  | 103.7911 | 21.6923 | 217.3198 | 163.0037 | 8.1654 | 255.1674
2010-08-23 | 11:14:17 | 9.681 | 103.971 | 30.6 | 5934 | 17.9 |  | 103.9714 | 21.73 | 217.6972 | 163.2868 | 8.1795 | 255.6105
2010-08-23 | 11:14:27 | 9.847 | 104.605 | 30.53 | 5938 | 17.9 |  | 104.605 | 21.8624 | 219.0239 | 164.2819 | 8.2294 | 257.1682
2010-08-23 | 11:14:37 | 10.014 | 104.242 | 30.57 | 5935 | 17.9 |  | 104.2424 | 21.7867 | 218.2646 | 163.7124 | 8.2009 | 256.2768
2010-08-23 | 11:14:47 | 10.181 | 104.423 | 30.55 | 5939 | 17.9 |  | 104.4235 | 21.8245 | 218.6439 | 163.9968 | 8.2151 | 256.7221
2010-08-23 | 11:14:57 | 10.348 | 102.895 | 30.72 | 5941 | 17.9 |  | 102.8952 | 21.5051 | 215.4439 | 161.5967 | 8.0949 | 252.9649