## convers.+compens.
file name: PSt1_Eq4_Sal |  |  | 2008-03-29
author: Dr. Christian Huber
programm for temperature compensated oxygen calculation of membrane PSt1
Fit of the calibration curve is optimized between 0 and 250 % air-saturation (0 and 388 Torr)
enter your values here
input of F0,T0 | 59.26 | phase angle of cal 0
input of F100,T100 | 30.25 | phase angle of cal 100 (water-vapor saturated air)
input of Fm,Tm | 28.6 | measured phase angle
input of T0 | 18 | temperature of cal 0
input of T100 | 17.8 | temperatur of cal100
input of Tm | 16.7 | temperature at measurement
air pressure patm | 1013 | air pressure
Salinity | 29.8 | Salinity
Results of calculation + compensation
[%] air saturation | 117.8706 | %
[%] O2 | 24.635 | %
pO2 [hPa] | 245.4419 | hPa
pO2 [Torr] | 184.0971 | Torr
cO2 [mg/L] | 9.5755 | mg/L
cO2 [ppm] | 9.5755 | ppm
cO2 [µmol/L] | 299.2331 | µmol/L
Internal calculated parameter (DO NOT CHANGE !!)
f1 | 0.801
DF/K | -0.08
DKSV/K | 0.000383
m | 22.9
tan(F0,T100) | 1.6826
tan(F0,Tm) | 1.6885
tan(F100,T100) | 0.58318
tan(Fm,Tm) | 0.54522
KSV,T100 | 0.035887
KSV,Tm | 0.035466
A | 151.3533
B | 12.7756
C | -0.6534
a | 1.7736e-05
b | 0.0036539
c | -0.6771
Chlorinity | 16.4931

## Recalculate Fibox values
file name: PSt1_Eq4_Sal |  |  |  | 2008-03-29
author: Dr. Christian Huber
programm for temperature compensated oxygen calculation of membrane PSt1
Fit of the calibration curve is optimized between 0 and 250 % air-saturation (0 and 388 Torr)
Insert your calibration values
cal0 | 59.26 | ° | T0 | 17.8 | °C
cal100 | 30.25 | ° | T100 | 17.7 | °C
air pressure | 1020 | mbar
Salinity | 29.8 | °/°°
Internal calculated parameter (DO NOT CHANGE !!)
f1 | 0.801 | tan(F0,T100) | 1.6821 | A | 151.4014 | KSV,T100 | 0.035862 | a | 1.7798e-05
DF/K | -0.08 | tan(F0,Tm) | 1.6874 | B | 12.7871 | KSV,Tm | 0.035479 | b | 0.0036877 | [5 empty] | INSERT YOUR RESPECTIVE DATA
DKSV/K | 0.000383 | tan(F100,T100) | 0.58318 | C | -0.65329 |  |  | c | -0.67622 |  |  |  | Blank Respiration | Blank (chamber 1) cO2 [µmol/ h] | chamber volume [L] | Fresh weight alga [g]
m | 22.9 | tan(Fm,Tm) | 0.54635 |  |  |  |  | Chlorinity | 16.4931 |  |  |  | -5.265 | -0.82291 | 0.025873 | 0.0426
Insert the data you get from the Ascii File recorded with the OxyView TX3-V5.31.exe software | [7 empty] | The oxygen contents are calculated automatically from the raw data
date | time/hh:mm:ss | logtime/min | oxygen/% airsatur. | phase/° | amp | temp/°C |  | % air-sat. | % oxygen | pO2 (hPa) | pO2 (Torr) | cO2 [mg/L] | cO2 [µmol/L] |  | cO2 [µmol/L 10 min] | cO2 [µmol/L h] | cO2 [µmol/ h] | cO2 [µmol/ h gFW]
2010-07-28 | 16:57:52 | 0 | 357.114 | 28.65 | 3336 | 16.7 |  | 117.1429 | 24.4829 | 245.6445 | 184.2491 | 9.5834 | 299.48 |  | 29.073 | 174.438 | 5.3361 | 125.2608
2010-07-28 | 16:58:04 | 0.201 | 353.956 | 28.75 | 3337 | 16.7 |  | 116.107 | 24.2664 | 243.4723 | 182.6198 | 9.4986 | 296.8318
2010-07-28 | 16:58:14 | 0.368 | 353.015 | 28.78 | 3341 | 16.7 |  | 115.7983 | 24.2019 | 242.8251 | 182.1343 | 9.4734 | 296.0426 |  | change data input according to column N
2010-07-28 | 16:58:24 | 0.535 | 355.215 | 28.71 | 3344 | 16.7 |  | 116.5201 | 24.3527 | 244.3385 | 183.2694 | 9.5324 | 297.8878
2010-07-28 | 16:58:34 | 0.702 | 353.642 | 28.76 | 3345 | 16.7 |  | 116.004 | 24.2448 | 243.2564 | 182.4578 | 9.4902 | 296.5685
2010-07-28 | 16:58:44 | 0.869 | 355.847 | 28.69 | 3345 | 16.7 |  | 116.7272 | 24.396 | 244.7729 | 183.5953 | 9.5494 | 298.4174
2010-07-28 | 16:58:54 | 1.036 | 353.328 | 28.77 | 3348 | 16.7 |  | 115.9011 | 24.2233 | 243.0406 | 182.2959 | 9.4818 | 296.3054
2010-07-28 | 16:59:04 | 1.202 | 353.015 | 28.78 | 3349 | 16.7 |  | 115.7983 | 24.2019 | 242.8251 | 182.1343 | 9.4734 | 296.0426
2010-07-28 | 16:59:14 | 1.369 | 355.531 | 28.7 | 3345 | 16.7 |  | 116.6236 | 24.3743 | 244.5556 | 183.4323 | 9.5409 | 298.1524
2010-07-28 | 16:59:24 | 1.536 | 353.015 | 28.78 | 3343 | 16.7 |  | 115.7983 | 24.2019 | 242.8251 | 182.1343 | 9.4734 | 296.0426
2010-07-28 | 16:59:34 | 1.703 | 354.585 | 28.73 | 3347 | 16.7 |  | 116.3133 | 24.3095 | 243.905 | 182.9443 | 9.5155 | 297.3592
2010-07-28 | 16:59:44 | 1.87 | 353.328 | 28.77 | 3348 | 16.7 |  | 115.9011 | 24.2233 | 243.0406 | 182.2959 | 9.4818 | 296.3054
2010-07-28 | 16:59:54 | 2.037 | 352.702 | 28.79 | 3352 | 16.7 |  | 115.6957 | 24.1804 | 242.6098 | 181.9728 | 9.465 | 295.7802
2010-07-28 | 17:00:04 | 2.204 | 353.956 | 28.75 | 3348 | 16.7 |  | 116.107 | 24.2664 | 243.4723 | 182.6198 | 9.4986 | 296.8318
2010-07-28 | 17:00:14 | 2.371 | 353.328 | 28.77 | 3352 | 16.7 |  | 115.9011 | 24.2233 | 243.0406 | 182.2959 | 9.4818 | 296.3054
2010-07-28 | 17:00:24 | 2.538 | 350.831 | 28.85 | 3343 | 16.7 |  | 115.0819 | 24.0521 | 241.3226 | 181.0074 | 9.4147 | 294.2109
2010-07-28 | 17:00:34 | 2.705 | 352.389 | 28.8 | 3351 | 16.7 |  | 115.5931 | 24.159 | 242.3947 | 181.8115 | 9.4566 | 295.5179
2010-07-28 | 17:00:44 | 2.871 | 354.793 | 28.68 | 3349 | 16.8 |  | 116.6254 | 24.3747 | 244.5297 | 183.4129 | 9.5226 | 297.5824
2010-07-28 | 17:00:54 | 3.038 | 351.345 | 28.79 | 3345 | 16.8 |  | 115.492 | 24.1378 | 242.1534 | 181.6305 | 9.4301 | 294.6905
2010-07-28 | 17:01:04 | 3.205 | 351.345 | 28.79 | 3345 | 16.8 |  | 115.492 | 24.1378 | 242.1534 | 181.6305 | 9.4301 | 294.6905
2010-07-28 | 17:01:14 | 3.372 | 348.862 | 28.87 | 3348 | 16.8 |  | 114.6758 | 23.9673 | 240.4421 | 180.3469 | 9.3635 | 292.6079
2010-07-28 | 17:01:24 | 3.539 | 351.345 | 28.79 | 3351 | 16.8 |  | 115.492 | 24.1378 | 242.1534 | 181.6305 | 9.4301 | 294.6905
2010-07-28 | 17:01:34 | 3.706 | 353.221 | 28.73 | 3354 | 16.8 |  | 116.1086 | 24.2667 | 243.4462 | 182.6002 | 9.4804 | 296.2638 |  |  | regression formula | time
2010-07-28 | 17:01:44 | 3.873 | 350.723 | 28.81 | 3354 | 16.8 |  | 115.2873 | 24.0951 | 241.7242 | 181.3086 | 9.4134 | 294.1682 |  | value for T=26 min. | 330.964 | T11
2010-07-28 | 17:01:54 | 4.04 | 350.412 | 28.82 | 3348 | 16.8 |  | 115.1852 | 24.0737 | 241.51 | 181.1479 | 9.405 | 293.9075 |  | value for T=1 min. | 307.156 | T1
2010-07-28 | 17:02:04 | 4.207 | 354.793 | 28.68 | 3354 | 16.8 |  | 116.6254 | 24.3747 | 244.5297 | 183.4129 | 9.5226 | 297.5824 |  | difference between T25 and T1 | 23.808 | 10minutes
2010-07-28 | 17:02:14 | 4.374 | 350.723 | 28.81 | 3350 | 16.8 |  | 115.2873 | 24.0951 | 241.7242 | 181.3086 | 9.4134 | 294.1682 |  | calculate from regression curve values for 10 minutes of photosynthesis or respiration
2010-07-28 | 17:02:24 | 4.541 | 348.245 | 28.89 | 3357 | 16.8 |  | 114.4729 | 23.9248 | 240.0165 | 180.0277 | 9.3469 | 292.09
2010-07-28 | 17:02:34 | 4.708 | 353.328 | 28.77 | 3349 | 16.7 |  | 115.9011 | 24.2233 | 243.0406 | 182.2959 | 9.4818 | 296.3054
2010-07-28 | 17:02:44 | 4.875 | 354.27 | 28.74 | 3356 | 16.7 |  | 116.2101 | 24.2879 | 243.6885 | 182.7819 | 9.5071 | 297.0953
2010-07-28 | 17:02:54 | 5.041 | 351.142 | 28.84 | 3354 | 16.7 |  | 115.1839 | 24.0734 | 241.5366 | 181.1678 | 9.4231 | 294.4718
2010-07-28 | 17:03:04 | 5.208 | 349.9 | 28.88 | 3355 | 16.7 |  | 114.7764 | 23.9883 | 240.682 | 180.5269 | 9.3898 | 293.43
2010-07-28 | 17:03:15 | 5.375 | 353.328 | 28.77 | 3362 | 16.7 |  | 115.9011 | 24.2233 | 243.0406 | 182.2959 | 9.4818 | 296.3054
2010-07-28 | 17:03:25 | 5.542 | 353.328 | 28.77 | 3351 | 16.7 |  | 115.9011 | 24.2233 | 243.0406 | 182.2959 | 9.4818 | 296.3054
2010-07-28 | 17:03:35 | 5.709 | 351.142 | 28.84 | 3360 | 16.7 |  | 115.1839 | 24.0734 | 241.5366 | 181.1678 | 9.4231 | 294.4718
2010-07-28 | 17:03:45 | 5.876 | 352.702 | 28.79 | 3361 | 16.7 |  | 115.6957 | 24.1804 | 242.6098 | 181.9728 | 9.465 | 295.7802
2010-07-28 | 17:03:55 | 6.043 | 353.956 | 28.75 | 3357 | 16.7 |  | 116.107 | 24.2664 | 243.4723 | 182.6198 | 9.4986 | 296.8318
2010-07-28 | 17:04:05 | 6.21 | 352.702 | 28.79 | 3358 | 16.7 |  | 115.6957 | 24.1804 | 242.6098 | 181.9728 | 9.465 | 295.7802
2010-07-28 | 17:04:15 | 6.377 | 352.389 | 28.8 | 3357 | 16.7 |  | 115.5931 | 24.159 | 242.3947 | 181.8115 | 9.4566 | 295.5179
2010-07-28 | 17:04:25 | 6.544 | 349.59 | 28.89 | 3366 | 16.7 |  | 114.6748 | 23.967 | 240.469 | 180.3671 | 9.3814 | 293.1702
2010-07-28 | 17:04:35 | 6.71 | 349.59 | 28.89 | 3364 | 16.7 |  | 114.6748 | 23.967 | 240.469 | 180.3671 | 9.3814 | 293.1702
2010-07-28 | 17:04:44 | 6.861 | 354.27 | 28.74 | 3356 | 16.7 |  | 116.2101 | 24.2879 | 243.6885 | 182.7819 | 9.5071 | 297.0953
2010-07-28 | 17:04:54 | 7.028 | 352.077 | 28.81 | 3361 | 16.7 |  | 115.4906 | 24.1375 | 242.1798 | 181.6503 | 9.4482 | 295.256
2010-07-28 | 17:05:04 | 7.195 | 353.642 | 28.76 | 3358 | 16.7 |  | 116.004 | 24.2448 | 243.2564 | 182.4578 | 9.4902 | 296.5685
2010-07-28 | 17:05:14 | 7.361 | 348.354 | 28.93 | 3363 | 16.7 |  | 114.2694 | 23.8823 | 239.6188 | 179.7294 | 9.3483 | 292.1337
2010-07-28 | 17:05:24 | 7.528 | 350.831 | 28.85 | 3362 | 16.7 |  | 115.0819 | 24.0521 | 241.3226 | 181.0074 | 9.4147 | 294.2109
2010-07-28 | 17:05:34 | 7.695 | 352.077 | 28.81 | 3356 | 16.7 |  | 115.4906 | 24.1375 | 242.1798 | 181.6503 | 9.4482 | 295.256
2010-07-28 | 17:05:44 | 7.862 | 350.21 | 28.87 | 3359 | 16.7 |  | 114.8781 | 24.0095 | 240.8953 | 180.6869 | 9.3981 | 293.69
2010-07-28 | 17:05:54 | 8.029 | 349.59 | 28.89 | 3369 | 16.7 |  | 114.6748 | 23.967 | 240.469 | 180.3671 | 9.3814 | 293.1702
2010-07-28 | 17:06:04 | 8.196 | 351.142 | 28.84 | 3359 | 16.7 |  | 115.1839 | 24.0734 | 241.5366 | 181.1678 | 9.4231 | 294.4718
2010-07-28 | 17:06:14 | 8.363 | 351.453 | 28.83 | 3361 | 16.7 |  | 115.286 | 24.0948 | 241.7508 | 181.3285 | 9.4315 | 294.7329
2010-07-28 | 17:06:24 | 8.53 | 347.738 | 28.95 | 3362 | 16.7 |  | 114.0673 | 23.8401 | 239.1951 | 179.4116 | 9.3317 | 291.6171
2010-07-28 | 17:06:34 | 8.697 | 348.046 | 28.94 | 3369 | 16.7 |  | 114.1683 | 23.8612 | 239.4069 | 179.5704 | 9.34 | 291.8753
2010-07-28 | 17:06:44 | 8.864 | 352.702 | 28.79 | 3367 | 16.7 |  | 115.6957 | 24.1804 | 242.6098 | 181.9728 | 9.465 | 295.7802
2010-07-28 | 17:06:54 | 9.031 | 351.453 | 28.83 | 3371 | 16.7 |  | 115.286 | 24.0948 | 241.7508 | 181.3285 | 9.4315 | 294.7329
2010-07-28 | 17:07:04 | 9.197 | 349.59 | 28.89 | 3368 | 16.7 |  | 114.6748 | 23.967 | 240.469 | 180.3671 | 9.3814 | 293.1702
2010-07-28 | 17:07:14 | 9.364 | 352.702 | 28.79 | 3361 | 16.7 |  | 115.6957 | 24.1804 | 242.6098 | 181.9728 | 9.465 | 295.7802
2010-07-28 | 17:07:24 | 9.531 | 351.142 | 28.84 | 3364 | 16.7 |  | 115.1839 | 24.0734 | 241.5366 | 181.1678 | 9.4231 | 294.4718
2010-07-28 | 17:07:34 | 9.698 | 349.59 | 28.89 | 3367 | 16.7 |  | 114.6748 | 23.967 | 240.469 | 180.3671 | 9.3814 | 293.1702
2010-07-28 | 17:07:44 | 9.865 | 351.142 | 28.84 | 3376 | 16.7 |  | 115.1839 | 24.0734 | 241.5366 | 181.1678 | 9.4231 | 294.4718
2010-07-28 | 17:07:54 | 10.032 | 350.21 | 28.87 | 3369 | 16.7 |  | 114.8781 | 24.0095 | 240.8953 | 180.6869 | 9.3981 | 293.69
2010-07-28 | 17:08:04 | 10.199 | 349.59 | 28.89 | 3369 | 16.7 |  | 114.6748 | 23.967 | 240.469 | 180.3671 | 9.3814 | 293.1702
2010-07-28 | 17:08:14 | 10.366 | 348.662 | 28.92 | 3363 | 16.7 |  | 114.3706 | 23.9034 | 239.831 | 179.8886 | 9.3566 | 292.3924
2010-07-28 | 17:08:24 | 10.532 | 352.702 | 28.79 | 3371 | 16.7 |  | 115.6957 | 24.1804 | 242.6098 | 181.9728 | 9.465 | 295.7802
2010-07-28 | 17:08:34 | 10.699 | 350.321 | 28.91 | 3370 | 16.6 |  | 114.6739 | 23.9668 | 240.4962 | 180.3875 | 9.3995 | 293.7349
2010-07-28 | 17:08:44 | 10.866 | 353.126 | 28.82 | 3373 | 16.6 |  | 115.5919 | 24.1587 | 242.4214 | 181.8315 | 9.4748 | 296.0863
2010-07-28 | 17:08:54 | 11.033 | 352.5 | 28.84 | 3366 | 16.6 |  | 115.3872 | 24.1159 | 241.992 | 181.5094 | 9.458 | 295.5619
2010-07-28 | 17:09:04 | 11.2 | 353.552 | 28.85 | 3370 | 16.5 |  | 115.4886 | 24.1371 | 242.2337 | 181.6907 | 9.4846 | 296.3944
2010-07-28 | 17:09:14 | 11.367 | 352.5 | 28.84 | 3376 | 16.6 |  | 115.3872 | 24.1159 | 241.992 | 181.5094 | 9.458 | 295.5619
2010-07-28 | 17:09:24 | 11.534 | 349.392 | 28.94 | 3373 | 16.6 |  | 114.3698 | 23.9033 | 239.8584 | 179.9091 | 9.3746 | 292.9559
2010-07-28 | 17:09:34 | 11.701 | 351.253 | 28.88 | 3372 | 16.6 |  | 114.979 | 24.0306 | 241.1359 | 180.8673 | 9.4245 | 294.5162
2010-07-28 | 17:09:44 | 11.868 | 348.158 | 28.98 | 3375 | 16.6 |  | 113.9658 | 23.8189 | 239.0111 | 179.2736 | 9.3415 | 291.921
2010-07-28 | 17:09:54 | 12.035 | 352.188 | 28.85 | 3372 | 16.6 |  | 115.285 | 24.0946 | 241.7777 | 181.3487 | 9.4496 | 295.3
2010-07-28 | 17:10:04 | 12.202 | 351.876 | 28.86 | 3372 | 16.6 |  | 115.1829 | 24.0732 | 241.5635 | 181.188 | 9.4412 | 295.0385
2010-07-28 | 17:10:14 | 12.368 | 348.774 | 28.96 | 3376 | 16.6 |  | 114.1676 | 23.861 | 239.4343 | 179.591 | 9.358 | 292.4379
2010-07-28 | 17:10:24 | 12.536 | 352.926 | 28.87 | 3374 | 16.5 |  | 115.2841 | 24.0944 | 241.8049 | 181.3691 | 9.4678 | 295.8696
2010-07-28 | 17:10:34 | 12.702 | 350.435 | 28.95 | 3375 | 16.5 |  | 114.4705 | 23.9243 | 240.0984 | 180.0891 | 9.401 | 293.7816
2010-07-28 | 17:10:44 | 12.869 | 351.253 | 28.88 | 3374 | 16.6 |  | 114.979 | 24.0306 | 241.1359 | 180.8673 | 9.4245 | 294.5162
2010-07-28 | 17:10:54 | 13.036 | 349.701 | 28.93 | 3373 | 16.6 |  | 114.4711 | 23.9245 | 240.0708 | 180.0684 | 9.3829 | 293.2153
2010-07-28 | 17:11:04 | 13.203 | 347.236 | 29.01 | 3372 | 16.6 |  | 113.6639 | 23.7558 | 238.3779 | 178.7986 | 9.3167 | 291.1477
2010-07-28 | 17:11:14 | 13.37 | 350.942 | 28.89 | 3377 | 16.6 |  | 114.8772 | 24.0093 | 240.9224 | 180.7072 | 9.4162 | 294.2555
2010-07-28 | 17:11:24 | 13.537 | 350.321 | 28.91 | 3378 | 16.6 |  | 114.6739 | 23.9668 | 240.4962 | 180.3875 | 9.3995 | 293.7349
2010-07-28 | 17:11:34 | 13.704 | 349.392 | 28.94 | 3377 | 16.6 |  | 114.3698 | 23.9033 | 239.8584 | 179.9091 | 9.3746 | 292.9559
2010-07-28 | 17:11:44 | 13.871 | 350.631 | 28.9 | 3375 | 16.6 |  | 114.7755 | 23.9881 | 240.7092 | 180.5472 | 9.4078 | 293.995
2010-07-28 | 17:11:54 | 14.038 | 347.543 | 29 | 3378 | 16.6 |  | 113.7644 | 23.7768 | 238.5888 | 178.9568 | 9.325 | 291.4052
2010-07-28 | 17:12:04 | 14.205 | 347.543 | 29 | 3380 | 16.6 |  | 113.7644 | 23.7768 | 238.5888 | 178.9568 | 9.325 | 291.4052
2010-07-28 | 17:12:14 | 14.371 | 349.083 | 28.95 | 3378 | 16.6 |  | 114.2686 | 23.8821 | 239.6462 | 179.75 | 9.3663 | 292.6968
2010-07-28 | 17:12:24 | 14.538 | 347.236 | 29.01 | 3378 | 16.6 |  | 113.6639 | 23.7558 | 238.3779 | 178.7986 | 9.3167 | 291.1477
2010-07-28 | 17:12:34 | 14.705 | 350.942 | 28.89 | 3378 | 16.6 |  | 114.8772 | 24.0093 | 240.9224 | 180.7072 | 9.4162 | 294.2555
2010-07-28 | 17:12:44 | 14.872 | 349.701 | 28.93 | 3376 | 16.6 |  | 114.4711 | 23.9245 | 240.0708 | 180.0684 | 9.3829 | 293.2153
2010-07-28 | 17:12:54 | 15.039 | 346.929 | 29.02 | 3378 | 16.6 |  | 113.5635 | 23.7348 | 238.1673 | 178.6407 | 9.3085 | 290.8905
2010-07-28 | 17:13:04 | 15.206 | 351.253 | 28.88 | 3374 | 16.6 |  | 114.979 | 24.0306 | 241.1359 | 180.8673 | 9.4245 | 294.5162
2010-07-28 | 17:13:14 | 15.373 | 350.321 | 28.91 | 3384 | 16.6 |  | 114.6739 | 23.9668 | 240.4962 | 180.3875 | 9.3995 | 293.7349
2010-07-28 | 17:13:24 | 15.54 | 350.011 | 28.92 | 3383 | 16.6 |  | 114.5724 | 23.9456 | 240.2834 | 180.2278 | 9.3912 | 293.4749
2010-07-28 | 17:13:34 | 15.707 | 348.466 | 28.97 | 3381 | 16.6 |  | 114.0666 | 23.8399 | 239.2226 | 179.4322 | 9.3497 | 292.1794
2010-07-28 | 17:13:44 | 15.874 | 346.929 | 29.02 | 3375 | 16.6 |  | 113.5635 | 23.7348 | 238.1673 | 178.6407 | 9.3085 | 290.8905
2010-07-28 | 17:13:54 | 16.041 | 350.942 | 28.89 | 3377 | 16.6 |  | 114.8772 | 24.0093 | 240.9224 | 180.7072 | 9.4162 | 294.2555
2010-07-28 | 17:14:04 | 16.208 | 348.158 | 28.98 | 3373 | 16.6 |  | 113.9658 | 23.8189 | 239.0111 | 179.2736 | 9.3415 | 291.921
2010-07-28 | 17:14:14 | 16.375 | 348.158 | 28.98 | 3372 | 16.6 |  | 113.9658 | 23.8189 | 239.0111 | 179.2736 | 9.3415 | 291.921
2010-07-28 | 17:14:24 | 16.541 | 352.188 | 28.85 | 3372 | 16.6 |  | 115.285 | 24.0946 | 241.7777 | 181.3487 | 9.4496 | 295.3
2010-07-28 | 17:14:34 | 16.708 | 350.011 | 28.92 | 3368 | 16.6 |  | 114.5724 | 23.9456 | 240.2834 | 180.2278 | 9.3912 | 293.4749
2010-07-28 | 17:14:45 | 16.875 | 349.392 | 28.94 | 3369 | 16.6 |  | 114.3698 | 23.9033 | 239.8584 | 179.9091 | 9.3746 | 292.9559
2010-07-28 | 17:14:55 | 17.042 | 347.85 | 28.99 | 3364 | 16.6 |  | 113.8651 | 23.7978 | 238.7998 | 179.1151 | 9.3332 | 291.663
2010-07-28 | 17:15:05 | 17.209 | 348.466 | 28.97 | 3369 | 16.6 |  | 114.0666 | 23.8399 | 239.2226 | 179.4322 | 9.3497 | 292.1794
2010-07-28 | 17:15:15 | 17.376 | 350.942 | 28.89 | 3369 | 16.6 |  | 114.8772 | 24.0093 | 240.9224 | 180.7072 | 9.4162 | 294.2555
2010-07-28 | 17:15:25 | 17.543 | 348.466 | 28.97 | 3370 | 16.6 |  | 114.0666 | 23.8399 | 239.2226 | 179.4322 | 9.3497 | 292.1794
2010-07-28 | 17:15:35 | 17.709 | 349.701 | 28.93 | 3368 | 16.6 |  | 114.4711 | 23.9245 | 240.0708 | 180.0684 | 9.3829 | 293.2153
2010-07-28 | 17:15:45 | 17.876 | 348.158 | 28.98 | 3366 | 16.6 |  | 113.9658 | 23.8189 | 239.0111 | 179.2736 | 9.3415 | 291.921
2010-07-28 | 17:15:55 | 18.043 | 349.701 | 28.93 | 3367 | 16.6 |  | 114.4711 | 23.9245 | 240.0708 | 180.0684 | 9.3829 | 293.2153
2010-07-28 | 17:16:05 | 18.21 | 350.125 | 28.96 | 3363 | 16.5 |  | 114.3693 | 23.9032 | 239.886 | 179.9298 | 9.3927 | 293.5218
2010-07-28 | 17:16:14 | 18.36 | 354.807 | 28.81 | 3360 | 16.5 |  | 115.8987 | 24.2228 | 243.094 | 182.336 | 9.5183 | 297.4471
2010-07-28 | 17:16:24 | 18.527 | 351.678 | 28.91 | 3361 | 16.5 |  | 114.8765 | 24.0092 | 240.9498 | 180.7277 | 9.4344 | 294.8235
2010-07-28 | 17:16:34 | 18.694 | 351.056 | 28.93 | 3367 | 16.5 |  | 114.6733 | 23.9667 | 240.5237 | 180.4081 | 9.4177 | 294.302
2010-07-28 | 17:16:44 | 18.861 | 355.122 | 28.8 | 3357 | 16.5 |  | 116.0016 | 24.2443 | 243.3097 | 182.4978 | 9.5267 | 297.7109
2010-07-28 | 17:16:54 | 19.028 | 353.865 | 28.84 | 3361 | 16.5 |  | 115.5909 | 24.1585 | 242.4485 | 181.8518 | 9.493 | 296.6571
2010-07-28 | 17:17:04 | 19.195 | 349.701 | 28.93 | 3357 | 16.6 |  | 114.4711 | 23.9245 | 240.0708 | 180.0684 | 9.3829 | 293.2153
2010-07-28 | 17:17:14 | 19.362 | 354.807 | 28.81 | 3358 | 16.5 |  | 115.8987 | 24.2228 | 243.094 | 182.336 | 9.5183 | 297.4471
2010-07-28 | 17:17:24 | 19.529 | 352.301 | 28.89 | 3358 | 16.5 |  | 115.0801 | 24.0517 | 241.3769 | 181.0481 | 9.4511 | 295.346
2010-07-28 | 17:17:34 | 19.696 | 356.701 | 28.75 | 3356 | 16.5 |  | 116.5172 | 24.3521 | 244.3913 | 183.3091 | 9.5691 | 299.0344
2010-07-28 | 17:17:44 | 19.863 | 351.678 | 28.91 | 3352 | 16.5 |  | 114.8765 | 24.0092 | 240.9498 | 180.7277 | 9.4344 | 294.8235
2010-07-28 | 17:17:54 | 20.029 | 355.753 | 28.78 | 3352 | 16.5 |  | 116.2075 | 24.2874 | 243.7417 | 182.8218 | 9.5437 | 298.2395
2010-07-28 | 17:18:04 | 20.197 | 352.613 | 28.88 | 3356 | 16.5 |  | 115.182 | 24.073 | 241.5908 | 181.2085 | 9.4594 | 295.6077
2010-07-28 | 17:18:14 | 20.363 | 356.701 | 28.75 | 3350 | 16.5 |  | 116.5172 | 24.3521 | 244.3913 | 183.3091 | 9.5691 | 299.0344
2010-07-28 | 17:18:24 | 20.53 | 355.753 | 28.78 | 3346 | 16.5 |  | 116.2075 | 24.2874 | 243.7417 | 182.8218 | 9.5437 | 298.2395
2010-07-28 | 17:18:34 | 20.697 | 357.018 | 28.74 | 3337 | 16.5 |  | 116.6207 | 24.3737 | 244.6083 | 183.4718 | 9.5776 | 299.2999
2010-07-28 | 17:18:44 | 20.864 | 357.335 | 28.73 | 3351 | 16.5 |  | 116.7242 | 24.3954 | 244.8255 | 183.6347 | 9.5861 | 299.5657
2010-07-28 | 17:18:54 | 21.031 | 357.335 | 28.73 | 3340 | 16.5 |  | 116.7242 | 24.3954 | 244.8255 | 183.6347 | 9.5861 | 299.5657
2010-07-28 | 17:19:04 | 21.198 | 361.376 | 28.56 | 3336 | 16.6 |  | 118.2927 | 24.7232 | 248.0856 | 186.08 | 9.6961 | 303.0043
2010-07-28 | 17:19:14 | 21.365 | 357.223 | 28.69 | 3328 | 16.6 |  | 116.9332 | 24.439 | 245.2343 | 183.9414 | 9.5847 | 299.5219
2010-07-28 | 17:19:24 | 21.532 | 355.325 | 28.75 | 3333 | 16.6 |  | 116.3119 | 24.3092 | 243.9313 | 182.964 | 9.5338 | 297.9305
2010-07-28 | 17:19:34 | 21.699 | 358.176 | 28.66 | 3335 | 16.6 |  | 117.2453 | 24.5043 | 245.8888 | 184.4323 | 9.6103 | 300.3213
2010-07-28 | 17:19:44 | 21.866 | 361.376 | 28.56 | 3328 | 16.6 |  | 118.2927 | 24.7232 | 248.0856 | 186.08 | 9.6961 | 303.0043
2010-07-28 | 17:19:54 | 22.033 | 363.636 | 28.49 | 3328 | 16.6 |  | 119.0325 | 24.8778 | 249.637 | 187.2437 | 9.7568 | 304.8992
2010-07-28 | 17:20:04 | 22.199 | 358.495 | 28.65 | 3319 | 16.6 |  | 117.3495 | 24.526 | 246.1075 | 184.5963 | 9.6188 | 300.5884
2010-07-28 | 17:20:14 | 22.366 | 359.133 | 28.63 | 3326 | 16.6 |  | 117.5583 | 24.5697 | 246.5455 | 184.9248 | 9.6359 | 301.1233
2010-07-28 | 17:20:24 | 22.533 | 361.376 | 28.56 | 3321 | 16.6 |  | 118.2927 | 24.7232 | 248.0856 | 186.08 | 9.6961 | 303.0043
2010-07-28 | 17:20:34 | 22.7 | 362.343 | 28.53 | 3317 | 16.6 |  | 118.6091 | 24.7893 | 248.7491 | 186.5777 | 9.7221 | 303.8147
2010-07-28 | 17:20:44 | 22.867 | 360.734 | 28.58 | 3318 | 16.6 |  | 118.0823 | 24.6792 | 247.6444 | 185.7491 | 9.6789 | 302.4655
2010-07-28 | 17:20:54 | 23.034 | 363.96 | 28.48 | 3310 | 16.6 |  | 119.1386 | 24.9 | 249.8595 | 187.4106 | 9.7655 | 305.171
2010-07-28 | 17:21:04 | 23.201 | 363.636 | 28.49 | 3307 | 16.6 |  | 119.0325 | 24.8778 | 249.637 | 187.2437 | 9.7568 | 304.8992
2010-07-28 | 17:21:14 | 23.368 | 362.666 | 28.52 | 3303 | 16.6 |  | 118.7147 | 24.8114 | 248.9707 | 186.7439 | 9.7307 | 304.0854
2010-07-28 | 17:21:24 | 23.535 | 362.02 | 28.54 | 3309 | 16.6 |  | 118.5035 | 24.7672 | 248.5277 | 186.4116 | 9.7134 | 303.5443
2010-07-28 | 17:21:34 | 23.702 | 363.636 | 28.49 | 3305 | 16.6 |  | 119.0325 | 24.8778 | 249.637 | 187.2437 | 9.7568 | 304.8992
2010-07-28 | 17:21:44 | 23.868 | 366.239 | 28.41 | 3306 | 16.6 |  | 119.8846 | 25.0559 | 251.424 | 188.5841 | 9.8266 | 307.0818
2010-07-28 | 17:21:54 | 24.035 | 364.61 | 28.46 | 3302 | 16.6 |  | 119.3512 | 24.9444 | 250.3054 | 187.745 | 9.7829 | 305.7155
2010-07-28 | 17:22:04 | 24.202 | 362.02 | 28.54 | 3303 | 16.6 |  | 118.5035 | 24.7672 | 248.5277 | 186.4116 | 9.7134 | 303.5443
2010-07-28 | 17:22:14 | 24.369 | 366.893 | 28.39 | 3294 | 16.6 |  | 120.0987 | 25.1006 | 251.8732 | 188.9209 | 9.8442 | 307.6304
2010-07-28 | 17:22:24 | 24.536 | 366.239 | 28.41 | 3297 | 16.6 |  | 119.8846 | 25.0559 | 251.424 | 188.5841 | 9.8266 | 307.0818
2010-07-28 | 17:22:34 | 24.703 | 363.312 | 28.5 | 3294 | 16.6 |  | 118.9264 | 24.8556 | 249.4146 | 187.0769 | 9.7481 | 304.6276
2010-07-28 | 17:22:44 | 24.87 | 364.61 | 28.46 | 3294 | 16.6 |  | 119.3512 | 24.9444 | 250.3054 | 187.745 | 9.7829 | 305.7155
2010-07-28 | 17:22:54 | 25.037 | 368.206 | 28.35 | 3289 | 16.6 |  | 120.5284 | 25.1904 | 252.7742 | 189.5968 | 9.8794 | 308.7309
2010-07-28 | 17:23:04 | 25.204 | 367.657 | 28.41 | 3286 | 16.5 |  | 120.096 | 25.1001 | 251.8977 | 188.9394 | 9.863 | 308.2192
2010-07-28 | 17:23:14 | 25.371 | 374.287 | 28.21 | 3290 | 16.5 |  | 122.2617 | 25.5527 | 256.4401 | 192.3464 | 10.0409 | 313.7772
2010-07-28 | 17:23:24 | 25.537 | 370.623 | 28.32 | 3282 | 16.5 |  | 121.0649 | 25.3026 | 253.93 | 190.4637 | 9.9426 | 310.7058
2010-07-28 | 17:23:34 | 25.704 | 369.301 | 28.36 | 3281 | 16.5 |  | 120.6332 | 25.2123 | 253.0244 | 189.7844 | 9.9071 | 309.5977
2010-07-28 | 17:23:44 | 25.871 | 372.283 | 28.27 | 3278 | 16.5 |  | 121.6072 | 25.4159 | 255.0673 | 191.3168 | 9.9871 | 312.0974
2010-07-28 | 17:23:54 | 26.038 | 366.348 | 28.45 | 3275 | 16.5 |  | 119.6684 | 25.0107 | 251.0007 | 188.2665 | 9.8279 | 307.1216
2010-07-28 | 17:24:04 | 26.205 | 373.283 | 28.24 | 3272 | 16.5 |  | 121.9339 | 25.4842 | 255.7526 | 191.8308 | 10.014 | 312.936
2010-07-28 | 17:24:14 | 26.372 | 373.952 | 28.22 | 3275 | 16.5 |  | 122.1523 | 25.5298 | 256.2107 | 192.1744 | 10.0319 | 313.4965
2010-07-28 | 17:24:24 | 26.539 | 370.954 | 28.31 | 3273 | 16.5 |  | 121.1731 | 25.3252 | 254.157 | 190.6339 | 9.9515 | 310.9835
2010-07-28 | 17:24:34 | 26.706 | 370.292 | 28.33 | 3267 | 16.5 |  | 120.9568 | 25.28 | 253.7032 | 190.2936 | 9.9337 | 310.4283
2010-07-28 | 17:24:44 | 26.873 | 374.287 | 28.21 | 3266 | 16.5 |  | 122.2617 | 25.5527 | 256.4401 | 192.3464 | 10.0409 | 313.7772
2010-07-28 | 17:24:54 | 27.04 | 375.293 | 28.18 | 3266 | 16.5 |  | 122.5905 | 25.6214 | 257.1298 | 192.8637 | 10.0679 | 314.621
2010-07-28 | 17:25:04 | 27.207 | 373.617 | 28.23 | 3268 | 16.5 |  | 122.043 | 25.507 | 255.9815 | 192.0025 | 10.0229 | 313.2161
2010-07-28 | 17:25:14 | 27.373 | 373.617 | 28.23 | 3259 | 16.5 |  | 122.043 | 25.507 | 255.9815 | 192.0025 | 10.0229 | 313.2161
2010-07-28 | 17:25:24 | 27.54 | 374.957 | 28.19 | 3257 | 16.5 |  | 122.4808 | 25.5985 | 256.8996 | 192.6911 | 10.0589 | 314.3394
2010-07-28 | 17:25:34 | 27.707 | 376.303 | 28.15 | 3254 | 16.5 |  | 122.9203 | 25.6903 | 257.8216 | 193.3827 | 10.095 | 315.4676
2010-07-28 | 17:25:44 | 27.874 | 373.617 | 28.23 | 3255 | 16.5 |  | 122.043 | 25.507 | 255.9815 | 192.0025 | 10.0229 | 313.2161
2010-07-28 | 17:25:54 | 28.041 | 379.692 | 28.05 | 3247 | 16.5 |  | 124.0275 | 25.9217 | 260.1438 | 195.1244 | 10.1859 | 318.3089
2010-07-28 | 17:26:04 | 28.208 | 373.952 | 28.22 | 3250 | 16.5 |  | 122.1523 | 25.5298 | 256.2107 | 192.1744 | 10.0319 | 313.4965
2010-07-28 | 17:26:14 | 28.375 | 376.64 | 28.14 | 3245 | 16.5 |  | 123.0305 | 25.7134 | 258.0527 | 193.556 | 10.104 | 315.7503
2010-07-28 | 17:26:24 | 28.542 | 379.012 | 28.07 | 3243 | 16.5 |  | 123.8051 | 25.8753 | 259.6774 | 194.7746 | 10.1676 | 317.7382
2010-07-28 | 17:26:35 | 28.709 | 376.64 | 28.14 | 3242 | 16.5 |  | 123.0305 | 25.7134 | 258.0527 | 193.556 | 10.104 | 315.7503
2010-07-28 | 17:26:45 | 28.876 | 377.654 | 28.11 | 3242 | 16.5 |  | 123.3618 | 25.7826 | 258.7475 | 194.0771 | 10.1312 | 316.6005
2010-07-28 | 17:26:55 | 29.043 | 376.64 | 28.14 | 3232 | 16.5 |  | 123.0305 | 25.7134 | 258.0527 | 193.556 | 10.104 | 315.7503
2010-07-28 | 17:27:05 | 29.21 | 382.43 | 27.97 | 3233 | 16.5 |  | 124.9217 | 26.1086 | 262.0193 | 196.5312 | 10.2593 | 320.6038
2010-07-28 | 17:27:15 | 29.377 | 379.352 | 28.06 | 3233 | 16.5 |  | 123.9162 | 25.8985 | 259.9104 | 194.9494 | 10.1767 | 318.0234
2010-07-28 | 17:27:25 | 29.543 | 377.654 | 28.11 | 3232 | 16.5 |  | 123.3618 | 25.7826 | 258.7475 | 194.0771 | 10.1312 | 316.6005
2010-07-28 | 17:27:35 | 29.71 | 384.844 | 27.9 | 3225 | 16.5 |  | 125.7104 | 26.2735 | 263.6736 | 197.772 | 10.3241 | 322.628
2010-07-28 | 17:27:45 | 29.877 | 382.086 | 27.98 | 3224 | 16.5 |  | 124.8095 | 26.0852 | 261.784 | 196.3547 | 10.2501 | 320.3159
2010-07-28 | 17:27:54 | 30.027 | 382.43 | 27.97 | 3221 | 16.5 |  | 124.9217 | 26.1086 | 262.0193 | 196.5312 | 10.2593 | 320.6038
2010-07-28 | 17:28:04 | 30.195 | 381.743 | 27.99 | 3223 | 16.5 |  | 124.6974 | 26.0618 | 261.549 | 196.1784 | 10.2409 | 320.0283
2010-07-28 | 17:28:14 | 30.361 | 385.885 | 27.87 | 3213 | 16.5 |  | 126.0502 | 26.3445 | 264.3864 | 198.3067 | 10.352 | 323.5001
2010-07-28 | 17:28:24 | 30.528 | 384.153 | 27.92 | 3215 | 16.5 |  | 125.4844 | 26.2262 | 263.1997 | 197.4166 | 10.3055 | 322.0481
2010-07-28 | 17:28:34 | 30.695 | 385.191 | 27.89 | 3210 | 16.5 |  | 125.8235 | 26.2971 | 263.911 | 197.95 | 10.3334 | 322.9184
2010-07-28 | 17:28:44 | 30.862 | 384.844 | 27.9 | 3208 | 16.5 |  | 125.7104 | 26.2735 | 263.6736 | 197.772 | 10.3241 | 322.628
2010-07-28 | 17:28:54 | 31.029 | 384.744 | 27.86 | 3207 | 16.6 |  | 125.9419 | 26.3219 | 264.1275 | 198.1125 | 10.3231 | 322.5975
2010-07-28 | 17:29:04 | 31.196 | 381.641 | 27.95 | 3204 | 16.6 |  | 124.9263 | 26.1096 | 261.9977 | 196.515 | 10.2399 | 319.9962
2010-07-28 | 17:29:14 | 31.363 | 385.437 | 27.84 | 3204 | 16.6 |  | 126.1689 | 26.3693 | 264.6036 | 198.4696 | 10.3417 | 323.179
2010-07-28 | 17:29:24 | 31.53 | 384.744 | 27.86 | 3202 | 16.6 |  | 125.9419 | 26.3219 | 264.1275 | 198.1125 | 10.3231 | 322.5975
2010-07-28 | 17:29:34 | 31.696 | 383.706 | 27.89 | 3203 | 16.6 |  | 125.6023 | 26.2509 | 263.4153 | 197.5783 | 10.2953 | 321.7276
2010-07-28 | 17:29:44 | 31.863 | 386.48 | 27.81 | 3193 | 16.6 |  | 126.5103 | 26.4407 | 265.3197 | 199.0067 | 10.3697 | 324.0536
2010-07-28 | 17:29:54 | 32.03 | 386.48 | 27.81 | 3195 | 16.6 |  | 126.5103 | 26.4407 | 265.3197 | 199.0067 | 10.3697 | 324.0536
2010-07-28 | 17:30:04 | 32.197 | 383.118 | 27.95 | 3197 | 16.5 |  | 125.1464 | 26.1556 | 262.4907 | 196.8848 | 10.2778 | 321.1806
2010-07-28 | 17:30:14 | 32.364 | 386.58 | 27.85 | 3193 | 16.5 |  | 126.2773 | 26.392 | 264.8628 | 198.664 | 10.3707 | 324.0831
2010-07-28 | 17:30:24 | 32.531 | 388.675 | 27.79 | 3188 | 16.5 |  | 126.9617 | 26.535 | 266.2983 | 199.7407 | 10.4269 | 325.8396
2010-07-28 | 17:30:34 | 32.698 | 390.784 | 27.73 | 3184 | 16.5 |  | 127.6505 | 26.679 | 267.7431 | 200.8244 | 10.4834 | 327.6073
2010-07-28 | 17:30:44 | 32.865 | 386.928 | 27.84 | 3191 | 16.5 |  | 126.3911 | 26.4157 | 265.1015 | 198.843 | 10.38 | 324.3751
2010-07-28 | 17:30:54 | 33.031 | 390.079 | 27.75 | 3180 | 16.5 |  | 127.4204 | 26.6309 | 267.2605 | 200.4624 | 10.4645 | 327.0168
2010-07-28 | 17:31:04 | 33.198 | 392.906 | 27.67 | 3186 | 16.5 |  | 128.3438 | 26.8239 | 269.1971 | 201.915 | 10.5404 | 329.3865
2010-07-28 | 17:31:14 | 33.365 | 388.325 | 27.8 | 3178 | 16.5 |  | 126.8474 | 26.5111 | 266.0584 | 199.5608 | 10.4175 | 325.546
2010-07-28 | 17:31:24 | 33.533 | 390.431 | 27.74 | 3183 | 16.5 |  | 127.5354 | 26.6549 | 267.5016 | 200.6433 | 10.474 | 327.3119
2010-07-28 | 17:31:34 | 33.699 | 391.843 | 27.7 | 3174 | 16.5 |  | 127.9966 | 26.7513 | 268.4689 | 201.3688 | 10.5119 | 328.4955
2010-07-28 | 17:31:44 | 33.866 | 392.552 | 27.68 | 3175 | 16.5 |  | 128.2279 | 26.7996 | 268.9541 | 201.7328 | 10.5309 | 329.0892
2010-07-28 | 17:31:54 | 34.033 | 397.649 | 27.58 | 3175 | 16.4 |  | 129.62 | 27.0906 | 271.9063 | 203.9471 | 10.6658 | 333.3068
2010-07-28 | 17:32:04 | 34.2 | 398.73 | 27.55 | 3174 | 16.4 |  | 129.9723 | 27.1642 | 272.6453 | 204.5014 | 10.6948 | 334.2127
2010-07-28 | 17:32:14 | 34.367 | 395.141 | 27.65 | 3170 | 16.4 |  | 128.8024 | 26.9197 | 270.1912 | 202.6606 | 10.5985 | 331.2044
2010-07-28 | 17:32:24 | 34.534 | 392.906 | 27.67 | 3165 | 16.5 |  | 128.3438 | 26.8239 | 269.1971 | 201.915 | 10.5404 | 329.3865
2010-07-28 | 17:32:34 | 34.701 | 397.29 | 27.59 | 3166 | 16.4 |  | 129.5028 | 27.0661 | 271.6605 | 203.7627 | 10.6562 | 333.0055
2010-07-28 | 17:32:44 | 34.867 | 395.4 | 27.6 | 3163 | 16.5 |  | 129.1582 | 26.9941 | 270.9054 | 203.1964 | 10.6073 | 331.4768
2010-07-28 | 17:32:54 | 35.034 | 395.757 | 27.59 | 3161 | 16.5 |  | 129.2751 | 27.0185 | 271.1505 | 203.3802 | 10.6169 | 331.7767
2010-07-28 | 17:33:04 | 35.201 | 398.009 | 27.57 | 3167 | 16.4 |  | 129.7373 | 27.1151 | 272.1524 | 204.1317 | 10.6755 | 333.6085
2010-07-28 | 17:33:14 | 35.368 | 398.369 | 27.56 | 3152 | 16.4 |  | 129.8547 | 27.1396 | 272.3987 | 204.3164 | 10.6851 | 333.9104
2010-07-28 | 17:33:24 | 35.535 | 400.806 | 27.45 | 3160 | 16.5 |  | 130.9244 | 27.3632 | 274.6098 | 205.9749 | 10.7523 | 336.0094
2010-07-28 | 17:33:34 | 35.702 | 398.272 | 27.52 | 3156 | 16.5 |  | 130.0966 | 27.1902 | 272.8736 | 204.6726 | 10.6843 | 333.885
2010-07-28 | 17:33:44 | 35.869 | 402.994 | 27.39 | 3153 | 16.5 |  | 131.6389 | 27.5125 | 276.1084 | 207.0989 | 10.811 | 337.8431
2010-07-28 | 17:33:54 | 36.036 | 398.994 | 27.5 | 3147 | 16.5 |  | 130.3325 | 27.2395 | 273.3683 | 205.0437 | 10.7037 | 334.4904
2010-07-28 | 17:34:04 | 36.203 | 398.272 | 27.52 | 3155 | 16.5 |  | 130.0966 | 27.1902 | 272.8736 | 204.6726 | 10.6843 | 333.885
2010-07-28 | 17:34:14 | 36.37 | 397.912 | 27.53 | 3148 | 16.5 |  | 129.9789 | 27.1656 | 272.6267 | 204.4874 | 10.6747 | 333.5829