## convers.+compens.
file name: PSt1_Eq4_Sal |  |  | 2008-03-29
author: Dr. Christian Huber
programm for temperature compensated oxygen calculation of membrane PSt1
Fit of the calibration curve is optimized between 0 and 250 % air-saturation (0 and 388 Torr)
enter your values here
input of F0,T0 | 58.62 | phase angle of cal 0
input of F100,T100 | 31.05 | phase angle of cal 100 (water-vapor saturated air)
input of Fm,Tm | 30.5 | measured phase angle
input of T0 | 23.3 | temperature of cal 0
input of T100 | 17.9 | temperatur of cal100
input of Tm | 17.9 | temperature at measurement
air pressure patm | 1013 | air pressure
Salinity | 32.3 | Salinity
Results of calculation + compensation
[%] air saturation | 104.8779 | %
[%] O2 | 21.9195 | %
pO2 [hPa] | 218.0572 | hPa
pO2 [Torr] | 163.5568 | Torr
cO2 [mg/L] | 8.1931 | mg/L
cO2 [ppm] | 8.1931 | ppm
cO2 [µmol/L] | 256.0334 | µmol/L
Internal calculated parameter (DO NOT CHANGE !!)
f1 | 0.801
DF/K | -0.08
DKSV/K | 0.000383
m | 22.9
tan(F0,T100) | 1.6677
tan(F0,Tm) | 1.6677
tan(F100,T100) | 0.60205
tan(Fm,Tm) | 0.58905
KSV,T100 | 0.032842
KSV,Tm | 0.032842
A | 157.6438
B | 14.279
C | -0.639
a | 1.6637e-05
b | 0.0044223
c | -0.64679
Chlorinity | 17.8781

## Recalculate Fibox values
file name: PSt1_Eq4_Sal |  |  |  | 2008-03-29
author: Dr. Christian Huber
programm for temperature compensated oxygen calculation of membrane PSt1
Fit of the calibration curve is optimized between 0 and 250 % air-saturation (0 and 388 Torr)
Insert your calibration values
cal0 | 58.62 | ° | T0 | 23.3 | °C
cal100 | 31.05 | ° | T100 | 17.9 | °C
air pressure | 1020 | mbar
Salinity | 32.2 | °/°°
Internal calculated parameter (DO NOT CHANGE !!)
f1 | 0.801 | tan(F0,T100) | 1.6677 | A | 157.6438 | KSV,T100 | 0.032842 | a | 1.669e-05
DF/K | -0.08 | tan(F0,Tm) | 1.6677 | B | 14.279 | KSV,Tm | 0.032842 | b | 0.004461 |  |  |  |  | INSERT YOUR RESPECTIVE DATA
DKSV/K | 0.000383 | tan(F100,T100) | 0.60205 | C | -0.639 |  |  | c | -0.64566 |  |  |  |  | chamber volume [L] | Fresh weight (g) | Blank (Chamber 2) cO2 [µmol/ h]
m | 22.9 | tan(Fm,Tm) | 0.59093 |  |  |  |  | Chlorinity | 17.8227 |  |  |  |  | 0.026 | 0.0236 | -0.13198
Insert the data you get from the Ascii File recorded with the OxyView TX3-V5.31.exe software | [7 empty] | The oxygen contents are calculated automatically from the raw data
date | time/hh:mm:ss | logtime/min | oxygen/% airsatur. | phase/° | amp | temp/°C |  | % air-sat. | % oxygen | pO2 (hPa) | pO2 (Torr) | cO2 [mg/L] | cO2 [µmol/L] | cO2 [µmol/L 10 min] | cO2 [µmol/L h] | cO2 [µmol/ h] | cO2 [µmol/ h gFW]
2010-08-23 | 12:25:26 | 0 | 308.47 | 30.58 | 5913 | 17.9 |  | 104.1519 | 21.7678 | 218.0753 | 163.5704 | 8.1991 | 256.221 | 12.12 | 72.72 | 2.0227 | 85.7075
2010-08-23 | 12:25:37 | 0.184 | 305.411 | 30.74 | 5905 | 17.8 |  | 102.908 | 21.5078 | 215.4986 | 161.6377 | 8.1165 | 253.6393
2010-08-23 | 12:25:47 | 0.35 | 307.269 | 30.67 | 5905 | 17.8 |  | 103.5341 | 21.6386 | 216.8097 | 162.6211 | 8.1658 | 255.1824 | change data input according to column N
2010-08-23 | 12:25:57 | 0.517 | 308.872 | 30.61 | 5920 | 17.8 |  | 104.0742 | 21.7515 | 217.9407 | 163.4694 | 8.2084 | 256.5136
2010-08-23 | 12:26:07 | 0.684 | 306.737 | 30.69 | 5901 | 17.8 |  | 103.3548 | 21.6012 | 216.4342 | 162.3394 | 8.1517 | 254.7405
2010-08-23 | 12:26:17 | 0.851 | 307.803 | 30.65 | 5917 | 17.8 |  | 103.7138 | 21.6762 | 217.186 | 162.9033 | 8.18 | 255.6253
2010-08-23 | 12:26:27 | 1.018 | 307.003 | 30.68 | 5906 | 17.8 |  | 103.4444 | 21.6199 | 216.6218 | 162.4802 | 8.1588 | 254.9613
2010-08-23 | 12:26:37 | 1.185 | 308.337 | 30.63 | 5898 | 17.8 |  | 103.8938 | 21.7138 | 217.563 | 163.1861 | 8.1942 | 256.069
2010-08-23 | 12:26:47 | 1.352 | 307.269 | 30.67 | 5905 | 17.8 |  | 103.5341 | 21.6386 | 216.8097 | 162.6211 | 8.1658 | 255.1824
2010-08-23 | 12:26:58 | 1.519 | 307.269 | 30.67 | 5900 | 17.8 |  | 103.5341 | 21.6386 | 216.8097 | 162.6211 | 8.1658 | 255.1824
2010-08-23 | 12:27:08 | 1.686 | 308.07 | 30.64 | 5896 | 17.8 |  | 103.8038 | 21.695 | 217.3744 | 163.0446 | 8.1871 | 255.847
2010-08-23 | 12:27:18 | 1.853 | 308.07 | 30.64 | 5891 | 17.8 |  | 103.8038 | 21.695 | 217.3744 | 163.0446 | 8.1871 | 255.847
2010-08-23 | 12:27:28 | 2.02 | 309.14 | 30.6 | 5893 | 17.8 |  | 104.1646 | 21.7704 | 218.1298 | 163.6113 | 8.2156 | 256.7362
2010-08-23 | 12:27:38 | 2.186 | 308.337 | 30.63 | 5882 | 17.8 |  | 103.8938 | 21.7138 | 217.563 | 163.1861 | 8.1942 | 256.069
2010-08-23 | 12:27:48 | 2.353 | 311.024 | 30.53 | 5892 | 17.8 |  | 104.7993 | 21.9031 | 219.4591 | 164.6083 | 8.2656 | 258.3007
2010-08-23 | 12:27:58 | 2.52 | 307.003 | 30.68 | 5876 | 17.8 |  | 103.4444 | 21.6199 | 216.6218 | 162.4802 | 8.1588 | 254.9613
2010-08-23 | 12:28:08 | 2.687 | 310.754 | 30.54 | 5884 | 17.8 |  | 104.7084 | 21.884 | 219.2686 | 164.4654 | 8.2584 | 258.0766
2010-08-23 | 12:28:18 | 2.854 | 312.378 | 30.48 | 5883 | 17.8 |  | 105.2554 | 21.9984 | 220.4141 | 165.3247 | 8.3016 | 259.4248
2010-08-23 | 12:28:28 | 3.021 | 308.872 | 30.61 | 5869 | 17.8 |  | 104.0742 | 21.7515 | 217.9407 | 163.4694 | 8.2084 | 256.5136
2010-08-23 | 12:28:38 | 3.188 | 309.14 | 30.6 | 5883 | 17.8 |  | 104.1646 | 21.7704 | 218.1298 | 163.6113 | 8.2156 | 256.7362
2010-08-23 | 12:28:48 | 3.355 | 309.946 | 30.57 | 5878 | 17.8 |  | 104.4361 | 21.8271 | 218.6984 | 164.0377 | 8.237 | 257.4054
2010-08-23 | 12:28:58 | 3.522 | 311.16 | 30.48 | 5875 | 17.9 |  | 105.0602 | 21.9576 | 219.9771 | 164.9968 | 8.2706 | 258.4554
2010-08-23 | 12:29:08 | 3.689 | 308.202 | 30.59 | 5860 | 17.9 |  | 104.0616 | 21.7489 | 217.8861 | 163.4285 | 8.192 | 255.9987 |  | regression formula | time
2010-08-23 | 12:29:18 | 3.856 | 306.868 | 30.64 | 5863 | 17.9 |  | 103.6112 | 21.6547 | 216.9431 | 162.7212 | 8.1565 | 254.8908 | value for T=26 min. | 266.472 | T11
2010-08-23 | 12:29:28 | 4.022 | 308.738 | 30.57 | 5873 | 17.9 |  | 104.2424 | 21.7867 | 218.2646 | 163.7124 | 8.2062 | 256.4434 | value for T=1 min. | 254.352 | T1
2010-08-23 | 12:29:38 | 4.189 | 308.202 | 30.59 | 5854 | 17.9 |  | 104.0616 | 21.7489 | 217.8861 | 163.4285 | 8.192 | 255.9987 | difference between T25 and T1 | 12.12 | 10minutes
2010-08-23 | 12:29:48 | 4.356 | 308.47 | 30.58 | 5858 | 17.9 |  | 104.1519 | 21.7678 | 218.0753 | 163.5704 | 8.1991 | 256.221 | calculate from regression curve values for 10 minutes of photosynthesis or respiration
2010-08-23 | 12:29:58 | 4.523 | 310.35 | 30.51 | 5864 | 17.9 |  | 104.7868 | 21.9004 | 219.4046 | 164.5674 | 8.249 | 257.7828
2010-08-23 | 12:30:08 | 4.69 | 307.935 | 30.6 | 5851 | 17.9 |  | 103.9714 | 21.73 | 217.6972 | 163.2868 | 8.1849 | 255.7767
2010-08-23 | 12:30:18 | 4.857 | 311.431 | 30.47 | 5851 | 17.9 |  | 105.1516 | 21.9767 | 220.1683 | 165.1403 | 8.2778 | 258.6801
2010-08-23 | 12:30:28 | 5.024 | 313.331 | 30.4 | 5850 | 17.9 |  | 105.7933 | 22.1108 | 221.5121 | 166.1482 | 8.3283 | 260.2589
2010-08-23 | 12:30:38 | 5.191 | 311.16 | 30.48 | 5841 | 17.9 |  | 105.0602 | 21.9576 | 219.9771 | 164.9968 | 8.2706 | 258.4554
2010-08-23 | 12:30:48 | 5.358 | 312.515 | 30.43 | 5842 | 17.9 |  | 105.5178 | 22.0532 | 220.935 | 165.7154 | 8.3066 | 259.581
2010-08-23 | 12:30:58 | 5.525 | 313.059 | 30.41 | 5840 | 17.9 |  | 105.7014 | 22.0916 | 221.3195 | 166.0038 | 8.321 | 260.0327
2010-08-23 | 12:31:08 | 5.692 | 314.15 | 30.37 | 5836 | 17.9 |  | 106.0698 | 22.1686 | 222.0908 | 166.5823 | 8.35 | 260.9389
2010-08-23 | 12:31:18 | 5.858 | 315.245 | 30.33 | 5840 | 17.9 |  | 106.4396 | 22.2459 | 222.8652 | 167.1631 | 8.3792 | 261.8487
2010-08-23 | 12:31:28 | 6.026 | 312.515 | 30.43 | 5836 | 17.9 |  | 105.5178 | 22.0532 | 220.935 | 165.7154 | 8.3066 | 259.581
2010-08-23 | 12:31:38 | 6.192 | 314.15 | 30.37 | 5836 | 17.9 |  | 106.0698 | 22.1686 | 222.0908 | 166.5823 | 8.35 | 260.9389
2010-08-23 | 12:31:48 | 6.359 | 316.07 | 30.3 | 5829 | 17.9 |  | 106.7179 | 22.304 | 223.448 | 167.6002 | 8.4011 | 262.5334
2010-08-23 | 12:31:58 | 6.526 | 313.877 | 30.38 | 5823 | 17.9 |  | 105.9775 | 22.1493 | 221.8977 | 166.4374 | 8.3428 | 260.712
2010-08-23 | 12:32:08 | 6.693 | 313.604 | 30.39 | 5830 | 17.9 |  | 105.8854 | 22.13 | 221.7048 | 166.2927 | 8.3355 | 260.4854
2010-08-23 | 12:32:18 | 6.86 | 314.423 | 30.36 | 5821 | 17.9 |  | 106.1621 | 22.1879 | 222.2841 | 166.7273 | 8.3573 | 261.166
2010-08-23 | 12:32:28 | 7.027 | 315.795 | 30.31 | 5817 | 17.9 |  | 106.625 | 22.2846 | 223.2535 | 167.4544 | 8.3938 | 262.305
2010-08-23 | 12:32:38 | 7.194 | 316.07 | 30.3 | 5820 | 17.9 |  | 106.7179 | 22.304 | 223.448 | 167.6002 | 8.4011 | 262.5334
2010-08-23 | 12:32:48 | 7.361 | 315.52 | 30.32 | 5823 | 17.9 |  | 106.5323 | 22.2652 | 223.0592 | 167.3087 | 8.3865 | 262.0767
2010-08-23 | 12:32:58 | 7.528 | 318.969 | 30.24 | 5813 | 17.8 |  | 107.4762 | 22.4625 | 225.0647 | 168.8129 | 8.4768 | 264.8985
2010-08-23 | 12:33:08 | 7.695 | 317.306 | 30.3 | 5820 | 17.8 |  | 106.916 | 22.3455 | 223.8917 | 167.933 | 8.4326 | 263.5179
2010-08-23 | 12:33:18 | 7.861 | 317.306 | 30.3 | 5808 | 17.8 |  | 106.916 | 22.3455 | 223.8917 | 167.933 | 8.4326 | 263.5179
2010-08-23 | 12:33:28 | 8.028 | 317.306 | 30.3 | 5810 | 17.8 |  | 106.916 | 22.3455 | 223.8917 | 167.933 | 8.4326 | 263.5179
2010-08-23 | 12:33:38 | 8.195 | 316.754 | 30.32 | 5797 | 17.8 |  | 106.7301 | 22.3066 | 223.5022 | 167.6409 | 8.4179 | 263.0595
2010-08-23 | 12:33:48 | 8.362 | 317.859 | 30.28 | 5799 | 17.8 |  | 107.1024 | 22.3844 | 224.2819 | 168.2257 | 8.4473 | 263.9772
2010-08-23 | 12:33:58 | 8.529 | 317.583 | 30.29 | 5802 | 17.8 |  | 107.0092 | 22.3649 | 224.0867 | 168.0793 | 8.4399 | 263.7474
2010-08-23 | 12:34:08 | 8.696 | 314.557 | 30.4 | 5796 | 17.8 |  | 105.9898 | 22.1519 | 221.9521 | 166.4782 | 8.3595 | 261.235
2010-08-23 | 12:34:17 | 8.846 | 317.859 | 30.28 | 5793 | 17.8 |  | 107.1024 | 22.3844 | 224.2819 | 168.2257 | 8.4473 | 263.9772
2010-08-23 | 12:34:27 | 9.013 | 319.247 | 30.23 | 5793 | 17.8 |  | 107.5699 | 22.4821 | 225.2609 | 168.96 | 8.4841 | 265.1294
2010-08-23 | 12:34:37 | 9.18 | 319.247 | 30.23 | 5800 | 17.8 |  | 107.5699 | 22.4821 | 225.2609 | 168.96 | 8.4841 | 265.1294
2010-08-23 | 12:34:47 | 9.347 | 322.042 | 30.13 | 5796 | 17.8 |  | 108.5119 | 22.679 | 227.2335 | 170.4396 | 8.5584 | 267.4511
2010-08-23 | 12:34:57 | 9.514 | 321.201 | 30.16 | 5802 | 17.8 |  | 108.2283 | 22.6197 | 226.6397 | 169.9942 | 8.5361 | 266.7522
2010-08-23 | 12:35:07 | 9.681 | 315.654 | 30.36 | 5787 | 17.8 |  | 106.3592 | 22.2291 | 222.7256 | 167.0584 | 8.3887 | 262.1454
2010-08-23 | 12:35:17 | 9.848 | 320.083 | 30.2 | 5789 | 17.8 |  | 107.8515 | 22.541 | 225.8506 | 169.4024 | 8.5064 | 265.8235
2010-08-23 | 12:35:27 | 10.014 | 320.362 | 30.19 | 5783 | 17.8 |  | 107.9456 | 22.5606 | 226.0476 | 169.5501 | 8.5138 | 266.0553
2010-08-23 | 12:35:37 | 10.181 | 317.583 | 30.29 | 5779 | 17.8 |  | 107.0092 | 22.3649 | 224.0867 | 168.0793 | 8.4399 | 263.7474
2010-08-23 | 12:35:47 | 10.348 | 320.362 | 30.19 | 5780 | 17.8 |  | 107.9456 | 22.5606 | 226.0476 | 169.5501 | 8.5138 | 266.0553
2010-08-23 | 12:35:57 | 10.515 | 319.247 | 30.23 | 5775 | 17.8 |  | 107.5699 | 22.4821 | 225.2609 | 168.96 | 8.4841 | 265.1294
2010-08-23 | 12:36:07 | 10.682 | 323.451 | 30.08 | 5763 | 17.8 |  | 108.9864 | 22.7782 | 228.2272 | 171.1849 | 8.5959 | 268.6207
2010-08-23 | 12:36:17 | 10.849 | 322.042 | 30.13 | 5769 | 17.8 |  | 108.5119 | 22.679 | 227.2335 | 170.4396 | 8.5584 | 267.4511
2010-08-23 | 12:36:27 | 11.016 | 321.762 | 30.14 | 5766 | 17.8 |  | 108.4172 | 22.6592 | 227.0354 | 170.291 | 8.551 | 267.2179
2010-08-23 | 12:36:37 | 11.183 | 320.083 | 30.2 | 5761 | 17.8 |  | 107.8515 | 22.541 | 225.8506 | 169.4024 | 8.5064 | 265.8235
2010-08-23 | 12:36:47 | 11.35 | 319.67 | 30.17 | 5773 | 17.9 |  | 107.9336 | 22.5581 | 225.9935 | 169.5095 | 8.4968 | 265.5242
2010-08-23 | 12:36:57 | 11.517 | 323.318 | 30.04 | 5758 | 17.9 |  | 109.1652 | 22.8155 | 228.5721 | 171.4437 | 8.5937 | 268.5539
2010-08-23 | 12:37:07 | 11.683 | 319.392 | 30.18 | 5755 | 17.9 |  | 107.8396 | 22.5385 | 225.7965 | 169.3618 | 8.4894 | 265.2928
2010-08-23 | 12:37:17 | 11.85 | 320.228 | 30.15 | 5756 | 17.9 |  | 108.1221 | 22.5975 | 226.388 | 169.8054 | 8.5116 | 265.9878
2010-08-23 | 12:37:27 | 12.017 | 321.628 | 30.1 | 5757 | 17.9 |  | 108.5948 | 22.6963 | 227.3778 | 170.5479 | 8.5488 | 267.1507
2010-08-23 | 12:37:37 | 12.184 | 323.6 | 30.03 | 5738 | 17.9 |  | 109.2606 | 22.8355 | 228.7719 | 171.5935 | 8.6012 | 268.7886
2010-08-23 | 12:37:47 | 12.351 | 321.348 | 30.11 | 5744 | 17.9 |  | 108.5001 | 22.6765 | 227.1795 | 170.3991 | 8.5414 | 266.9177
2010-08-23 | 12:37:57 | 12.518 | 321.348 | 30.11 | 5744 | 17.9 |  | 108.5001 | 22.6765 | 227.1795 | 170.3991 | 8.5414 | 266.9177
2010-08-23 | 12:38:07 | 12.685 | 320.788 | 30.13 | 5738 | 17.9 |  | 108.3109 | 22.637 | 226.7833 | 170.102 | 8.5265 | 266.4522
2010-08-23 | 12:38:17 | 12.852 | 325.017 | 29.98 | 5737 | 17.9 |  | 109.739 | 22.9355 | 229.7736 | 172.3449 | 8.6389 | 269.9656
2010-08-23 | 12:38:28 | 13.019 | 325.017 | 29.98 | 5743 | 17.9 |  | 109.739 | 22.9355 | 229.7736 | 172.3449 | 8.6389 | 269.9656
2010-08-23 | 12:38:38 | 13.186 | 323.036 | 30.05 | 5733 | 17.9 |  | 109.0699 | 22.7956 | 228.3726 | 171.294 | 8.5862 | 268.3195
2010-08-23 | 12:38:48 | 13.353 | 322.754 | 30.06 | 5726 | 17.9 |  | 108.9747 | 22.7757 | 228.1732 | 171.1445 | 8.5787 | 268.0853
2010-08-23 | 12:38:58 | 13.519 | 325.871 | 29.95 | 5727 | 17.9 |  | 110.0272 | 22.9957 | 230.3771 | 172.7975 | 8.6616 | 270.6746
2010-08-23 | 12:39:08 | 13.687 | 325.586 | 29.96 | 5729 | 17.9 |  | 109.9311 | 22.9756 | 230.1757 | 172.6465 | 8.654 | 270.438
2010-08-23 | 12:39:18 | 13.853 | 323.6 | 30.03 | 5727 | 17.9 |  | 109.2606 | 22.8355 | 228.7719 | 171.5935 | 8.6012 | 268.7886
2010-08-23 | 12:39:28 | 14.02 | 326.727 | 29.92 | 5721 | 17.9 |  | 110.3163 | 23.0561 | 230.9824 | 173.2515 | 8.6843 | 271.3857
2010-08-23 | 12:39:38 | 14.187 | 325.586 | 29.96 | 5709 | 17.9 |  | 109.9311 | 22.9756 | 230.1757 | 172.6465 | 8.654 | 270.438
2010-08-23 | 12:39:48 | 14.355 | 328.16 | 29.87 | 5717 | 17.9 |  | 110.8 | 23.1572 | 231.9952 | 174.0112 | 8.7224 | 272.5758
2010-08-23 | 12:39:58 | 14.521 | 325.586 | 29.96 | 5718 | 17.9 |  | 109.9311 | 22.9756 | 230.1757 | 172.6465 | 8.654 | 270.438
2010-08-23 | 12:40:08 | 14.688 | 325.586 | 29.96 | 5710 | 17.9 |  | 109.9311 | 22.9756 | 230.1757 | 172.6465 | 8.654 | 270.438
2010-08-23 | 12:40:18 | 14.856 | 325.302 | 29.97 | 5705 | 17.9 |  | 109.835 | 22.9555 | 229.9746 | 172.4956 | 8.6465 | 270.2017
2010-08-23 | 12:40:28 | 15.022 | 323.883 | 30.02 | 5699 | 17.9 |  | 109.3561 | 22.8554 | 228.9718 | 171.7435 | 8.6088 | 269.0235
2010-08-23 | 12:40:38 | 15.189 | 327.432 | 29.94 | 5711 | 17.8 |  | 110.3277 | 23.0585 | 231.0361 | 173.2918 | 8.7017 | 271.9267
2010-08-23 | 12:40:48 | 15.356 | 328.292 | 29.91 | 5705 | 17.8 |  | 110.6176 | 23.1191 | 231.6431 | 173.7471 | 8.7245 | 272.6412
2010-08-23 | 12:40:58 | 15.523 | 327.145 | 29.95 | 5705 | 17.8 |  | 110.2313 | 23.0383 | 230.8342 | 173.1403 | 8.694 | 271.6891
2010-08-23 | 12:41:08 | 15.69 | 330.889 | 29.82 | 5698 | 17.8 |  | 111.4925 | 23.3019 | 233.4753 | 175.1214 | 8.7935 | 274.7977
2010-08-23 | 12:41:18 | 15.857 | 326.574 | 29.97 | 5697 | 17.8 |  | 110.0387 | 22.9981 | 230.4309 | 172.8378 | 8.6789 | 271.2144
2010-08-23 | 12:41:28 | 16.024 | 330.599 | 29.83 | 5692 | 17.8 |  | 111.3949 | 23.2815 | 233.2709 | 174.9681 | 8.7858 | 274.5571
2010-08-23 | 12:41:38 | 16.191 | 329.155 | 29.88 | 5692 | 17.8 |  | 110.9084 | 23.1799 | 232.252 | 174.2038 | 8.7475 | 273.3579
2010-08-23 | 12:41:48 | 16.358 | 331.469 | 29.8 | 5689 | 17.8 |  | 111.6881 | 23.3428 | 233.8847 | 175.4284 | 8.8089 | 275.2795
2010-08-23 | 12:41:58 | 16.525 | 334.681 | 29.69 | 5691 | 17.8 |  | 112.7704 | 23.569 | 236.1513 | 177.1285 | 8.8943 | 277.9472
2010-08-23 | 12:42:08 | 16.692 | 329.155 | 29.88 | 5692 | 17.8 |  | 110.9084 | 23.1799 | 232.252 | 174.2038 | 8.7475 | 273.3579
2010-08-23 | 12:42:18 | 16.859 | 330.02 | 29.85 | 5680 | 17.8 |  | 111.2 | 23.2408 | 232.8627 | 174.6619 | 8.7705 | 274.0767
2010-08-23 | 12:42:28 | 17.025 | 331.759 | 29.79 | 5686 | 17.8 |  | 111.786 | 23.3633 | 234.0897 | 175.5822 | 8.8167 | 275.5208
2010-08-23 | 12:42:38 | 17.192 | 330.889 | 29.82 | 5685 | 17.8 |  | 111.4925 | 23.3019 | 233.4753 | 175.1214 | 8.7935 | 274.7977
2010-08-23 | 12:42:48 | 17.359 | 331.469 | 29.8 | 5671 | 17.8 |  | 111.6881 | 23.3428 | 233.8847 | 175.4284 | 8.8089 | 275.2795
2010-08-23 | 12:42:58 | 17.526 | 328.579 | 29.9 | 5668 | 17.8 |  | 110.7145 | 23.1393 | 231.8459 | 173.8992 | 8.7322 | 272.8799
2010-08-23 | 12:43:08 | 17.693 | 328.292 | 29.91 | 5662 | 17.8 |  | 110.6176 | 23.1191 | 231.6431 | 173.7471 | 8.7245 | 272.6412
2010-08-23 | 12:43:18 | 17.86 | 330.309 | 29.84 | 5674 | 17.8 |  | 111.2974 | 23.2612 | 233.0667 | 174.8149 | 8.7781 | 274.3168
2010-08-23 | 12:43:28 | 18.027 | 330.599 | 29.83 | 5674 | 17.8 |  | 111.3949 | 23.2815 | 233.2709 | 174.9681 | 8.7858 | 274.5571
2010-08-23 | 12:43:38 | 18.194 | 333.216 | 29.74 | 5664 | 17.8 |  | 112.2769 | 23.4659 | 235.1179 | 176.3534 | 8.8554 | 276.731
2010-08-23 | 12:43:48 | 18.361 | 329.443 | 29.87 | 5659 | 17.8 |  | 111.0055 | 23.2002 | 232.4554 | 174.3564 | 8.7551 | 273.5972
2010-08-23 | 12:43:58 | 18.528 | 332.924 | 29.75 | 5658 | 17.8 |  | 112.1785 | 23.4453 | 234.9118 | 176.1989 | 8.8476 | 276.4884
2010-08-23 | 12:44:08 | 18.695 | 333.509 | 29.73 | 5654 | 17.8 |  | 112.3754 | 23.4865 | 235.3241 | 176.5081 | 8.8632 | 276.9737
2010-08-23 | 12:44:18 | 18.861 | 331.919 | 29.74 | 5658 | 17.9 |  | 112.0692 | 23.4225 | 234.6526 | 176.0044 | 8.8223 | 275.698
2010-08-23 | 12:44:28 | 19.028 | 334.256 | 29.66 | 5654 | 17.9 |  | 112.8585 | 23.5874 | 236.3053 | 177.244 | 8.8845 | 277.6398
2010-08-23 | 12:44:38 | 19.195 | 329.889 | 29.81 | 5652 | 17.9 |  | 111.3837 | 23.2792 | 233.2174 | 174.9279 | 8.7684 | 274.0117
2010-08-23 | 12:44:48 | 19.362 | 335.727 | 29.61 | 5647 | 17.9 |  | 113.3551 | 23.6912 | 237.345 | 178.0239 | 8.9236 | 278.8614
2010-08-23 | 12:44:58 | 19.529 | 334.55 | 29.65 | 5652 | 17.9 |  | 112.9576 | 23.6081 | 236.5128 | 177.3997 | 8.8923 | 277.8836
2010-08-23 | 12:45:08 | 19.696 | 334.256 | 29.66 | 5646 | 17.9 |  | 112.8585 | 23.5874 | 236.3053 | 177.244 | 8.8845 | 277.6398
2010-08-23 | 12:45:18 | 19.863 | 336.613 | 29.58 | 5642 | 17.9 |  | 113.6543 | 23.7537 | 237.9714 | 178.4937 | 8.9471 | 279.5974
2010-08-23 | 12:45:28 | 20.03 | 334.55 | 29.65 | 5636 | 17.9 |  | 112.9576 | 23.6081 | 236.5128 | 177.3997 | 8.8923 | 277.8836
2010-08-23 | 12:45:38 | 20.197 | 332.21 | 29.73 | 5630 | 17.9 |  | 112.1675 | 23.443 | 234.8584 | 176.1588 | 8.8301 | 275.9398
2010-08-23 | 12:45:47 | 20.347 | 333.67 | 29.68 | 5617 | 17.9 |  | 112.6606 | 23.5461 | 235.8909 | 176.9332 | 8.8689 | 277.1528
2010-08-23 | 12:45:57 | 20.514 | 335.432 | 29.62 | 5633 | 17.9 |  | 113.2556 | 23.6704 | 237.1367 | 177.8676 | 8.9157 | 278.6166
2010-08-23 | 12:46:07 | 20.681 | 336.022 | 29.6 | 5628 | 17.9 |  | 113.4547 | 23.712 | 237.5536 | 178.1804 | 8.9314 | 279.1065
2010-08-23 | 12:46:17 | 20.848 | 336.022 | 29.6 | 5619 | 17.9 |  | 113.4547 | 23.712 | 237.5536 | 178.1804 | 8.9314 | 279.1065
2010-08-23 | 12:46:27 | 21.014 | 336.022 | 29.6 | 5623 | 17.9 |  | 113.4547 | 23.712 | 237.5536 | 178.1804 | 8.9314 | 279.1065
2010-08-23 | 12:46:37 | 21.181 | 335.01 | 29.59 | 5624 | 18 |  | 113.3446 | 23.689 | 237.2922 | 177.9843 | 8.9059 | 278.3102
2010-08-23 | 12:46:47 | 21.348 | 335.304 | 29.58 | 5612 | 18 |  | 113.4442 | 23.7098 | 237.5008 | 178.1408 | 8.9138 | 278.5548
2010-08-23 | 12:46:57 | 21.515 | 334.128 | 29.62 | 5609 | 18 |  | 113.0463 | 23.6267 | 236.6677 | 177.5158 | 8.8825 | 277.5776
2010-08-23 | 12:47:07 | 21.682 | 337.968 | 29.49 | 5601 | 18 |  | 114.3456 | 23.8982 | 239.3878 | 179.5561 | 8.9846 | 280.768
2010-08-23 | 12:47:17 | 21.849 | 334.716 | 29.6 | 5610 | 18 |  | 113.245 | 23.6682 | 237.0838 | 177.828 | 8.8981 | 278.0657
2010-08-23 | 12:47:27 | 22.016 | 334.716 | 29.6 | 5605 | 18 |  | 113.245 | 23.6682 | 237.0838 | 177.828 | 8.8981 | 278.0657
2010-08-23 | 12:47:37 | 22.183 | 337.078 | 29.52 | 5595 | 18 |  | 114.0442 | 23.8352 | 238.7569 | 179.0829 | 8.9609 | 280.028
2010-08-23 | 12:47:47 | 22.35 | 342.59 | 29.38 | 5602 | 17.9 |  | 115.6721 | 24.1755 | 242.1965 | 181.6628 | 9.106 | 284.5615
2010-08-23 | 12:47:57 | 22.517 | 334.55 | 29.65 | 5595 | 17.9 |  | 112.9576 | 23.6081 | 236.5128 | 177.3997 | 8.8923 | 277.8836
2010-08-23 | 12:48:07 | 22.683 | 340.784 | 29.44 | 5595 | 17.9 |  | 115.0625 | 24.0481 | 240.9199 | 180.7053 | 9.058 | 283.0616
2010-08-23 | 12:48:17 | 22.85 | 339.586 | 29.48 | 5591 | 17.9 |  | 114.6581 | 23.9635 | 240.0732 | 180.0702 | 9.0261 | 282.0668
2010-08-23 | 12:48:27 | 23.017 | 339.885 | 29.47 | 5599 | 17.9 |  | 114.759 | 23.9846 | 240.2846 | 180.2288 | 9.0341 | 282.3151
2010-08-23 | 12:48:37 | 23.184 | 338.096 | 29.53 | 5587 | 17.9 |  | 114.1549 | 23.8584 | 239.0197 | 179.28 | 8.9865 | 280.8289
2010-08-23 | 12:48:47 | 23.351 | 340.484 | 29.45 | 5594 | 17.9 |  | 114.9612 | 24.0269 | 240.7079 | 180.5463 | 9.05 | 282.8125
2010-08-23 | 12:48:57 | 23.518 | 338.393 | 29.52 | 5582 | 17.9 |  | 114.2553 | 23.8794 | 239.2299 | 179.4377 | 8.9944 | 281.076
2010-08-23 | 12:49:07 | 23.685 | 324.582 | 30.04 | 5650 | 17.8 |  | 109.3677 | 22.8579 | 229.0257 | 171.7839 | 8.6259 | 269.5606
2010-08-23 | 12:49:17 | 23.852 | 322.887 | 30.1 | 5684 | 17.8 |  | 108.7963 | 22.7384 | 227.8291 | 170.8864 | 8.5809 | 268.1522
2010-08-23 | 12:49:28 | 24.019 | 317.583 | 30.29 | 5763 | 17.8 |  | 107.0092 | 22.3649 | 224.0867 | 168.0793 | 8.4399 | 263.7474
2010-08-23 | 12:49:37 | 24.185 | 312.378 | 30.48 | 5819 | 17.8 |  | 105.2554 | 21.9984 | 220.4141 | 165.3247 | 8.3016 | 259.4248
2010-08-23 | 12:49:47 | 24.352 | 307.803 | 30.65 | 5832 | 17.8 |  | 103.7138 | 21.6762 | 217.186 | 162.9033 | 8.18 | 255.6253
2010-08-23 | 12:49:58 | 24.519 | 307.269 | 30.67 | 5852 | 17.8 |  | 103.5341 | 21.6386 | 216.8097 | 162.6211 | 8.1658 | 255.1824
2010-08-23 | 12:50:08 | 24.686 | 305.411 | 30.74 | 5853 | 17.8 |  | 102.908 | 21.5078 | 215.4986 | 161.6377 | 8.1165 | 253.6393
2010-08-23 | 12:50:18 | 24.853 | 306.737 | 30.69 | 5854 | 17.8 |  | 103.3548 | 21.6012 | 216.4342 | 162.3394 | 8.1517 | 254.7405
2010-08-23 | 12:50:28 | 25.019 | 308.872 | 30.61 | 5841 | 17.8 |  | 104.0742 | 21.7515 | 217.9407 | 163.4694 | 8.2084 | 256.5136
2010-08-23 | 12:50:38 | 25.186 | 307.269 | 30.67 | 5844 | 17.8 |  | 103.5341 | 21.6386 | 216.8097 | 162.6211 | 8.1658 | 255.1824
2010-08-23 | 12:50:48 | 25.353 | 308.872 | 30.61 | 5818 | 17.8 |  | 104.0742 | 21.7515 | 217.9407 | 163.4694 | 8.2084 | 256.5136
2010-08-23 | 12:50:58 | 25.52 | 308.337 | 30.63 | 5830 | 17.8 |  | 103.8938 | 21.7138 | 217.563 | 163.1861 | 8.1942 | 256.069
2010-08-23 | 12:51:08 | 25.687 | 310.35 | 30.51 | 5830 | 17.9 |  | 104.7868 | 21.9004 | 219.4046 | 164.5674 | 8.249 | 257.7828
2010-08-23 | 12:51:18 | 25.854 | 310.35 | 30.51 | 5749 | 17.9 |  | 104.7868 | 21.9004 | 219.4046 | 164.5674 | 8.249 | 257.7828
2010-08-23 | 12:51:28 | 26.021 | 310.62 | 30.5 | 5735 | 17.9 |  | 104.8779 | 21.9195 | 219.5952 | 164.7104 | 8.2562 | 258.0068
2010-08-23 | 12:51:38 | 26.188 | 321.068 | 30.12 | 5666 | 17.9 |  | 108.4054 | 22.6567 | 226.9813 | 170.2505 | 8.5339 | 266.6848
2010-08-23 | 12:51:48 | 26.356 | 305.276 | 30.7 | 5810 | 17.9 |  | 103.0737 | 21.5424 | 215.8176 | 161.877 | 8.1142 | 253.5684
2010-08-23 | 12:51:58 | 26.521 | 317.449 | 30.25 | 5697 | 17.9 |  | 107.1836 | 22.4014 | 224.4231 | 168.3316 | 8.4377 | 263.6792
2010-08-23 | 12:52:08 | 26.689 | 323.036 | 30.05 | 5680 | 17.9 |  | 109.0699 | 22.7956 | 228.3726 | 171.294 | 8.5862 | 268.3195
2010-08-23 | 12:52:18 | 26.856 | 299.518 | 30.92 | 5885 | 17.9 |  | 101.1295 | 21.1361 | 211.7468 | 158.8236 | 7.9611 | 248.7855
2010-08-23 | 12:52:28 | 27.022 | 316.621 | 30.28 | 5729 | 17.9 |  | 106.9039 | 22.3429 | 223.8374 | 167.8924 | 8.4157 | 262.991
2010-08-23 | 12:52:38 | 27.189 | 326.441 | 29.93 | 5627 | 17.9 |  | 110.2199 | 23.0359 | 230.7804 | 173.1 | 8.6768 | 271.1485
2010-08-23 | 12:52:48 | 27.357 | 335.727 | 29.61 | 5574 | 17.9 |  | 113.3551 | 23.6912 | 237.345 | 178.0239 | 8.9236 | 278.8614