## convers.+compens.
file name: PSt1_Eq4_Sal |  |  | 2008-03-29
author: Dr. Christian Huber
programm for temperature compensated oxygen calculation of membrane PSt1
Fit of the calibration curve is optimized between 0 and 250 % air-saturation (0 and 388 Torr)
enter your values here
input of F0,T0 | 59.41 | phase angle of cal 0
input of F100,T100 | 31.06 | phase angle of cal 100 (water-vapor saturated air)
input of Fm,Tm | 28.56 | measured phase angle
input of T0 | 18 | temperature of cal 0
input of T100 | 17.8 | temperatur of cal100
input of Tm | 20 | temperature at measurement
air pressure patm | 1013 | air pressure
Salinity | 29.8 | Salinity
Results of calculation + compensation
[%] air saturation | 119.6019 | %
[%] O2 | 24.9968 | %
pO2 [hPa] | 247.9494 | hPa
pO2 [Torr] | 185.9779 | Torr
cO2 [mg/L] | 9.1269 | mg/L
cO2 [ppm] | 9.1269 | ppm
cO2 [µmol/L] | 285.2172 | µmol/L
Internal calculated parameter (DO NOT CHANGE !!)
f1 | 0.801
DF/K | -0.08
DKSV/K | 0.000383
m | 22.9
tan(F0,T100) | 1.6927
tan(F0,Tm) | 1.6808
tan(F100,T100) | 0.60229
tan(Fm,Tm) | 0.54431
KSV,T100 | 0.033895
KSV,Tm | 0.034738
A | 155.3811
B | 13.7383
C | -0.64418
a | 1.7064e-05
b | 0.0036126
c | -0.67617
Chlorinity | 16.4931

## Recalculate Fibox values
file name: PSt1_Eq4_Sal |  |  |  | 2008-03-29
author: Dr. Christian Huber
programm for temperature compensated oxygen calculation of membrane PSt1
Fit of the calibration curve is optimized between 0 and 250 % air-saturation (0 and 388 Torr)
Insert your calibration values
cal0 | 59.41 | ° | T0 | 18 | °C
cal100 | 31.06 | ° | T100 | 17.8 | °C
air pressure | 1020 | mbar
Salinity | 29.8 | °/°°
Internal calculated parameter (DO NOT CHANGE !!)
f1 | 0.801 | tan(F0,T100) | 1.6927 | A | 155.3811 | KSV,T100 | 0.033895 | a | 1.7151e-05
DF/K | -0.08 | tan(F0,Tm) | 1.6798 | B | 13.7383 | KSV,Tm | 0.034814 | b | 0.003628 | [5 empty] | INSERT YOUR RESPECTIVE DATA
DKSV/K | 0.000383 | tan(F100,T100) | 0.60229 | C | -0.64418 |  |  | c | -0.67596 |  |  |  |  | Blank (chamber 3) cO2 [µmol/ h] | chamber volume [L] | Fresh weight [g]
m | 22.9 | tan(Fm,Tm) | 0.54431 |  |  |  |  | Chlorinity | 16.4931 |  |  |  |  | -0.41107 | 0.02607 | 0.049
Insert the data you get from the Ascii File recorded with the OxyView TX3-V5.31.exe software | [7 empty] | The oxygen contents are calculated automatically from the raw data
date | time/hh:mm:ss | logtime/min | oxygen/% airsatur. | phase/° | amp | temp/°C |  | % air-sat. | % oxygen | pO2 (hPa) | pO2 (Torr) | cO2 [mg/L] | cO2 [µmol/L] |  | cO2 [µmol/L 10 min] | cO2 [µmol/L h] | cO2 [µmol/ h] | cO2 [µmol/ h gFW]
2010-07-28 | 16:24:54 | 0 | 337.057 | 28.56 | 5461 | 20.2 |  | 119.1764 | 24.9079 | 248.7421 | 186.5724 | 9.1247 | 285.1465 |  | 26.982 | 161.892 | 4.6315 | 94.5213
2010-07-28 | 16:25:05 | 0.184 | 339.167 | 28.49 | 5462 | 20.2 |  | 119.9223 | 25.0638 | 250.2989 | 187.7401 | 9.1818 | 286.9312
2010-07-28 | 16:25:15 | 0.351 | 337.358 | 28.55 | 5455 | 20.2 |  | 119.2826 | 24.9301 | 248.9638 | 186.7387 | 9.1328 | 285.4007 |  | change data input according to column N
2010-07-28 | 16:25:25 | 0.518 | 344.663 | 28.31 | 5455 | 20.2 |  | 121.8656 | 25.4699 | 254.355 | 190.7825 | 9.3306 | 291.581
2010-07-28 | 16:25:35 | 0.684 | 339.469 | 28.48 | 5444 | 20.2 |  | 120.0293 | 25.0861 | 250.5222 | 187.9076 | 9.19 | 287.1872
2010-07-28 | 16:25:45 | 0.851 | 338.563 | 28.51 | 5445 | 20.2 |  | 119.7086 | 25.0191 | 249.8529 | 187.4056 | 9.1654 | 286.42
2010-07-28 | 16:25:55 | 1.018 | 340.379 | 28.45 | 5443 | 20.2 |  | 120.3509 | 25.1533 | 251.1936 | 188.4112 | 9.2146 | 287.9569
2010-07-28 | 16:26:05 | 1.185 | 342.819 | 28.37 | 5443 | 20.2 |  | 121.2137 | 25.3337 | 252.9944 | 189.7619 | 9.2807 | 290.0212
2010-07-28 | 16:26:15 | 1.352 | 343.433 | 28.35 | 5444 | 20.2 |  | 121.4306 | 25.379 | 253.447 | 190.1014 | 9.2973 | 290.5401
2010-07-28 | 16:26:25 | 1.519 | 340.076 | 28.46 | 5446 | 20.2 |  | 120.2436 | 25.1309 | 250.9696 | 188.2432 | 9.2064 | 287.7001
2010-07-28 | 16:26:35 | 1.686 | 338.261 | 28.52 | 5443 | 20.2 |  | 119.6019 | 24.9968 | 249.6303 | 187.2386 | 9.1573 | 286.1647
2010-07-28 | 16:26:45 | 1.853 | 342.819 | 28.37 | 5436 | 20.2 |  | 121.2137 | 25.3337 | 252.9944 | 189.7619 | 9.2807 | 290.0212
2010-07-28 | 16:26:55 | 2.02 | 341.597 | 28.41 | 5450 | 20.2 |  | 120.7814 | 25.2433 | 252.0921 | 189.0852 | 9.2476 | 288.9869
2010-07-28 | 16:27:05 | 2.187 | 338.865 | 28.5 | 5449 | 20.2 |  | 119.8154 | 25.0414 | 250.0758 | 187.5728 | 9.1736 | 286.6754
2010-07-28 | 16:27:15 | 2.353 | 339.772 | 28.47 | 5445 | 20.2 |  | 120.1364 | 25.1085 | 250.7458 | 188.0753 | 9.1982 | 287.4435
2010-07-28 | 16:27:25 | 2.52 | 337.658 | 28.54 | 5452 | 20.2 |  | 119.3889 | 24.9523 | 249.1857 | 186.9052 | 9.141 | 285.6551
2010-07-28 | 16:27:35 | 2.687 | 339.469 | 28.48 | 5443 | 20.2 |  | 120.0293 | 25.0861 | 250.5222 | 187.9076 | 9.19 | 287.1872
2010-07-28 | 16:27:45 | 2.854 | 342.513 | 28.38 | 5449 | 20.2 |  | 121.1055 | 25.311 | 252.7685 | 189.5925 | 9.2724 | 289.7622
2010-07-28 | 16:27:55 | 3.021 | 338.563 | 28.51 | 5453 | 20.2 |  | 119.7086 | 25.0191 | 249.8529 | 187.4056 | 9.1654 | 286.42
2010-07-28 | 16:28:05 | 3.188 | 339.228 | 28.53 | 5433 | 20.1 |  | 119.7084 | 25.0191 | 249.8892 | 187.4329 | 9.1825 | 286.9542
2010-07-28 | 16:28:15 | 3.355 | 341.352 | 28.46 | 5446 | 20.1 |  | 120.458 | 25.1757 | 251.4539 | 188.6064 | 9.24 | 288.7509
2010-07-28 | 16:28:25 | 3.522 | 342.573 | 28.42 | 5443 | 20.1 |  | 120.8887 | 25.2657 | 252.3531 | 189.2809 | 9.2731 | 289.7836
2010-07-28 | 16:28:35 | 3.689 | 338.323 | 28.56 | 5456 | 20.1 |  | 119.3889 | 24.9523 | 249.2222 | 186.9325 | 9.158 | 286.1883 |  |  | regression formula | time
2010-07-28 | 16:28:45 | 3.856 | 340.744 | 28.48 | 5445 | 20.1 |  | 120.2432 | 25.1308 | 251.0057 | 188.2703 | 9.2236 | 288.2363 |  | value for T=26 min. | 336.866 | T11
2010-07-28 | 16:28:55 | 4.022 | 336.82 | 28.61 | 5451 | 20.1 |  | 118.8585 | 24.8414 | 248.1151 | 186.1021 | 9.1173 | 284.9169 |  | value for T=1 min. | 309.884 | T1
2010-07-28 | 16:29:05 | 4.19 | 341.048 | 28.47 | 5456 | 20.1 |  | 120.3505 | 25.1533 | 251.2297 | 188.4383 | 9.2318 | 288.4935 |  | difference between T25 and T1 | 26.982 | 10minutes
2010-07-28 | 16:29:15 | 4.356 | 337.42 | 28.59 | 5456 | 20.1 |  | 119.0703 | 24.8857 | 248.5572 | 186.4338 | 9.1336 | 285.4247 |  | calculate from regression curve values for 10 minutes of photosynthesis or respiration
2010-07-28 | 16:29:25 | 4.523 | 341.415 | 28.5 | 5453 | 20 |  | 120.2431 | 25.1308 | 251.0421 | 188.2976 | 9.2408 | 288.7744
2010-07-28 | 16:29:35 | 4.69 | 338.086 | 28.61 | 5456 | 20 |  | 119.0707 | 24.8858 | 248.5942 | 186.4615 | 9.1507 | 285.9586
2010-07-28 | 16:29:45 | 4.857 | 338.688 | 28.59 | 5471 | 20 |  | 119.2828 | 24.9301 | 249.0372 | 186.7938 | 9.167 | 286.4681
2010-07-28 | 16:29:55 | 5.024 | 341.111 | 28.51 | 5464 | 20 |  | 120.136 | 25.1084 | 250.8184 | 188.1298 | 9.2325 | 288.5171
2010-07-28 | 16:30:05 | 5.191 | 341.111 | 28.51 | 5460 | 20 |  | 120.136 | 25.1084 | 250.8184 | 188.1298 | 9.2325 | 288.5171
2010-07-28 | 16:30:15 | 5.358 | 339.897 | 28.55 | 5468 | 20 |  | 119.7085 | 25.0191 | 249.9259 | 187.4604 | 9.1997 | 287.4904
2010-07-28 | 16:30:25 | 5.525 | 341.415 | 28.5 | 5464 | 20 |  | 120.2431 | 25.1308 | 251.0421 | 188.2976 | 9.2408 | 288.7744
2010-07-28 | 16:30:35 | 5.692 | 340.873 | 28.56 | 5468 | 19.9 |  | 119.8155 | 25.0414 | 250.1856 | 187.6551 | 9.2251 | 288.2852
2010-07-28 | 16:30:45 | 5.859 | 338.454 | 28.64 | 5472 | 19.9 |  | 118.9653 | 24.8638 | 248.4104 | 186.3236 | 9.1597 | 286.2396
2010-07-28 | 16:30:55 | 6.025 | 340.569 | 28.57 | 5474 | 19.9 |  | 119.7088 | 25.0191 | 249.9629 | 187.4881 | 9.2169 | 288.0286
2010-07-28 | 16:31:05 | 6.192 | 338.153 | 28.65 | 5471 | 19.9 |  | 118.8596 | 24.8416 | 248.1895 | 186.1579 | 9.1515 | 285.9851
2010-07-28 | 16:31:15 | 6.359 | 342.09 | 28.52 | 5481 | 19.9 |  | 120.2433 | 25.1308 | 251.0788 | 188.3251 | 9.2581 | 289.3144
2010-07-28 | 16:31:25 | 6.526 | 339.963 | 28.59 | 5489 | 19.9 |  | 119.4958 | 24.9746 | 249.5181 | 187.1545 | 9.2005 | 287.5161
2010-07-28 | 16:31:35 | 6.693 | 339.428 | 28.65 | 5490 | 19.8 |  | 119.072 | 24.886 | 248.669 | 186.5176 | 9.185 | 287.0321
2010-07-28 | 16:31:45 | 6.86 | 337.922 | 28.7 | 5487 | 19.8 |  | 118.5439 | 24.7757 | 247.5661 | 185.6903 | 9.1443 | 285.759
2010-07-28 | 16:31:55 | 7.027 | 338.523 | 28.68 | 5481 | 19.8 |  | 118.7548 | 24.8198 | 248.0066 | 186.0207 | 9.1606 | 286.2675
2010-07-28 | 16:32:05 | 7.194 | 338.223 | 28.69 | 5480 | 19.8 |  | 118.6493 | 24.7977 | 247.7862 | 185.8555 | 9.1524 | 286.0131
2010-07-28 | 16:32:15 | 7.361 | 339.126 | 28.66 | 5490 | 19.8 |  | 118.9662 | 24.8639 | 248.448 | 186.3518 | 9.1769 | 286.777
2010-07-28 | 16:32:25 | 7.528 | 338.896 | 28.71 | 5497 | 19.7 |  | 118.6505 | 24.7979 | 247.8243 | 185.884 | 9.1696 | 286.5513
2010-07-28 | 16:32:35 | 7.694 | 338.896 | 28.71 | 5487 | 19.7 |  | 118.6505 | 24.7979 | 247.8243 | 185.884 | 9.1696 | 286.5513
2010-07-28 | 16:32:45 | 7.861 | 337.694 | 28.75 | 5489 | 19.7 |  | 118.2297 | 24.71 | 246.9453 | 185.2247 | 9.1371 | 285.535
2010-07-28 | 16:32:55 | 8.028 | 338.896 | 28.71 | 5497 | 19.7 |  | 118.6505 | 24.7979 | 247.8243 | 185.884 | 9.1696 | 286.5513
2010-07-28 | 16:33:05 | 8.195 | 338.595 | 28.72 | 5502 | 19.7 |  | 118.5451 | 24.7759 | 247.6042 | 185.7189 | 9.1615 | 286.2968
2010-07-28 | 16:33:15 | 8.362 | 344.133 | 28.58 | 5507 | 19.6 |  | 120.2451 | 25.1312 | 251.1909 | 188.4092 | 9.3103 | 290.9464
2010-07-28 | 16:33:25 | 8.529 | 340.782 | 28.69 | 5504 | 19.6 |  | 119.0743 | 24.8865 | 248.7451 | 186.5747 | 9.2196 | 288.1136
2010-07-28 | 16:33:35 | 8.696 | 339.875 | 28.72 | 5508 | 19.6 |  | 118.7573 | 24.8203 | 248.0829 | 186.078 | 9.1951 | 287.3466
2010-07-28 | 16:33:45 | 8.863 | 341.996 | 28.65 | 5508 | 19.6 |  | 119.4985 | 24.9752 | 249.6312 | 187.2393 | 9.2525 | 289.1399
2010-07-28 | 16:33:55 | 9.03 | 339.272 | 28.74 | 5525 | 19.6 |  | 118.5466 | 24.7762 | 247.6426 | 185.7477 | 9.1788 | 286.8366
2010-07-28 | 16:34:05 | 9.197 | 340.48 | 28.7 | 5512 | 19.6 |  | 118.9686 | 24.8644 | 248.5241 | 186.4089 | 9.2114 | 287.8576
2010-07-28 | 16:34:14 | 9.347 | 337.171 | 28.81 | 5526 | 19.6 |  | 117.8123 | 24.6228 | 246.1087 | 184.5972 | 9.1219 | 285.06
2010-07-28 | 16:34:25 | 9.514 | 336.275 | 28.84 | 5522 | 19.6 |  | 117.4993 | 24.5573 | 245.4548 | 184.1067 | 9.0977 | 284.3025
2010-07-28 | 16:34:35 | 9.681 | 339.875 | 28.72 | 5526 | 19.6 |  | 118.7573 | 24.8203 | 248.0829 | 186.078 | 9.1951 | 287.3466
2010-07-28 | 16:34:45 | 9.847 | 340.177 | 28.71 | 5525 | 19.6 |  | 118.8629 | 24.8423 | 248.3034 | 186.2434 | 9.2033 | 287.602
2010-07-28 | 16:34:55 | 10.014 | 335.977 | 28.85 | 5535 | 19.6 |  | 117.3951 | 24.5356 | 245.2372 | 183.9436 | 9.0896 | 284.0505
2010-07-28 | 16:35:05 | 10.181 | 337.47 | 28.8 | 5543 | 19.6 |  | 117.9169 | 24.6446 | 246.3272 | 184.7611 | 9.13 | 285.313
2010-07-28 | 16:35:15 | 10.348 | 336.872 | 28.82 | 5535 | 19.6 |  | 117.7079 | 24.6009 | 245.8905 | 184.4336 | 9.1138 | 284.8072
2010-07-28 | 16:35:25 | 10.515 | 336.573 | 28.83 | 5536 | 19.6 |  | 117.6035 | 24.5791 | 245.6725 | 184.2701 | 9.1058 | 284.5547
2010-07-28 | 16:35:35 | 10.682 | 338.069 | 28.78 | 5538 | 19.6 |  | 118.1263 | 24.6884 | 246.7647 | 185.0893 | 9.1462 | 285.8198
2010-07-28 | 16:35:45 | 10.849 | 336.872 | 28.82 | 5542 | 19.6 |  | 117.7079 | 24.6009 | 245.8905 | 184.4336 | 9.1138 | 284.8072
2010-07-28 | 16:35:55 | 11.016 | 337.171 | 28.81 | 5543 | 19.6 |  | 117.8123 | 24.6228 | 246.1087 | 184.5972 | 9.1219 | 285.06
2010-07-28 | 16:36:05 | 11.183 | 335.757 | 28.9 | 5540 | 19.5 |  | 117.0856 | 24.4709 | 244.6253 | 183.4846 | 9.0826 | 283.8325
2010-07-28 | 16:36:15 | 11.35 | 339.048 | 28.79 | 5539 | 19.5 |  | 118.233 | 24.7107 | 247.0227 | 185.2828 | 9.1717 | 286.6142
2010-07-28 | 16:36:25 | 11.517 | 337.248 | 28.85 | 5538 | 19.5 |  | 117.6055 | 24.5796 | 245.7117 | 184.2994 | 9.123 | 285.093
2010-07-28 | 16:36:35 | 11.683 | 335.757 | 28.9 | 5548 | 19.5 |  | 117.0856 | 24.4709 | 244.6253 | 183.4846 | 9.0826 | 283.8325
2010-07-28 | 16:36:45 | 11.85 | 338.747 | 28.8 | 5546 | 19.5 |  | 118.1282 | 24.6888 | 246.8036 | 185.1185 | 9.1635 | 286.36
2010-07-28 | 16:36:55 | 12.017 | 339.952 | 28.76 | 5547 | 19.5 |  | 118.5483 | 24.7766 | 247.6813 | 185.7768 | 9.1961 | 287.3783
2010-07-28 | 16:37:05 | 12.184 | 333.978 | 28.96 | 5549 | 19.5 |  | 116.4652 | 24.3412 | 243.3291 | 182.5124 | 9.0345 | 282.3286
2010-07-28 | 16:37:15 | 12.351 | 335.757 | 28.9 | 5550 | 19.5 |  | 117.0856 | 24.4709 | 244.6253 | 183.4846 | 9.0826 | 283.8325
2010-07-28 | 16:37:25 | 12.518 | 337.847 | 28.83 | 5567 | 19.5 |  | 117.8143 | 24.6232 | 246.1478 | 184.6265 | 9.1392 | 285.599
2010-07-28 | 16:37:35 | 12.685 | 336.353 | 28.88 | 5550 | 19.5 |  | 117.2932 | 24.5143 | 245.0592 | 183.81 | 9.0987 | 284.3359
2010-07-28 | 16:37:45 | 12.852 | 337.628 | 28.88 | 5570 | 19.4 |  | 117.5036 | 24.5583 | 245.5334 | 184.1657 | 9.1322 | 285.3802
2010-07-28 | 16:37:55 | 13.019 | 340.031 | 28.8 | 5555 | 19.4 |  | 118.34 | 24.7331 | 247.2812 | 185.4766 | 9.1972 | 287.4116
2010-07-28 | 16:38:05 | 13.186 | 335.838 | 28.94 | 5557 | 19.4 |  | 116.8809 | 24.4281 | 244.2321 | 183.1896 | 9.0838 | 283.8677
2010-07-28 | 16:38:15 | 13.353 | 337.628 | 28.88 | 5559 | 19.4 |  | 117.5036 | 24.5583 | 245.5334 | 184.1657 | 9.1322 | 285.3802
2010-07-28 | 16:38:25 | 13.52 | 337.628 | 28.88 | 5569 | 19.4 |  | 117.5036 | 24.5583 | 245.5334 | 184.1657 | 9.1322 | 285.3802
2010-07-28 | 16:38:35 | 13.687 | 337.03 | 28.9 | 5572 | 19.4 |  | 117.2956 | 24.5148 | 245.0987 | 183.8397 | 9.116 | 284.875
2010-07-28 | 16:38:45 | 13.853 | 335.625 | 28.99 | 5570 | 19.3 |  | 116.5738 | 24.3639 | 243.6246 | 182.734 | 9.0769 | 283.6534
2010-07-28 | 16:38:55 | 14.02 | 338.009 | 28.91 | 5569 | 19.3 |  | 117.4021 | 24.537 | 245.3558 | 184.0325 | 9.1414 | 285.669
2010-07-28 | 16:39:05 | 14.187 | 336.815 | 28.95 | 5576 | 19.3 |  | 116.9871 | 24.4503 | 244.4884 | 183.3819 | 9.1091 | 284.6591
2010-07-28 | 16:39:15 | 14.354 | 339.811 | 28.85 | 5583 | 19.3 |  | 118.0279 | 24.6678 | 246.6636 | 185.0134 | 9.1901 | 287.1917
2010-07-28 | 16:39:25 | 14.521 | 338.009 | 28.91 | 5582 | 19.3 |  | 117.4021 | 24.537 | 245.3558 | 184.0325 | 9.1414 | 285.669
2010-07-28 | 16:39:35 | 14.688 | 334.735 | 29.02 | 5582 | 19.3 |  | 116.2649 | 24.2994 | 242.9791 | 182.2498 | 9.0529 | 282.9018
2010-07-28 | 16:39:45 | 14.855 | 337.411 | 28.93 | 5585 | 19.3 |  | 117.1944 | 24.4936 | 244.9216 | 183.7068 | 9.1252 | 285.1635
2010-07-28 | 16:39:55 | 15.022 | 338.009 | 28.91 | 5588 | 19.3 |  | 117.4021 | 24.537 | 245.3558 | 184.0325 | 9.1414 | 285.669
2010-07-28 | 16:40:05 | 15.189 | 335.031 | 29.01 | 5598 | 19.3 |  | 116.3677 | 24.3209 | 243.194 | 182.411 | 9.0609 | 283.1521
2010-07-28 | 16:40:15 | 15.356 | 338.009 | 28.91 | 5592 | 19.3 |  | 117.4021 | 24.537 | 245.3558 | 184.0325 | 9.1414 | 285.669
2010-07-28 | 16:40:25 | 15.522 | 336.219 | 28.97 | 5595 | 19.3 |  | 116.7802 | 24.4071 | 244.056 | 183.0576 | 9.093 | 284.1557
2010-07-28 | 16:40:35 | 15.689 | 337.113 | 28.94 | 5596 | 19.3 |  | 117.0907 | 24.472 | 244.7049 | 183.5443 | 9.1172 | 284.9112
2010-07-28 | 16:40:45 | 15.856 | 334.144 | 29.04 | 5594 | 19.3 |  | 116.0595 | 24.2564 | 242.5498 | 181.9278 | 9.0369 | 282.402
2010-07-28 | 16:40:55 | 16.023 | 336.219 | 28.97 | 5589 | 19.3 |  | 116.7802 | 24.4071 | 244.056 | 183.0576 | 9.093 | 284.1557
2010-07-28 | 16:41:05 | 16.19 | 335.625 | 28.99 | 5601 | 19.3 |  | 116.5738 | 24.3639 | 243.6246 | 182.734 | 9.0769 | 283.6534
2010-07-28 | 16:41:15 | 16.357 | 336.219 | 28.97 | 5606 | 19.3 |  | 116.7802 | 24.4071 | 244.056 | 183.0576 | 9.093 | 284.1557
2010-07-28 | 16:41:25 | 16.524 | 334.735 | 29.02 | 5604 | 19.3 |  | 116.2649 | 24.2994 | 242.9791 | 182.2498 | 9.0529 | 282.9018
2010-07-28 | 16:41:35 | 16.691 | 333.554 | 29.06 | 5605 | 19.3 |  | 115.8545 | 24.2136 | 242.1215 | 181.6066 | 9.0209 | 281.9033
2010-07-28 | 16:41:45 | 16.858 | 332.671 | 29.09 | 5601 | 19.3 |  | 115.5479 | 24.1495 | 241.4806 | 181.1259 | 8.997 | 281.1571
2010-07-28 | 16:41:55 | 17.025 | 334.735 | 29.02 | 5598 | 19.3 |  | 116.2649 | 24.2994 | 242.9791 | 182.2498 | 9.0529 | 282.9018
2010-07-28 | 16:42:05 | 17.192 | 337.71 | 28.92 | 5602 | 19.3 |  | 117.2982 | 24.5153 | 245.1386 | 183.8696 | 9.1333 | 285.4161
2010-07-28 | 16:42:15 | 17.358 | 335.031 | 29.01 | 5608 | 19.3 |  | 116.3677 | 24.3209 | 243.194 | 182.411 | 9.0609 | 283.1521
2010-07-28 | 16:42:25 | 17.525 | 337.71 | 28.92 | 5598 | 19.3 |  | 117.2982 | 24.5153 | 245.1386 | 183.8696 | 9.1333 | 285.4161
2010-07-28 | 16:42:35 | 17.692 | 336.219 | 28.97 | 5602 | 19.3 |  | 116.7802 | 24.4071 | 244.056 | 183.0576 | 9.093 | 284.1557
2010-07-28 | 16:42:45 | 17.859 | 336.517 | 28.96 | 5595 | 19.3 |  | 116.8836 | 24.4287 | 244.2721 | 183.2197 | 9.101 | 284.4073
2010-07-28 | 16:42:55 | 18.026 | 333.554 | 29.06 | 5594 | 19.3 |  | 115.8545 | 24.2136 | 242.1215 | 181.6066 | 9.0209 | 281.9033
2010-07-28 | 16:43:05 | 18.193 | 335.922 | 28.98 | 5588 | 19.3 |  | 116.6769 | 24.3855 | 243.8402 | 182.8957 | 9.0849 | 283.9044
2010-07-28 | 16:43:15 | 18.36 | 335.711 | 29.03 | 5578 | 19.2 |  | 116.3709 | 24.3215 | 243.2346 | 182.4414 | 9.0781 | 283.6917
2010-07-28 | 16:43:25 | 18.527 | 340.498 | 28.87 | 5591 | 19.2 |  | 118.0305 | 24.6684 | 246.7035 | 185.0434 | 9.2076 | 287.7376
2010-07-28 | 16:43:35 | 18.694 | 338.694 | 28.93 | 5574 | 19.2 |  | 117.4049 | 24.5376 | 245.3959 | 184.0626 | 9.1588 | 286.2126
2010-07-28 | 16:43:45 | 18.86 | 336.008 | 29.02 | 5584 | 19.2 |  | 116.4738 | 24.343 | 243.4497 | 182.6028 | 9.0862 | 283.9426
2010-07-28 | 16:43:55 | 19.027 | 337.796 | 28.96 | 5566 | 19.2 |  | 117.0936 | 24.4726 | 244.7452 | 183.5745 | 9.1345 | 285.4536
2010-07-28 | 16:44:05 | 19.194 | 341.102 | 28.85 | 5568 | 19.2 |  | 118.2399 | 24.7121 | 247.1412 | 185.3716 | 9.2239 | 288.2481
2010-07-28 | 16:44:15 | 19.361 | 340.8 | 28.86 | 5557 | 19.2 |  | 118.1352 | 24.6903 | 246.9222 | 185.2074 | 9.2158 | 287.9927
2010-07-28 | 16:44:25 | 19.528 | 343.836 | 28.76 | 5564 | 19.2 |  | 119.1876 | 24.9102 | 249.122 | 186.8574 | 9.2979 | 290.5584
2010-07-28 | 16:44:35 | 19.695 | 341.794 | 28.87 | 5562 | 19.1 |  | 118.2427 | 24.7127 | 247.1813 | 185.4017 | 9.2415 | 288.7969
2010-07-28 | 16:44:45 | 19.862 | 342.4 | 28.85 | 5550 | 19.1 |  | 118.4525 | 24.7566 | 247.6198 | 185.7306 | 9.2579 | 289.3092
2010-07-28 | 16:44:55 | 20.029 | 343.008 | 28.83 | 5544 | 19.1 |  | 118.6627 | 24.8005 | 248.0592 | 186.0602 | 9.2743 | 289.8226
2010-07-28 | 16:45:05 | 20.196 | 347.297 | 28.69 | 5544 | 19.1 |  | 120.1463 | 25.1106 | 251.1608 | 188.3866 | 9.3903 | 293.4463
2010-07-28 | 16:45:15 | 20.363 | 342.704 | 28.84 | 5535 | 19.1 |  | 118.5575 | 24.7785 | 247.8394 | 185.8953 | 9.2661 | 289.5658
2010-07-28 | 16:45:25 | 20.53 | 346.372 | 28.72 | 5539 | 19.1 |  | 119.8266 | 25.0438 | 250.4923 | 187.8852 | 9.3653 | 292.6654
2010-07-28 | 16:45:35 | 20.697 | 346.988 | 28.7 | 5538 | 19.1 |  | 120.0396 | 25.0883 | 250.9377 | 188.2193 | 9.3819 | 293.1857
2010-07-28 | 16:45:45 | 20.863 | 346.988 | 28.7 | 5519 | 19.1 |  | 120.0396 | 25.0883 | 250.9377 | 188.2193 | 9.3819 | 293.1857
2010-07-28 | 16:45:54 | 21.014 | 348.534 | 28.65 | 5533 | 19.1 |  | 120.5742 | 25.2 | 252.0553 | 189.0575 | 9.4237 | 294.4914
2010-07-28 | 16:46:05 | 21.181 | 345.451 | 28.75 | 5513 | 19.1 |  | 119.5078 | 24.9771 | 249.826 | 187.3854 | 9.3404 | 291.8868
2010-07-28 | 16:46:15 | 21.348 | 349.776 | 28.61 | 5501 | 19.1 |  | 121.0039 | 25.2898 | 252.9535 | 189.7312 | 9.4573 | 295.5409
2010-07-28 | 16:46:25 | 21.514 | 348.45 | 28.61 | 5506 | 19.2 |  | 120.7869 | 25.2445 | 252.4649 | 189.3647 | 9.4226 | 294.4573
2010-07-28 | 16:46:35 | 21.681 | 347.213 | 28.65 | 5515 | 19.2 |  | 120.358 | 25.1548 | 251.5683 | 188.6923 | 9.3892 | 293.4116
2010-07-28 | 16:46:45 | 21.848 | 350.628 | 28.54 | 5494 | 19.2 |  | 121.5419 | 25.4023 | 254.0429 | 190.5484 | 9.4815 | 296.2978
2010-07-28 | 16:46:55 | 22.015 | 348.76 | 28.6 | 5489 | 19.2 |  | 120.8945 | 25.2669 | 252.6896 | 189.5333 | 9.431 | 294.7194
2010-07-28 | 16:47:05 | 22.182 | 349.693 | 28.57 | 5481 | 19.2 |  | 121.2177 | 25.3345 | 253.3652 | 190.04 | 9.4562 | 295.5073
2010-07-28 | 16:47:15 | 22.349 | 350.628 | 28.54 | 5483 | 19.2 |  | 121.5419 | 25.4023 | 254.0429 | 190.5484 | 9.4815 | 296.2978
2010-07-28 | 16:47:25 | 22.516 | 350.941 | 28.53 | 5478 | 19.2 |  | 121.6502 | 25.4249 | 254.2692 | 190.7181 | 9.49 | 296.5618
2010-07-28 | 16:47:35 | 22.683 | 351.253 | 28.52 | 5468 | 19.2 |  | 121.7586 | 25.4476 | 254.4959 | 190.8881 | 9.4984 | 296.8261
2010-07-28 | 16:47:45 | 22.85 | 353.767 | 28.44 | 5475 | 19.2 |  | 122.6301 | 25.6297 | 256.3173 | 192.2544 | 9.5664 | 298.9506
2010-07-28 | 16:47:55 | 23.017 | 351.253 | 28.52 | 5466 | 19.2 |  | 121.7586 | 25.4476 | 254.4959 | 190.8881 | 9.4984 | 296.8261
2010-07-28 | 16:48:05 | 23.183 | 353.452 | 28.45 | 5455 | 19.2 |  | 122.5208 | 25.6068 | 256.0888 | 192.0829 | 9.5579 | 298.684
2010-07-28 | 16:48:15 | 23.35 | 354.716 | 28.41 | 5452 | 19.2 |  | 122.9588 | 25.6984 | 257.0044 | 192.7697 | 9.5921 | 299.7518
2010-07-28 | 16:48:25 | 23.517 | 352.508 | 28.48 | 5459 | 19.2 |  | 122.1934 | 25.5384 | 255.4047 | 191.5698 | 9.5324 | 297.8861
2010-07-28 | 16:48:35 | 23.684 | 355.667 | 28.38 | 5449 | 19.2 |  | 123.2885 | 25.7673 | 257.6935 | 193.2866 | 9.6178 | 300.5556
2010-07-28 | 16:48:45 | 23.851 | 353.767 | 28.44 | 5437 | 19.2 |  | 122.6301 | 25.6297 | 256.3173 | 192.2544 | 9.5664 | 298.9506
2010-07-28 | 16:48:55 | 24.018 | 355.349 | 28.39 | 5441 | 19.2 |  | 123.1785 | 25.7443 | 257.4636 | 193.1141 | 9.6092 | 300.2874
2010-07-28 | 16:49:05 | 24.185 | 356.621 | 28.35 | 5426 | 19.2 |  | 123.6193 | 25.8364 | 258.3849 | 193.8052 | 9.6436 | 301.362
2010-07-28 | 16:49:15 | 24.352 | 354.716 | 28.41 | 5433 | 19.2 |  | 122.9588 | 25.6984 | 257.0044 | 192.7697 | 9.5921 | 299.7518
2010-07-28 | 16:49:25 | 24.519 | 355.985 | 28.37 | 5414 | 19.2 |  | 123.3986 | 25.7903 | 257.9238 | 193.4593 | 9.6264 | 300.8241
2010-07-28 | 16:49:35 | 24.686 | 357.657 | 28.36 | 5422 | 19.1 |  | 123.7306 | 25.8597 | 258.6535 | 194.0067 | 9.6704 | 302.2006
2010-07-28 | 16:49:45 | 24.852 | 357.019 | 28.38 | 5412 | 19.1 |  | 123.5098 | 25.8136 | 258.192 | 193.6605 | 9.6532 | 301.6614
2010-07-28 | 16:49:55 | 25.019 | 363.137 | 28.19 | 5411 | 19.1 |  | 125.6262 | 26.2559 | 262.6161 | 196.9788 | 9.8186 | 306.8303
2010-07-28 | 16:50:05 | 25.186 | 357.019 | 28.38 | 5409 | 19.1 |  | 123.5098 | 25.8136 | 258.192 | 193.6605 | 9.6532 | 301.6614
2010-07-28 | 16:50:15 | 25.353 | 361.515 | 28.24 | 5399 | 19.1 |  | 125.0651 | 26.1386 | 261.4433 | 196.0991 | 9.7747 | 305.46
2010-07-28 | 16:50:25 | 25.52 | 361.515 | 28.24 | 5399 | 19.1 |  | 125.0651 | 26.1386 | 261.4433 | 196.0991 | 9.7747 | 305.46
2010-07-28 | 16:50:35 | 25.687 | 358.617 | 28.33 | 5395 | 19.1 |  | 124.0627 | 25.9291 | 259.3477 | 194.5273 | 9.6964 | 303.0116
2010-07-28 | 16:50:45 | 25.854 | 361.515 | 28.24 | 5381 | 19.1 |  | 125.0651 | 26.1386 | 261.4433 | 196.0991 | 9.7747 | 305.46
2010-07-28 | 16:50:55 | 26.021 | 363.137 | 28.19 | 5399 | 19.1 |  | 125.6262 | 26.2559 | 262.6161 | 196.9788 | 9.8186 | 306.8303
2010-07-28 | 16:51:05 | 26.188 | 366.156 | 28.14 | 5386 | 19 |  | 126.4168 | 26.4211 | 264.3053 | 198.2458 | 9.899 | 309.3423
2010-07-28 | 16:51:15 | 26.355 | 365.173 | 28.17 | 5380 | 19 |  | 126.0774 | 26.3502 | 263.5958 | 197.7136 | 9.8724 | 308.5119
2010-07-28 | 16:51:25 | 26.521 | 365.173 | 28.17 | 5374 | 19 |  | 126.0774 | 26.3502 | 263.5958 | 197.7136 | 9.8724 | 308.5119
2010-07-28 | 16:51:35 | 26.688 | 366.484 | 28.13 | 5373 | 19 |  | 126.5301 | 26.4448 | 264.5423 | 198.4236 | 9.9078 | 309.6197
2010-07-28 | 16:51:45 | 26.855 | 363.541 | 28.22 | 5366 | 19 |  | 125.5142 | 26.2325 | 262.4182 | 196.8304 | 9.8283 | 307.1337
2010-07-28 | 16:51:55 | 27.022 | 366.813 | 28.12 | 5366 | 19 |  | 126.6436 | 26.4685 | 264.7796 | 198.6016 | 9.9167 | 309.8974
2010-07-28 | 16:52:05 | 27.189 | 366.156 | 28.14 | 5362 | 19 |  | 126.4168 | 26.4211 | 264.3053 | 198.2458 | 9.899 | 309.3423
2010-07-28 | 16:52:15 | 27.356 | 369.123 | 28.05 | 5361 | 19 |  | 127.4414 | 26.6352 | 266.4475 | 199.8526 | 9.9792 | 311.8496
2010-07-28 | 16:52:25 | 27.523 | 367.142 | 28.11 | 5356 | 19 |  | 126.7572 | 26.4923 | 265.0171 | 198.7797 | 9.9256 | 310.1754
2010-07-28 | 16:52:35 | 27.69 | 366.484 | 28.13 | 5340 | 19 |  | 126.5301 | 26.4448 | 264.5423 | 198.4236 | 9.9078 | 309.6197
2010-07-28 | 16:52:45 | 27.857 | 369.455 | 28.04 | 5336 | 19 |  | 127.5558 | 26.6592 | 266.6868 | 200.0321 | 9.9881 | 312.1296
2010-07-28 | 16:52:55 | 28.024 | 371.785 | 27.97 | 5342 | 19 |  | 128.3604 | 26.8273 | 268.369 | 201.2938 | 10.0511 | 314.0984
2010-07-28 | 16:53:05 | 28.191 | 369.787 | 28.03 | 5339 | 19 |  | 127.6704 | 26.6831 | 266.9263 | 200.2118 | 9.9971 | 312.41
2010-07-28 | 16:53:15 | 28.357 | 374.47 | 27.89 | 5333 | 19 |  | 129.2873 | 27.021 | 270.3068 | 202.7474 | 10.1237 | 316.3665
2010-07-28 | 16:53:25 | 28.525 | 370.709 | 27.96 | 5328 | 19.1 |  | 128.2457 | 26.8033 | 268.0921 | 201.0862 | 10.0233 | 313.2282
2010-07-28 | 16:53:35 | 28.691 | 368.385 | 28.03 | 5321 | 19.1 |  | 127.4416 | 26.6353 | 266.4113 | 199.8255 | 9.9605 | 311.2645
2010-07-28 | 16:53:45 | 28.858 | 373.05 | 27.89 | 5313 | 19.1 |  | 129.0557 | 26.9726 | 269.7855 | 202.3563 | 10.0866 | 315.2067
2010-07-28 | 16:53:55 | 29.025 | 370.376 | 27.97 | 5314 | 19.1 |  | 128.1304 | 26.7793 | 267.8512 | 200.9055 | 10.0143 | 312.9468
2010-07-28 | 16:54:05 | 29.192 | 372.38 | 27.91 | 5304 | 19.1 |  | 128.8237 | 26.9241 | 269.3004 | 201.9924 | 10.0685 | 314.6399
2010-07-28 | 16:54:15 | 29.359 | 372.715 | 27.9 | 5305 | 19.1 |  | 128.9396 | 26.9484 | 269.5428 | 202.1743 | 10.0775 | 314.9232
2010-07-28 | 16:54:25 | 29.526 | 372.38 | 27.91 | 5287 | 19.1 |  | 128.8237 | 26.9241 | 269.3004 | 201.9924 | 10.0685 | 314.6399
2010-07-28 | 16:54:35 | 29.693 | 374.396 | 27.85 | 5298 | 19.1 |  | 129.5213 | 27.07 | 270.7588 | 203.0864 | 10.123 | 316.3439
2010-07-28 | 16:54:45 | 29.86 | 374.059 | 27.86 | 5285 | 19.1 |  | 129.4047 | 27.0456 | 270.5151 | 202.9036 | 10.1139 | 316.0592
2010-07-28 | 16:54:55 | 30.027 | 377.105 | 27.77 | 5286 | 19.1 |  | 130.4586 | 27.2658 | 272.718 | 204.5559 | 10.1963 | 318.633
2010-07-28 | 16:55:05 | 30.194 | 377.105 | 27.77 | 5296 | 19.1 |  | 130.4586 | 27.2658 | 272.718 | 204.5559 | 10.1963 | 318.633
2010-07-28 | 16:55:15 | 30.36 | 372.045 | 27.92 | 5285 | 19.1 |  | 128.7078 | 26.8999 | 269.0582 | 201.8108 | 10.0594 | 314.357
2010-07-28 | 16:55:25 | 30.527 | 376.426 | 27.79 | 5273 | 19.1 |  | 130.2235 | 27.2167 | 272.2266 | 204.1873 | 10.1779 | 318.0589
2010-07-28 | 16:55:35 | 30.694 | 374.734 | 27.84 | 5275 | 19.1 |  | 129.6381 | 27.0944 | 271.0028 | 203.2694 | 10.1321 | 316.629
2010-07-28 | 16:55:45 | 30.861 | 376.426 | 27.79 | 5268 | 19.1 |  | 130.2235 | 27.2167 | 272.2266 | 204.1873 | 10.1779 | 318.0589
2010-07-28 | 16:55:55 | 31.028 | 378.81 | 27.72 | 5260 | 19.1 |  | 131.0484 | 27.3891 | 273.9511 | 205.4808 | 10.2424 | 320.0737
2010-07-28 | 16:56:05 | 31.195 | 376.16 | 27.84 | 5261 | 19 |  | 129.8706 | 27.143 | 271.5265 | 203.6622 | 10.1694 | 317.794
2010-07-28 | 16:56:15 | 31.362 | 378.199 | 27.78 | 5254 | 19 |  | 130.5748 | 27.2901 | 272.9986 | 204.7664 | 10.2245 | 319.517
2010-07-28 | 16:56:25 | 31.529 | 380.939 | 27.7 | 5254 | 19 |  | 131.5207 | 27.4878 | 274.9763 | 206.2498 | 10.2986 | 321.8317
2010-07-28 | 16:56:35 | 31.696 | 379.909 | 27.73 | 5246 | 19 |  | 131.165 | 27.4135 | 274.2327 | 205.692 | 10.2708 | 320.9613
2010-07-28 | 16:56:45 | 31.863 | 379.566 | 27.74 | 5245 | 19 |  | 131.0467 | 27.3888 | 273.9854 | 205.5065 | 10.2615 | 320.6718
2010-07-28 | 16:56:55 | 32.03 | 380.252 | 27.72 | 5241 | 19 |  | 131.2834 | 27.4382 | 274.4803 | 205.8777 | 10.28 | 321.2511
2010-07-28 | 16:57:05 | 32.197 | 385.093 | 27.58 | 5235 | 19 |  | 132.9549 | 27.7876 | 277.9749 | 208.4989 | 10.4109 | 325.3412
2010-07-28 | 16:57:15 | 32.363 | 383.009 | 27.64 | 5235 | 19 |  | 132.2355 | 27.6372 | 276.4708 | 207.3707 | 10.3546 | 323.5808
2010-07-28 | 16:57:24 | 32.514 | 384.397 | 27.6 | 5228 | 19 |  | 132.7146 | 27.7373 | 277.4725 | 208.122 | 10.3921 | 324.7531
2010-07-28 | 16:57:35 | 32.68 | 385.093 | 27.58 | 5231 | 19 |  | 132.9549 | 27.7876 | 277.9749 | 208.4989 | 10.4109 | 325.3412
2010-07-28 | 16:57:45 | 32.847 | 385.442 | 27.57 | 5223 | 19 |  | 133.0753 | 27.8127 | 278.2266 | 208.6877 | 10.4203 | 325.6357
2010-07-28 | 16:57:55 | 33.014 | 382.663 | 27.65 | 5213 | 19 |  | 132.116 | 27.6123 | 276.221 | 207.1834 | 10.3452 | 323.2885
2010-07-28 | 16:58:05 | 33.181 | 384.397 | 27.6 | 5212 | 19 |  | 132.7146 | 27.7373 | 277.4725 | 208.122 | 10.3921 | 324.7531
2010-07-28 | 16:58:15 | 33.348 | 387.541 | 27.51 | 5210 | 19 |  | 133.8001 | 27.9642 | 279.7421 | 209.8244 | 10.4771 | 327.4095
2010-07-28 | 16:58:25 | 33.515 | 386.84 | 27.53 | 5213 | 19 |  | 133.558 | 27.9136 | 279.2358 | 209.4447 | 10.4581 | 326.817
2010-07-28 | 16:58:35 | 33.682 | 388.596 | 27.48 | 5200 | 19 |  | 134.1643 | 28.0403 | 280.5035 | 210.3955 | 10.5056 | 328.3007
2010-07-28 | 16:58:45 | 33.849 | 387.19 | 27.52 | 5203 | 19 |  | 133.679 | 27.9389 | 279.4888 | 209.6344 | 10.4676 | 327.1131
2010-07-28 | 16:58:55 | 34.016 | 387.541 | 27.51 | 5197 | 19 |  | 133.8001 | 27.9642 | 279.7421 | 209.8244 | 10.4771 | 327.4095
2010-07-28 | 16:59:05 | 34.182 | 389.654 | 27.45 | 5199 | 19 |  | 134.5297 | 28.1167 | 281.2675 | 210.9685 | 10.5342 | 329.1948
2010-07-28 | 16:59:15 | 34.349 | 389.301 | 27.46 | 5191 | 19 |  | 134.4078 | 28.0912 | 281.0125 | 210.7773 | 10.5247 | 328.8964
2010-07-28 | 16:59:25 | 34.516 | 390.716 | 27.42 | 5184 | 19 |  | 134.8963 | 28.1933 | 282.0339 | 211.5434 | 10.5629 | 330.0918
2010-07-28 | 16:59:35 | 34.683 | 392.85 | 27.36 | 5180 | 19 |  | 135.633 | 28.3473 | 283.5742 | 212.6987 | 10.6206 | 331.8946
2010-07-28 | 16:59:45 | 34.85 | 389.654 | 27.45 | 5178 | 19 |  | 134.5297 | 28.1167 | 281.2675 | 210.9685 | 10.5342 | 329.1948
2010-07-28 | 16:59:55 | 35.017 | 390.362 | 27.43 | 5166 | 19 |  | 134.774 | 28.1678 | 281.7781 | 211.3516 | 10.5534 | 329.7925
2010-07-28 | 17:00:05 | 35.184 | 391.426 | 27.4 | 5174 | 19 |  | 135.1413 | 28.2445 | 282.5462 | 211.9277 | 10.5821 | 330.6914
2010-07-28 | 17:00:15 | 35.351 | 391.781 | 27.39 | 5153 | 19 |  | 135.2641 | 28.2702 | 282.8028 | 212.1201 | 10.5917 | 330.9917